## Gemba Requests
ID | Short Description | Contact / Organizer | Where (building/room #/lab)? | When did the event occur? | When was the event detected? | Expected Results | Based on process conf., Gemba 100% effective | Outcome of Gemba
6227 | 51NYA Gemba: annual vial shipper PM failed | adam.rowe@sanofi.com | 45NYA/101 | 2025-10-02 | 2025-10-02 | evaporation rate should be 0.045L/hr or less
6226 | 8NYA-Pompe- FGHAF012-VI Mixer 2 not agitating | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2025-10-01 | 2025-10-02 | It is expected that Mixer 2 weight remain within range to allow for automated mixing of contents (back up mixer
6225 | 76 NYA Biosurgery Fill/Finish - Improper running of Bowie Dick Test | brian.roy2@sanofi.com | 76 NYA/Room 352 | 2025-09-17 | 2025-10-01 | Bowie Dick Test is to be loaded and ran properly before running vacuum leak test and any other runs in the autoclave.
6224 | GEMBA for discoloration/spotting of ceiling caulk observed in 51 NYA Room 125 (Weigh Room) | megh.vakharia@sanofi.com | 51 New York Ave | 2025-10-01 | 2025-10-01 | TBD
6223 | 8NYA Gemba - 805010 batch FGHBQ004 failed pH testing. pH OOS. | jack.laplante@sanofi.com | 8NYA 1350 | 2025-10-01 | 2025-10-01 | Path forward
6222 | Aseptic Qual Turbidity Semiannual sample 2084055 marked turbid on day 7 and negative on day 14 | Stefanie.Erickson@sanofi.com | 68 NYA, QC Micro Counting bench | 2025-10-09 | 2025-10-01 | for there to be consistency at the day 7 and 14 read, or if a turbid sample was observed for there to be a note made and submitted for ID.
6221 | GEMBA; 8NYA: Fabrazyme Butyl: Equil going into Alarm | William.Bradley@sanofi.com | 8NYA 1323 | 2025-10-01 | 2025-10-01 | Recipe should run as intended without gaps.
6220 | Gemba Suite 1 Downstream Column Packing Pest Discovered during Room 1303 Maximum Occupancy Study | gary.dodakian3@sanofi.com | 8 nya room 1303 suite 1 | 2025-09-30 | 2025-09-30 | there should be no insects in the suite
6219 | 11 Forbes / Room 1315 Issues noted during Q3 GoGemba related to lack of good documentation practice, filing, chain of custody for chart recorder data, etc. General Lack of 5S and Good Housekeeping | John.sicurella@sanofi.com | 11 Forbes / Room 1315 | 2025-09-30 | 2025-09-30 | Expected that 11 Forbes / Room 1315 would comply with QU-SOP-0074684-Good Documentation Practices and have good 5S and housekeeping
6218 | 11 Forbes / Room 1311 Lives computer system used for Temperature related qualification work does not have appropriate security | JOhn.sicurella@sanofi.com | 11 Forbes / Room 1311 | 2025-09-30 | 2025-09-30 | Expected that the 11 Forbes / Room 1311 Lives computer system used for Temperature related qualification would comply with QU-SOP-0013232-Computer System Validation Life Cycle |  | Deviation
6217 | 8NYA - Support service - Violation of gowning entering solution prep PAL | Bailey.gleason@sanofi.com | 8NYA 1361 | 2025-09-30 | 2025-09-30 | Follow gowning procedure
6216 | WO 3402717 investigate WFI toc no flow caused by WFI GMP loop shutdown. | jesus.acosta2@sanofi.com | 45-181 | 2025-09-30 | 2025-09-30 | system should be running all the time with no issues. TOC should have flow all the time
6215 | 8NYA - Support service - DSP Butyl load filter with bleed valve broken | bailey.gleason@sanofi.com | 8NYA RM 1290 | 2025-09-30 | 2025-09-30 | filter not have a broken bleed valve
6214 | 8NYA - Support service - RF573F136 - 3000L sum bag failed inspection | bailey.gleason@sanofi.com | 8NYA 1350 | 2025-09-29 | 2025-09-29 | Hold Gemba to see path forward.
6213 | 8NYA - Support Service - RF573F136 - Failed pre use inspection for 1000L tote bag | bailey.gleason@sanofi.com | CNC | 2025-09-30 | 2025-09-30 | no punctures
6212 | 8NYA - Support Service - PH OOS RF1NHF015 P/N 805003 | bailey.gleason@sanofi.com | RM 1350 | 2025-09-29 | 2025-09-29 | Ph spec be within range
6211 | 8NYA GEMBA Downstream Suite 2 - Sartorius 2000L SUM Port Seals - Pre-Use Visual Inspection Failure | alex.tello@sanofi.com | 1503 | 2025-09-29 | 2025-09-29 | Per MABC-QU-SOP-0084572-Single Use Containers at 8 NYA (v10.0), Section 6.2.19, it is expected there are no voids on the port seals. It would not be expected to see a bubble or divot present on the port seals.
6210 | MES executed without training | Pratik.kharecha@sanofi.com | 8NYA/Room 1333 | 2025-09-21 | 2025-09-29 | Only trained operator should execute steps in MES record
6209 | Personal Phone was carried in Grade D Suite 1 | sushama.bhagat@sanofi.com | Room 1320, Suite 1, 8NYA | 2025-09-29 | 2025-09-29 | MABC-QU-SOP-0084643-Gowning procedure for entry into aseptic processing area at 8NYA (v4.0) Step 6.2.2 Only work-issued cell phones, radios or badges are allowed in the controlled areas.
6208 | 8NYA-No Product-Phase 2 Upstream-Tote containing unknown solution/ unlabeled discovered in vestibule Rm1404 | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1404 | 2025-09-28 | 2025-09-28 | it is expected that solutions prepared for use in Upstream GMP production be labeled according to the asssociated Solution Prep SOP
6207 | GEMBA - 8 NYA Suite Two DSP - Discrepancies in UF/DF-A SOP (Membrane Install Portion) | adam.keister@sanofi.com | 8 NYA room 1503 | 2025-09-27 | 2025-09-27 | Alignment between SOP, MES, and physical skid. | Yes | No Deviation
6206 | Pompe, upstream, RF1NAF012 SUT SUM Bag p/n: 692168 fail visual inspection with embedded particulate | marshall.phan@sanofi.com | 1403 | 2025-09-27 | 2025-09-27 | no particulates to found for bag to pass inspection
6205 | Failed Pre-Use SUT from Fabrazyme Butyl (RF1BRF051) | william.bradley@sanofi.com | 8NYA /RM1323 | 2025-09-27 | 2025-09-27 | SUT is to be in good standing condition without deformities
6204 | ANX CIP Buffer Manifold setup | charles.salgadinho@sanofi.com | 8NYA, 1503 | 2025-09-27 | 2025-09-27 | The Standalone CIP records, should state the correct manifold ports to scan, so the buffer can go into the skid |  | Deviation
6203 | TSH - FGF3B04 - 51 NYA - Filter test performed under incorrect username | MARIE.VALLIERE@SANOFI.COM | 51 NYA, RM 140 | 2025-08-28 | 2025-09-27 | It is expected that users are signed in to machines they are using.
6202 | 8NYA-Pompe- Automation Intervention needed for MFC control post X2250 Inoculation | john.lytle@sanofi.com | Rm 1403, 8NYA | 2025-09-26 | 2025-09-26 | It is expected the MFC functions as intended
6201 | WFI loop temp too high during UF/DF-A CIP | David.Loik@sanofi.com | 8NYA | 2025-09-26 | 2025-08-26 | Recipe would not have a tight temperature range for the WFI loop, or wouldn't look at the loop at that phase. |  | No Deviation
6200 | 8 NYA Support Services - Glycan (657833) TOR tracking | nicholas.grote@sanofi.com | 08-1421 and 08-1364 | RF1NJE002 - 5/1/2024, RF1NJE003 - 5/6/2024, and RF1NJE004 - 7/2/204 | 2025-09-25 | If limited time out of refrigeration is allowed the time must be tracked to ensure the limit is not reached.
6199 | Viable cells did not meet criteria | valeria.moreno@sanofi.com | 8NYA-1403 | 2025-09-26 | 2025-09-26 | Take a new sample
6198 | Viable cells did not meet criteria | valeria.moreno@sanofi.com | 8NYA-1403 | 2025-09-26 | 2025-09-26 | Take a new sample
6197 | 8 NYA Suite 1 Downstream Bag Failed Pre-Inspection | Brian.Phan@Sanofi.com | 8 NYA, Room 1323 | 2025-09-25 | 2025-09-25 | SUT free from all particles or defects.
6196 | Wrong date recorded in Batch record. | Luis.dominguez@sanofi.com | 51NYA room 138 | 2025-09-07 | 2025-09-25 | Date should be the date of occurrence | Yes | Deviation
6195 | Rationale for worst-case soil for clarified harvest loaded onto the multi-product Suite 2 MCC skids is not documented QU-OPE-0250476-Soil matrix rationale for 8 New York Ave Manufacturing Facility | nicholas.lapointe@sanofi.com | 8NYA suite 2 upstream | 2021-08-20 | 2025-09-16 | Rationale for all worst-case soils should have been documented in soil matrix report QU-OPE-0250476, or the NeoGAA cleaning validation project plan should have required MCC skid cleaning PQ to be performed using NeoGAA harvest material
6194 | QE-1544328 already opened- P/N:813907 Batch: 25GN00996 Flexsafe Pro Mixer Bag 200L Failed inspection | brian.madigan@sanofi.com | Suite 1 DS | 2025-09-21 | 2025-09-21 | Bag is expected to be particle free
6193 | Form coming from Zinc drain in CNC area from US suite 2 | dennis.bazinet@sanofi.com | 8NYA CNC / 1210 | 2025-09-24 | 2025-09-24 | No forming at the drain
6192 | Rationale for worst-case soil for clarified harvest loaded onto the multi-product Suite 2 MCC skids is not documented QU-OPE-0250476-Soil matrix rationale for 8 New York Ave Manufacturing Facility | nicholas.lapointe@sanofi.com | 8NYA suite 2 upstream | 2021-08-20 | 2025-09-16 | Rationale for all worst-case soils should have been documented in soil matrix report QU-OPE-0250476, or the NeoGAA cleaning validation project plan should have required MCC skid cleaning PQ to be performed using NeoGAA harvest material
6191 | GEMBA to discuss switching the controlling pH probe in the VI-Skid X-2765 | Bhargav.Patel@sanofi.com | Room 1403 | 2025-09-24 | 2025-09-24 | We expect the active vi probe to function properly
6190 | PH out of range During SP1 Elution Step | Luis.dominguez@sanofi.com | 51NYA room 138 | 09/15/2025 and 09/19/2025 | 2025-09-24 | per QU-SOP-0061107 attachment 3, Ph should stabilize from 4.8-5.6 and during the batch record review it was discovered that the PH dropped to 4.40 on both lots. |  | No Deviation
6189 | Incorrect Probe installation during setup | Pavel.dobrovolskiy@sanofi.com | 8RM1323 | 2025-08-14 | 2025-09-24 | its expected that during skid setup the correct probes will be installed into the correct locations
6188 | 51-NYA: PM on GMP critical equipment completed after scheduled time | ines.bica@sanofi.com | 80NYA | 2025-09-23 | 2025-09-24 | PM to be performed on time |  | Deviation
6187 | Manifold tube connection (CNC) not signed for | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2025-09-23 | 2025-09-23 | Step will be signed for before beginning a process on the skid, per SOP-0084514(V2.0) - step 6.8.2
6186 | 51NYA Annual PM not assigned to Global Cell Banking MVE Cryoshippers | Adriana.Alban@sanofi.com | 45-101 | 2022-12-18 | 2025-09-23 | It is expected that an annual Preventative maintenance be performed on the MVE cryoshippers. | Yes | Deviation
6185 | 8NYA Gemba - Support services - 705667 1.0 Sodium Carbonate / RF1NVF014, FGHBW014 barrels incorrectly labeled | JACK.LAPLANTE@SANOFI.COM | 8NYA 1350 | 2025-09-22 | 2025-09-23 | permission to remove old labels from the barrels and keep the temporary correctly written labels
6184 | GEMBA for Engineering best practices on drawing approval for SUT | Benjamin.Larson@sanofi.com | NA | 2025-09-19 | 2025-09-23 | N/A
6183 | Quarterly Badge Access Audit for QC Raw Materials is unable to be completed per QU-SOP-0074027-Facility Security and Access Control for the Northborough Facility | amy.roy@sanofi.com | 11 Forbes/ QC Raw Materials | 2025-09-23 | 2025-09-23 | Quarterly Badge Access Audit for QC Raw Materials is unable to be completed per QU-SOP-0074027-Facility Security and Access Control for the Northborough Facility, step 6.8.3- there are no training requirements listed in the referenced form nor SOP for the QC Raw Materials Lab
6182 | Suite 2 - Train 2 - BGA did not show PH value, only C02 and PC02 | joseph.gonzalez@sanofi.com | 1403, 8NYA | 2025-09-23 | 2025-09-23 | BGA to show all results, PH, C02, and O2.
6181 | Manual mixing SOP does NOT include POMPE nor Sartorius SUMs | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2025-09-22 | 2025-09-22 | Sartorius SUMs will have its own section (along with Millipore and Thermo) and there will be RPM settings table for Pompe product
6180 | ANX Blank Run; Eluate-- No Mixing Recipe in MES | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2025-09-22 | 2025-09-22 | A typical process record will have a step where an operator can send a mixing recipe to DeltaV.
6179 | Pompe Cedex sample need to re-run | thanhtruc.le@sanofi.com | 1403 | 2025-09-22 | 2025-09-22 | Have enough sample to re-sample
6178 | 51NYA gemba: LOTO not used as 'group LOTO' | adam.rowe@sanofi.com | 51NYA, 140 | 2025-08-25 | 2025-09-22 | Group LOTO should have been selected | Yes | Deviation
6177 | 51NYA Gemba: cooler sample log was used when out of date | adam.rowe@sanofi.com | 45NYA/1111 | 2025-09-19 | 2025-09-20 | The new version should have been used
6176 | GEMBA: Failed SUT Suite 1 P/N:813907 | William.Bradley@sanofi.com | 8NYA /RM1323 | 2025-09-21 | 2025-09-22 | It is expected that bags are in good integrity with no particulates |  | Deviation
6175 | Brief loss of 76-AHU 2A | JAKE.STANTON@SANOFI.COM | 76-MEZZ 200 | 2025-09-19 | 2025-09-19 | Action notice created per QU-SOP-0022403
6174 | Brief loss of 76-AHU 2A | JAKE.STANTON@SANOFI.COM | 76-MEZZ 200 | 2025-09-19 | 2025-09-19 | Action notice created per QU-SOP-0022403
6173 | 8NYA GEMBA - Support Services - Incorrect batch size for 705667 | JACK.LAPLANTE@SANOFI.COM | 8RM1350 | 2025-09-21 | 2025-09-21 | Path forward
6172 | 68NYA-Metrology-No product impact-Metrology standard S-1740-001 Pressure Module out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2025-09-08 | 2025-09-22 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
6171 | Brief loss of 76-AHU 2A | JAKE.STANTON@SANOFI.COM | 76-MEZZ 200 | 2025-09-19 | 2025-09-19 | Action notice created per QU-SOP-0022403
6170 | 8NYA - Support Services - Incorrect batch size. | William.Lee@sanofi.com | 8NYA/ 1350 | 2025-09-21 | 2025-09-21 | Batch size must be 950L
6169 | No lot associated 250L SUM Bag | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-09-21 | 2025-09-21 | SUM Bag should be free of particles
6168 | No lot associated ATF | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-09-21 | 2025-09-21 | ATF10s should be free of particles
6167 | Storage condition discrepancy for Eluate retain samples for Pompe | adam.hildebrand@sanofi.com | 8NYA RM 1403 | 2025-09-21 | 2025-09-21 | It is expected that the storage condition for the sample retains would be the same in both the sample plan SOP and BioLIMS.
6166 | 8NYA-Gemba-Downstream Suite 2 Pompe-Copper Water Run Overview | aidan.mula@sanofi.com | 8RM1503 | 2025-09-19 | 2025-09-19 | Process should run as defined by control strategy
6165 | Deviation Intake breakdown 08AUG25 | brandon.haffner@sanofi.com | 1113 | 2025-08-08 | 2025-09-18 | Deviations are required to be opened 24 hours after the GEMBA
6164 | slices from someone doing some cutting in the floor by MAL 1204 | dennis.bazinet@sanofi.com | 8 NYA CNC space by MAL 1204 | 2025-09-19 | 2025-09-19 | NO slices cut into the floor
6163 | Black box for HOF shelves have storiform stored in them in the clean room | dennis.bazinet@sanofi.com | 8 NYA CNC space Phase 2 downstream | 2025-09-19 | 2025-09-19 | No storiform to be inside the suite
6162 | Dry residue on the walls at 8NYA CNC space | dennis.bazinet@sanofi.com | 8 NYA CNC space | 2025-09-19 | 2025-09-19 | Walls to be residue free
6161 | SP transfer timer exceeded about 20 min during SP run on 18SEP25 | christian.venegas@sanofi.com | 8 NYA Suite 1 Downstream room 1323 | 2025-09-18 | 2025-09-18 | To get the Phath forward and confirm the product is safe
6160 | SP transfer timer exceeded about 20 min during SP run on 18SEP25 | christian.venegas@sanofi.com | 8 NYA Suite 1 Downstream room 1323 | 2025-09-18 | 2025-09-18 | To get the Phath forward and confirm the product is safe
6159 | 8NYA, Suite1, SP Maximum pH hold time exceeded 210minutes | Robert.pino@sanofi.com | 8NYA, Suite1, 1323, SP Chromatography | 2025-09-18 | 2025-09-18 | It is expected that the SP eluate would be finished mixing and adjusted before the 210 minute hold time has passed.
6158 | GEMBA to discuss movement between CNC corridor and room 127 at 51NYA | Sylvester.Kegbeh@sanofi.com | 51 NYA ROOMS corridor 82 | 2025-09-17 | 2025-09-17 | Refer to QU-SOP-0006952
6157 | 8NYA-Pompe-Vaccuum Pump Failure | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2025-09-18 | 2025-09-18 | It is expected the Vaccuum pump remain running through ATF process use
6156 | 8NYA-Pompe- FGHAF012- ATF B Low Flow | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2025-09-18 | 2025-09-18 | it is expected that ATFs maintain flowrates based off CST allowances
6155 | Insulated chilled water lines are shedding protective coating. | joel.henchar@sanofi.com | 8 NYA, all rooms utilizing these insulated hoses. | 2025-09-17 | 2025-09-17 | N/A
6154 | GEMBA: Suite 1: SP: Viral Inactivation | William.Bradley@sanofi.com | 8NYA /RM1323 | 2025-09-18 | 2025-09-18 | Recipe should resume to a 15 minute mix after VI addition
6153 | Password identified on sticky notes on engineering laptop | James.Renzullo@sanofi.com | 8NYA, Suite 2 | 2025-09-16 | 2025-09-16 | Not to use stickynotes as per MABC-QU-SOP-0074684-Good Documentation Practices |  | Deviation
6152 | GEMBA: Suite 1 Fabrazyme: Butyl: Low Low alarm | william.bradley@sanofi.com | 8NYA 1323 | 2025-09-17 | 2025-09-17 | Process should run without issues or alarms
6151 | Dedicated product contact equipment stored in CNC Area | Kelley.Gibson@sanofi.com | 8NYA-CNC1211 Corridor | 2025-09-17 | 2025-09-17 | No storing product contact dedicated equipment in CNC areas.
6150 | 8NYA Gemba - Support Services - breakdown step missing from EBR | jack.laplante@sanofi.com | 8NYA RM1350 | 2025-09-17 | 2025-09-17 | get a breakdown step added to the EBR
6149 | 8NYA - Support Service- EMS alarm for HH room 1350 SOL prep to Neut DP | Bailey.gleason@sanofi.com | RM1350 | 2025-09-16 | 2025-09-16 | N/A
6148 | 8NYA- Gemba - Downstream suite 2 - Pompe skid leak | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 room 1503 | 2025-09-16 | 2025-09-16 | Adjust triclamp connection restart cleaning recipe |  | No Deviation
6147 | 8NYA Gemba - downstream suite 2 - Pompe prepacked column valves do not align with SOP | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 room 1503 | 2025-09-16 | 2025-09-16 | Need to identify correct valves to attach to sid
6146 | 8NYA Gemba - downstream suite 2 - Pompe prepacked column prime record not capturing previous cycles | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 room 1503 | 2025-09-16 | 2025-09-16 | record need to identify previous cycles on the column |  | No Deviation
6145 | 68NYA-Support APU-Room 4700-QCM-product impact unknown-Temperature Indicator F-6980-001 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-4700 | 2025-09-12 | 2025-09-16 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
6144 | Pompe - RF1NAF012 - 8NYA - Train 1 DO MFC IO input failure | michael.knapp@sanofi.com | 8NYA 1403 | 2025-09-15 | 2025-09-15 | MFC operates as intended
6143 | 8NYA - Support Service - RF1YIF053 3k SUM bag failed inspection | bailey.gleason@sanofi.com | 8NYA | 2025-09-15 | 2025-09-15 | SUT be ready to be used by MFG with no issues
6142 | Media Valve close during Top off | 8NYAUpstream1@sanofi.onmicrosoft.com | Room 1403 | 2025-09-15 | 2025-09-15 | Media valve to stay opened when the recipe is on Hold | Yes | Deviation
6141 | 8NYA-DS/CNC-Column packing. unable to do a WFI take out due to expiry. | genesis.torres2@sanofi.com | 8nya RM1210 | 2025-09-15 | 2025-09-15 | Expected result is for WFi not to go expired within the timeframe of operations.
6140 | 76Nya Dewpoint alarm | andy.le@sanofi.com | 76Nya | 2025-09-15 | 2025-09-15 | MABC-QU-SOP-0022403-Action Notification Procedure
6139 | SUT Upstream RF1NAf012 ATF failed pre-use visual Inspection with particulate | jeffrey.boldizar@sanofi.com | 8NYA1403 | 2025-09-14 | 2025-09-14 | no particulates present | Yes | Deviation
6138 | 8NYA-Pompe-FGHAF011-X2230 pH Controller Hi Hi Alarm | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2025-09-14 | 2025-09-14 | pH controller function within process alarm limits
6137 | Pompe upstream RF1BGF019 SIP record Takeout not available when remaining quantity is acceptable | jeffrey.boldizar@sanofi.com | 8NYA1403 | 2025-09-14 | 2025-09-14 | Availability to takeout addition items to the BOM
6136 | Unable to Print Harvest 1 Labels | jonathan.lucas@sanofi.com | 51NYA, rm 140 | 2025-09-14 | 2025-09-14 | WIP labels should print for H1.
6135 | CIP record missing SUT and WFI tote | david.loik@sanofi.com | 1503 | 2025-09-13 | 2025-09-13 | MES would have required information and takeouts
6134 | GEMBA for LAL Sample 2073348 Cracked Sample Tube-8 NYA Pompe PN 805012 Lot FGHSD008 | roxanne.vacha@sanofi.com | 68 NYA, 4700, QC Micro | 2025-09-13 | 2025-09-13 | Per QU-SOP-0070558 Step 6.2.3, confirm sample containers are not cracked, leaking. Expectation is that sample tubes are not cracked.
6133 | 8NYA Gemba - Support services - BOM missing pump tubing. | JACK.LAPLANTE@SANOFI.COM | 8NYA 1350 | 2025-09-13 | 2025-09-13 | get pump tubing added to formulation BOM
6132 | Material revived from 11Forbes mislabeled | michael.montiverdi@sanofi.com | 8NYA | 2025-09-13 | 2025-09-13 | Material received from 11 Forbes with the correct material ID tag
6131 | 51-NYA: Cracked column tube on Q-Column | ines.bica@sanofi.com | 51-132 | 2025-09-08 | 2025-09-08 | Column Tube should be fully intact
6130 | 8NYA-No Product- L&D received uncontrolled version of MABC-QU-FOR-0152676 for proof of training | JOHN.LYTLE@SANOFI.COM | 8NYA- General | 2025-09-09 | 2025-09-11 | It is expected that controlled versions of document be submitted for updates to curricula
6129 | 8NYA GEMBA Pompe media lot FGHMB008 detecting as expired in MES despite expiry being 22SEP25 | jacob.roman@sanofi.com | 8RM1343 | 2025-09-12 | 2025-09-12 | Media that aren't expired should be usable
6128 | Pompe, FGHAF012, 8 NYA, Upstream, Pompe X2220 MFC not able to maintain %DO | eric.horten@sanofi.com | 8NYA RM1403 | 2025-09-12 | 2025-09-12 | Oxygen is added to the SUB via Mass Flow Controller (MFC) connected to the micro (MFC-007) and macro (MFC-008) sparge lines. The DO requirement for bioreactor N-1 stage is 40 – 100% per MABC-QU-OPE-0522556-Control Strategy for Pompe Integrated Continuous Biomanufacturing (Pompe-ICB) Pr (v10.0) Table 8-5. | Yes | Deviation
6127 | 51 NYA - TSH - Cell Culture - Incorrect time entry on Portable Media Storage and Usage Log | MARIE.VALLIERE@SANOFI.COM | 51 NYA, room 147 | 2025-08-02 | 2025-09-12 | It is expected that the time is accurately entered on the Portable Media Storage and Usage Log.
6126 | 8NYA Gemba Support services urgent - 2k tote scanning as 3k tote | jack.laplante@sanofi.com | 8NYA CNC | 2025-09-12 | 2025-09-12 | Path forward | No | No Deviation
6125 | 8NYA Gemba Support Services - Missing DeltaV prompt, repeat of QE-1543609 | jack.laplante@sanofi.com | 8NYA 1350 | 2025-09-12 | 2025-09-12 | correct these parameters
6124 | TSH Gemba: Warm Room 148 Temp OOR | cindy.guerra@sanofi.com | 51-148 | 2025-09-11 | 2025-09-11 | Temperature would stay within range. | Yes | No Deviation
6123 | 8NYA -RF1N0F002 - exceeded 1% tolerance of Hex glycol charge | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-09-11 | 2025-09-11 | Recipe driven material charges be correct
6122 | 8NYA - Support Service - Prior to connection it was noticed the connection on the filter was not aligned and it fell off | Bailey.gleason@sanofi.com | 8NYA | 2025-09-11 | 2025-09-11 | SUT be ready to be used by MFG with no issues | Yes | Deviation
6121 | 8NYA Support Services - 805007 Wrong value used for glycan purity | jesse.welch@sanofi.com | 8NYA | 2024-09-03 | 2025-09-11 | Purity value would be defined. | No | Deviation
6120 | X-2543 Bubble Trap causing holds during buffer prime | Dana.hopkins@sanofi.com | 8NYA room 1333 | 2025-09-11 | 2025-09-11 | Recipe should run with no holds
6119 | 2K tote BHT-1584001 when scanned says it's a 3K tote in MES | dennis.bazinet@sanofi.com | 8 NYA / CNC space | 2025-09-11 | 2025-09-11 | Have MES put in the correct information in MES for a 2K
6118 | critical observation (Critical GMP Deficiencies in Contamination Control) for Avantor Paris, KY (supplier) | Nashmin.Hojatti@sanofi.com | n/a | 2025-09-11 | 2025-09-11 | TBD
6117 | 8NYA-US-SS(CNC)-Expired Media. | Genesis.torres2@Sanofi.com | 8nya CNC | 2025-09-11 | 2025-09-11 | Path forward
6116 | GEMBA-No Product -No Lots-EandM 8NYA-Unable to locate IV/IQ walkdown reports for three BVs on chilled water distribution. | jose.rodriguez@sanofi.com | 8NYA |  | 2025-09-10 | These BVs should have been verified during commissioning and qualification during an IV/IQ walkdown
6115 | 8NYA - Support Service - Pompe bioreactor probes autoclave with incorrect test probes | Bailey.gleason@sanofi.com | 8NYA | 2025-09-11 | 2025-09-11 | Probe load be set up correctly
6114 | X2230 PH Probe Deviation | sean.gormley@sanofi.com | 8RM1403 | 2025-09-10 | 2025-09-10 | Probes should not become fouled during course of run.
6113 | Incorrected exp. entered | alexandre.silva2@sanofi.com | 76 NYA APU | 2025-04-22 | 2025-09-05 | EXp. date on ERP should match the COA.
6112 | Incorrected exp. entered | alexandre.silva2@sanofi.com | 76 NYA APU | 2025-04-22 | 2025-09-05 | EXp. date on ERP should match the COA.
6111 | Multiple products - 8 NYA - WFI used prior to release of Action Notice 2025-I-119 | Jeremy.Wooden@Sanofi.com | 8 NYA | 2025-09-09 | 2025-09-10 | That the Action Notice would be released prior to WFI usage
6110 | FZ Suite 1 Upstream Interim Buffer Temperature Control 705686 Needed for Use | gary.dodakian3@sanofi.com | 8 nya room 1333 suite 1 | 2025-09-10 | 2025-09-10 | Buffer should be released for buffer prime so it can be scanned in
6109 | 76NYA - Incorrect version of paperwork used for two manufacturing processes | megan.botti-marino@sanofi.com | 76 NYA | 2025-08-27 | 2025-09-10 | Per MABC-QU-SOP-0074684-Good Documentation Practices (v8.0), the working copy block is signed by the individual who is verifying that the effective version of the paperwork is being used.
6108 | 3392963. Investigate WFI TOC loop 4 not starting after reboot. 8 NYA WFO LOOP 4 TOC ANALYZER. | jesus.acosta2@sanofi.com | 08-1170 | 2025-09-09 | 2025-09-09 | TOC analyzer on Loop 4 should be running as expected with no issues.
6107 | MFG Cleaning record Incomplete. | Dana.Hopkins@sanofi.com | 8 NYA RM 1333 | 2025-05-23 | 2025-09-10 | MES Team To close out record. QE To be generated.
6106 | x2141 Wave reactor OCP Communication failure. | ricardo.rodriguez@sanofi.com | 8 NYA RM 1403 | 2025-09-09 | 2025-09-09 | Automation to clear alarm.
6105 | EMS for CO2 incubator, F-16140, offline and out of specification. | heidy.ruano@sanofi.com | 68 New York Ave, QC Micro, 4735 | 2025-09-09 | 2025-09-09 | It is expected that the EMS probes are reading at all times and that the CO2 levels remain within range.
6104 | 8NYA- Support Services Gemba-805012 media stored in incorrect cold room | william.lee@sanofi.com | 8NYA/CNC | 2025-09-09 | 2025-09-09 | 805012 media should be placed in correct cold room
6103 | 8NYA-Pompe FGHAF012-Base tote found to have SUT incorrectly installed | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2025-09-09 | 2025-09-09 | It is expected that the elbow on the 200L bags face towards the barrel outlet
6102 | 76 NYA - Room 507 (Corrugate Room) - Seprafilm - Good Documentation Practices - Observation from Quarterly Walkthrough | yamini.patcha@sanofi.com | 76 - 507 | 2025-09-09 | 2025-09-09 | N/A
6101 | URGENT Pompe 50L wave X2163 for production lost communication with DeltaV | jacob.roman@sanofi.com | 8RM1403 | 2025-09-09 | 2025-09-09 | Expected the wave does not disconnect from OPC server
6100 | Filter Integrity Testing failure | julie.auguste@sanofi.com | Solution prep, 1290 | 2025-09-09 | 2025-09-09 | Expected FIT result is PASS per MABC-QU-SOP-0084585-Filter Preparation and Testing (v4.0)
6099 | Gemba – 8NYA – Prefiltration timer observed at start of 805012 batch | william.lee@sanofi.com | 8NYA/1350 | 2025-09-08 | 2025-09-08 | Alarm should not have triggered during batching.
6098 | Gemba – 8NYA – Leak noticed on SUT during buffer preparation | william.lee@sanofi.com | 8NYA/1350 | 2025-09-08 | 2025-09-08 | SUM bag should not have leaked
6097 | Drying room temperature malfunction | Kenneth.morgan@sanofi.com | 76-234 | 2025-09-05 | 2025-09-05 | It is expected that the temperature would be 55C +/- 5C
6096 | 8 NYA Gemba - Support Services - 805012 Bottle vent filters not crimped off | Nicholas.grote@sanofi.com | 08-1220 | 2025-12-07 | 2025-09-08 | Vent filters are crimped off bottles during the manifold breakdown.
6095 | FZ SUB X-2200 Hi Hi SUB Weight alarm / Low Low & Hi Hi ROC Alarm, due to media tote change and kink in media line | elon.elzam@sanofi.com | Phase 1 Upstream room # 1333 | 2025-09-08 | 2025-09-08 | MSAT to provide assessment. | Yes | No Deviation
6094 | 68NYA Support APU-Metrology-Room 1830-no potential product impact-Metrology standard FBS-0098 a Digital Aerosol Photometer found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2025-08-14 | 2025-09-08 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
6093 | 8NYA - Support Service - RF1NLF007 Perform fill was not executed contemporaneously | bailey.gleason@sanofi.com | 8NYA | 2025-09-07 | 2025-09-07 | steps be signed contemporaneously
6092 | 8NYA - Support Service - Cleaners carts with non GMP writing | bailey.gleason@sanofi.com | 8NYA phase 1 hallway | 2025-09-07 | 2025-09-07 | N/A
6091 | 500L SUB bag inspection failure due to free-flowing particulates inside bag | adam.hildebrand@sanofi.com | 8NYA RM 1403 | 2025-09-07 | 2025-09-07 | It is expected that the 500L SUB bag would not have any particulates inside.
6090 | Phase 1 Upstream Fz SUB X-2210 perfusion start issues | dana.hopkins@sanofi.com | 8NYA 1333 | 2025-09-06 | 2025-09-06 | Perfusion should have started without holds
6089 | 8NYA - Support Service - Possible violation of gowning procedure | Bailey.gleason@sanofi.com | 8NYA 1350 | 2025-09-07 | 2025-09-07 | N/A |  | Deviation
6088 | 8NYA - Support Service - RF1NLF007 Error message with scale for liquid weigh out | maria.palmar@sanofi.com | 8NYA / 1350. | 2025-09-07 | 2025-09-07 | Record allows operators to perform liquid weigh out
6087 | Percent viability lower than specification on Day 5 sample for 500L SUB on POMPE | adam.hildebrand@sanofi.com | 8NYA RM 1403 | 2025-09-07 | 2025-09-07 | It is expected that the culture viability to remain in specification. |  | No Deviation
6086 | Incorrect samples were taken during butyl end of hold sampling | Pavel.dobrovolskiy@sanofi.com | 8NYA 1323 | 2025-09-06 | 2025-09-06 | Its expected that the correct samples will be taken per FZ sample plan
6085 | 8NYA - Support Service - RF1NLF007 Error message with scale for liquid weigh out | Bailey.gleason@sanofi.com | 8NYA | 2025-09-07 | 2025-09-07 | Scale be compatible with MES
6084 | Incorrect Biolims labels used for MCC end holdsampling | christian.venegas@sanofi.com | 8NYA upstream MCC eluate | 2025-09-05 | 2025-09-06 | To get path forward to get Butyl elution samples
6083 | Incorrect Biolims labels used for MCC end holdsampling | christian.venegas@sanofi.com | 8NYA upstream MCC eluate | 2025-09-05 | 2025-09-06 | To get path forward to get Butyl elution samples
6082 | 8NYA - Support Services - Batch overshoot expected weight of 3000L, had high, high alarmm | Evans.kamau@sanofi.com | 8NYA1350 | 2025-09-06 | 2025-09-06 | Automation to acknowledge alarm and final qs to proceed, test and transfer is within spec, gemba to be held.
6081 | 8NYA - GEMBA - Downstream P2, Decrease between MES and DeltaV on buffer ports | Alex.tello@sanofi.com | 8NYA RM1503 | 2025-09-06 | 2025-09-06 | for MES to give you the correct ports to connect to.
6080 | 8NYA- Support Service- Pompe VI probes autoclaved incorrectly | Bailey.gleason@sanofi.com | 8NYA RM 1350 | 2025-09-06 | 2025-09-06 | Probes be autoclaved without compromising the sterility
6079 | X2230 pCO2 probe selection deviation alarm not enabled | adam.hildebrand@sanofi.com | 8NYA RM 1403 | 2025-09-05 | 2025-09-05 | It is expected that the correct parameter is loaded with recipe at beginning of run.
6078 | Inadequate Gowning | beau.mccready@sanofi.com | 51-140 | 2025-09-04 | 2025-09-04 | Proper Gowning Procedure Followed
6077 | Column Packing SOP is missing Axichrom ASM setting for ASM Columns | pavel.dobrovolskiy@sanofi.com | 8NYA 1303 | 2025-09-05 | 2025-09-05 | its expected that the SOP will provide information for every column thats being worked on
6076 | 8NYA APU-room 1503-No product impact-UV Indicators FB-4023-006 and FB-4023-016 failed calibration on MEMBRANE FILTRATION SKID X-3332001 | robert.bauch@sanofi.com | 8-1503 | 2025-09-05 | 2025-09-05 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results |  | Deviation
6075 | Sweco nitrogen cool down chart recorder failed to recordproperly | Kenneth.morgan@sanofi.com | 76-3103 | 2025-09-04 | 2025-09-05 | The chart should legibly record all data
6074 | 51NYA Thyrogen GMP NCRIT clock exceeded calibration due date | Adriana.Alban@sanofi.com | 51-135 | 2025-09-01 | 2025-09-02 | It is expected that portable equipment is calibrated by it's due date or have an calibration extension request completed prior to the cal due date.
6073 | Butyl buffer set up for Blank Run needed to be used for Ops run | caroline.depaula2@sanofi.com | 8NYA/room 1323/ Butyl Skid X3001 and CNC space | 2025-09-05 | 2025-09-05 | Maintain the CNC setup as is, considering that a breakdown is not possible since it’s the last available line. Close the blank run record as is, given that all CNC steps are identical to those in an operational run. Re-do the skid setup using the operational run MO RF1BRF046, and reference the CNC setup in the record. | Yes | No Deviation
6072 | 51 NYA TSH incorrect version of QU-FOR-0115149 used for vial transfer from 45NYA to 51NYA | Nicole.lufkin@sanofi.com | 51 NYA room 152 | 2025-09-04 | 2025-09-05 | it is expected that the current version of the logbook be used
6071 | Daily WFI Work Orders have not been tech completed at time of action | paul.mclaughlinjr@sanofi.com | 8NYA | 2025-09-04 | 2025-09-04 | Work orders are to be updated contemporaneously
6070 | pH control task started before pH 1 point and pCO2 1 point | jared.watson@sanofi.com | 8RM1333 | 2025-09-04 | 2025-09-04 | 1point to be performed prior to pH control task
6069 | Vendor expiration labels expired for drive mixers (P1 DS); metrology labels are fine. | elias.mcdonnell@sanofi.com | 8NYA, 1323 | 2025-09-03 | 2025-09-03 | Calibration will be verified and made sure that it does not expire before the end of the process.
6068 | Gas alarm on wave X2142, low low gas flow rate | Jacob.Roman@sanofi.com | 8RM1403 | 2025-09-03 | 2025-09-03 | That gas flow rate stays constant at 0.25SLPM |  | No Deviation
6067 | Thyrogen - FGFBPB006 - 51 NYA - Downstream - Unexpected result from the pooling formulated bulk | ines.bica@sanofi.com | 51NYA | 2025-09-02 | 2025-09-02 | Expected weight would be around 38-37kg. | No | No Deviation
6066 | 8NYA-Weigh Booth and CNC area flooring flakes | Daniel.bati@sanofi.com | 8NYA, CNC, Solution prep Weigh booth | 2025-09-02 | 2025-09-03 | no flooring Flakes left after work
6065 | Cryostats were not enrolled in the RQ Program in a timely manner | stephen.bohenko2@sanofi.com | Refer to table in Associated Documentation | 2023-12-01 | 2025-08-29 | Cryostats should have been enrolled in the RQ Program in a timely manner and should have been executed by 30NOV23.
6064 | Checklist from Quarterly GEMBA Not Submitted on Due Date | tohkti.wayet@sanofi.com | 76 NYA and 51 NYA | 2025-08-29 | 2025-09-02 | Per MABC-QU-SOP-0075222, Step 6.2.3, the Quarterly GEMBA should have been completed before the end of the month it occured.
6063 | Checklist from Quarterly GEMBA Not Submitted on Due Date | tohkti.wayet@sanofi.com | 76 NYA and 51 NYA | 2025-08-29 | 2025-09-02 | Per MABC-QU-SOP-0075222, Step 6.2.3, the Quarterly GEMBA should have been completed before the end of the month it occured.
6062 | 8NYA-CNC-Sol prep-flooring flakes found during a walkthrough | daniel.bati@sanofi.com | 8NYA/ CNC/ Sol prep | 2025-09-02 | 2025-09-02 | no dust or flooring flakes to be around
6061 | Pompe - RF1NAF011 - 8NYA - Liquid identified in micro sparge tubing | michael.knapp@sanofi.com | 8NYA 1403 | 2025-09-02 | 2025-09-02 | No liquid in the sparge line | Yes | Deviation
6060 | Investigate RO Main low tank level. TOC RO triggered alarm also. Metrology notified no RODI water while performing flow test in suite room 190. | jesus.acosta2@sanofi.com | 45-181 | 2025-09-02 | 2025-09-02 | Tank should 've keep up with RO/DI usage.
6059 | 8NYA APU-Room 1350-4 GMP Critical instruments went overdue for calibration | robert.bauch@sanofi.com | 8-1350 | 2025-09-01 | 2025-09-02 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
6058 | Pompe, FGHMD046, FBLBIOGAA-000000368- Pompe bioburden plates were left in 20-25C incubation for 5 days. Pompe plates incubate 3-4 days at 20-25C and are then moved into our 30-35 incubator F-15594 for 48-72 hours. The plates were moved in the onelims tracker but not physically moved into 30-35C incubation. | kathryn.lentoni@sanofi.com | 68NYA 4700, QCM | 2025-08-30 | 2025-09-01 | For plates to be removed on day 3 from 20-25C and then incubated in 30-35 for 48-72 hours
6057 | 8NYA Gemba- Downstream, Recipe trying to pull buffer from the wrong port. | alex.tello@sanofi.com | CNC | 2025-09-01 | 2025-09-01 | for a port to be called in the SOP and in the MES batch record. All to have what SUT to use |  | Deviation
6056 | 8NYA GEMBA - DOWNSTREAM P2, No instruction on WFI connection for MMC skid clean | alex.tello@sanofi.com | CNC | 2025-09-01 | 2025-09-01 | for a port to be called in the SOP and in the MES batch record. All to have what SUT to use
6055 | Witness step missed mixer speed | kevin.ornell@sanofi.com | 45-190 | 2025-08-27 | 2025-08-31 | Witness step would be performed by Operator
6054 | Wave Bag P/N 8NYA-No Product-SUTInspection Failure/ Wave Bag P/N 669424 | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2025-08-30 | 2025-08-30 | SUT pass criteria defined in SOP
6053 | Fz P1 Embedded particulate in SUM bag | elaine.jansen@sanofi.com | 1333 | 2025-08-31 | 2025-08-31 | If failed, Failed Bag Inspection Checklist needs to be completed for SUT
6052 | Particulate found during ATF10 inspection | joseph.gonzalez@sanofi.com | 8NYA room 1403 | 2025-08-31 | 2025-08-31 | New ATF with no particulate's
6051 | 8 NYA - Support Services - 705661 RF545F128 failed 2 post FIT attempts | william.lee@sanofi.com | 8NYA/1350 | 2025-08-31 | 2025-08-31 | Post FIT should pass.
6050 | 8NYA - GEMBA -Downstream Suite 1 - SNOW Ticket CHG008656 | danny.martinez@sanofi.com | 8NYA/ RM 1323 | 2025-08-31 | 2025-08-31 | Clarification of the snow ticket is needed
6049 | GEMBA: 8NYA Suite 1 Fabrazyme Zinc: pH discrepancy comment in MES record | William.Bradley@sanofi.com | 8NYA /RM1323 | 2025-08-14 | 2025-08-30 | Comments made in the record should be verified and correct.
6048 | 8NYA-Pompe-X2220 MFC not able to maintain %DO | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2025-08-29 | 2025-08-29 | It is expected that the Mass Flow Controller maintains the setpoint range for %DO
6047 | GEMBA 8NYA Suite 1 Wave X-2121 Prompt Adjustment | GARY.DODAKIAN3@SANOFI.COM | 8 nya suite 1 room 1333 | 2025-08-30 | 2025-08-30 | inoculation prompt should be present
6046 | SUB X2230 DO HIHI MFC Leak By | brian.o'rourke@sanofi.com | 8NYA Room1403 | 2025-08-29 | 2025-08-29 | MFC-007 should be closed to allow MFC-008 to control process air.
6045 | 8 NYA Gemba - Support Services - 805003 - Incorrect Delta V Formula does not include Liquid Addition | nicholas.grote@sanfofi.com | 08-1350 | 2025-08-29 | 2025-08-29 | The recipe should include an intermediate mix prompt after the dry addition, following that prompt the team would do the liquid addition and kick of the second timer before the batch moves to QS.
6044 | 8NYA-SS-GEMBA, Execution steps not documented on Action Notice 2025-I-098 | Kelley.Gibson@sanofi.com | 8 NYA-CNC area 1210, MALs, and PALs | 2025-08-25 | 2025-08-29 | Step 5, area clearance and date/time of signage postings should have been signed at time of execution.
6043 | 8NYA-Gemba-Downstream Suite 2 Cz freeze bags failed pre-use inspection for exterior embedded black particulates throughout | aidan.mula@sanofi.com | 8RM1703 | 2025-08-26 | 2025-08-26 | Bags should not contain particulates
6042 | 8 NYA - Support Services - Media Expiry Misalignment between MES/Athena and potential storage misalignment | nicholas.grote@sanofi.com | NA | 8/15/25 and 8/26/25 | 2025-08-27 | Expiry in MES and Athena Align.
6041 | 80 NY-8 NYA APU-Room 122-no product impact-EMS readout FB-5105-005 on -80 C freezer out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 80-122 | 2025-08-26 | 2025-08-28 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
6040 | X2230 Agitator did not turn on following media charge | ralston.najarro@sanofi.com | 8NYA1403 | 2025-08-28 | 2025-08-28 | Agitator should turn on when recipe command |  | No Deviation
6039 | SAP MES Component Discrepancy | Thomas.Loik@sanofi.com | 8RM 1350 | 2025-08-28 | 2025-07-07 | Values align
6038 | Received SCADA call at 12:56AM for 45NYA RO system alarm. Investigation found out LOW level at the main RO tank . Distribution pumps were still running. 45-RODI loop 1 penthouse tank was low as well. In order not to drain the main tank, The fill valve to the loop 1 was shut off for investigation . At this point the distribution pump also shut off since thank was empty within few munities. WE investigated all the labs for leaks and possible POUs left open and could not find the cause. At around 1:30am , We found the drain valve switch was left open instead of auto position . This turn out to be the cause of the low level alarm . Switch turn off and system is running normal. WO # 3386701 | ANDYLE@SANOFI.COM | 45 NYA RODI. WO # 3386701 | 2025-08-28 | 2025-08-28 | N/A
6037 | WHO discovered the poly bag was breached during the HU labeling process | michael.montiverdi@sanofi.com | 8NYA | 2025-08-26 | 2025-08-26 | SUT material does not impact the integrity of the outer poly bag
6036 | 8NYA - Support Services - Conductivity SOP unclear about units we need to use | maria.palmar@sanofi.com | 8NYA / 1350. | 2025-08-27 | 2025-08-27 | SOP to have a clear path
6035 | 76 NYA - HA - FBYFRM036 - Organism of concern detected on RODAC negative control plate for EM under dynamic conditions | larissa.fairweather@sanofi.com | 68 NYA / 4700 / QCM | 2025-08-19 | 2025-08-27 | The expectation is for negative control results to have 0 CFU.
6034 | Seprafilm - FBYSEP060 - 76NYA – AHU 5 Shutdown during Seprafilm Casting process | Yamini.patcha@sanofi.com | 76-238 (casting room) | 2025-08-27 | 2025-08-27 | Casting should be performed after 26 hours of resuspension
6033 | 8NYA - Support Services - Conductivity meter 2008025 put out of service | maria.palmar@sanofi.com | 8NYA / 1350. | 2025-08-27 | 2025-08-27 | Calibration testing to not fail.
6032 | FZ Suite 1 RF503F012 MES Auto step failure | jared.watson@Saonfi.com | 8RM1333 | 2025-08-26 | 2025-08-26 | the auto step is supposed to execute automatically
6031 | 51 NYA - TSH - Cell Culture - Alternate filter part number not able to be issued to Athena Process Order | MARIE.VALLIERE@SANOFI.COM | 51 NYA 140 | 2025-08-25 | 2025-08-25 | It is expected that the MIGO Athena transaction would be successful and parts used in the process would be correctly issued to BOM. | Yes | No Deviation
6030 | Downstream Suite 2 - X-3303 Bubble Trap Level Sensors out of tolerance | bo.caldwell@sanofi.com | 8 NYA downstream suite 2 | unknown | 2025-08-26 | LS-3304061 is at 32+/- 5 mm, LS-3304062 is at 200+/- 5 mm, LS-3304080 is at 349+/- 5 mm,
6029 | FZ Upstream suite 1 RF503F012/FGHCE012 Cloudy ATF | jared.watson@Sanofi.com | 8RM1333 | 2025-08-25 | 2025-08-26 | non rising pressure and clear ATF
6028 | 68 NYA - Thyrogen - Multiple Lots - QC unable to perform source data review on LAL Plate-Reader FB-1154. | larissa.fairweather@sanofi.com | 68 NYA / 4700 / QCM | 2025-08-20 | 2025-08-26 | It is expected that data is retrievable upon performing source data review.
6027 | CZ Freeze Operation, Two Bags Leaking from the Crimp | zachary.rickard@sanofi.com | 8NYA/1703 | 2025-08-26 | 2025-08-26 | Crimp will not leak.
6026 | 8NYA-US-SS- Daily checklist Area clearance not on SOP. | genesis.torres2@sanofi.com | 8nya CNC | 2025-08-24 | 2025-08-24 | update SOP to reflect changes.
6025 | GEMBA - 8 NYA - Suite One - Adapter Stopping Before Column Unpack Is Done | adam.keister@sanofi.com | 8 NYA, Room 1303 | 2025-08-26 | 2025-08-26 | To be able to finish Unpacking procedure.
6024 | GEMBA: 8NYA: Morning Sample Run | William.Bradley@sanofi.com | 8NYA 1410 | 2025-08-26 | 2025-08-26 | Samples delivered. Calling a GEMBA to understand how multiple buildings communicate with sample runs if not using the same logbook when buildings need assistance
6023 | GEMBA - 8 NYA - Suite One - Not All Product Specific Hoses Changed on Qual Skid Prior to CIP | adam.keister@sanofi.com | 8 NYA, Room 1303 | 2025-08-26 | 2025-08-26 | For all hoses to be changed out prior to CIP.
6022 | MABC-QU-SOP-0084585-Filter Preparation and Testing (v3.0) does not include information for 805015 Cystine / Tyrosine Solution. | julie.auguste@sanofi.com | Solution prep, 1290 | 2025-08-25 | 2025-08-25 | It is expected that the SOP used to prep and perform filter testing and wetting contains the appropriate information for every buffer and media that Solution prep formulates.
6021 | Upstream Pompe, RF1NAF011, MES record missing take out step for PN: 692307 in growth media setup | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2025-08-25 | 2025-08-25 | all SUT takeout spots to be available for SUT use
6020 | MES Misalignment with Control Strategy | ilene.robinson@sanofi.com | 8 NYA | unknown | 2025-08-25 | MES product loading values are aligned with the Control strategy
6019 | 8NYA - Support Services - Missed Weekly Walkthrough for week of 18AUG25 | christopher.arnold@sanofi.com | 8RM1290, 8RM1364 | 2025-08-23 | 2025-08-25 | That a scheduled walkthrough inspection will be held.
6018 | Damaged GZ4199-01 shipper for lot FBYSEP009 unlabeled part number 768282 | Alan.rousseau@sanofi.com | 74NYA/RM 1017 | 2025-08-25 | 2025-08-25 | Team will make a decision regarding damage to films.
6017 | SOP/MES Nano Equil Buffer Discrepancy, 8NYA, DS, P2 | david.loik@sanofi.com | 8nya | 2025-08-25 | 2025-08-25 | SOP/MES align
6016 | 8NYA - Support Services - Bag leaked during initial WFI fill | william.lee@sanofi.com | 8NYA/1350 | 2025-08-25 | 2025-08-25 | 3K SUM bag should not leak
6015 | 8NYA-SS-GEMBA, Execution steps not documented on Action Notice 2025-I-094. | Kelley.Gibson@sanofi.com | 8NYA-CNC1210 | 2025-08-19 | 2025-08-25 | per QU-FOR-0092319, Section 5-Work execution Completion section must be completed at time of execution (name, date and time).
6014 | 8NYA- Support Services Gemba-breakdown step not completed | william.lee@sanofi.com | 8NYA/1350 | 2025-08-25 | 2025-08-25 | Step for breakdown should have been completed and verified.
6013 | 8NYA - Support Services - Bag leaked during initial WFI fill | maria.palmar@sanofi.com | 8 New York Ave | 2025-08-24 | 2025-08-24 | The bag to not leak
6012 | 10L Wave Bag failed for particle found in Middle Innoc Line | ryan.devanna@sanofi.com | 8NYA / 1333 | 2025-08-24 | 2025-08-24 | SUT Gemba to determine path forward
6011 | Cleaning Verification Attachments is missing instructions for ASM | Pavel.dobrovolskiy@sanofi.com | 8nya 1303 | 2025-08-23 | 2025-08-23 | its expected that ASM CV sampling would be covered under the same SOP
6010 | WFI TOC analyzer lost communication with EMS from 0924 AM to 0942 AM on 18aug2025 | jesus.acosta2@sanofi.com | 45-240 | 2025-08-18 | 2025-08-18 | EMS should have received data all the time. No gap should have exists.
6009 | 8NYA-Gemba-Downstream Suite 2 Cz Phenyl Load filter FIT was performed under incorrect parameters | aidan.mula@sanofi.com | 8RM1350 | 2025-08-21 | 2025-08-22 | Parameters should not be changed
6008 | 8NYA-FZ-US-liquid on media tray. | genesis.torres2@sanofi.com | 8nya CNC | 2025-08-22 | 2025-08-22 | Drain media (?) and complete next install.
6007 | Mixed Scans of TSA labels with RODAC labels during BSC Operation | jonathan.lucas@sanofi.com | 51NYA / 152B | 2025-08-21 | 2025-08-21 | Expected that correct label IDs be scanned.
6006 | Vendor (Accugenix) Sent Data to Different Company's Account | kyle.lebherz@sanofi.com | QC Micro | 2025-08-20 | 2025-08-20 | Our Data not be sent to other business accounts.
6005 | MES is not recording a takeout a quantity. | jose.lugo2@sanofi.com | 8 NYA CNC | 2025-08-21 | 2025-08-21 | We will need to perform a takeover and the storage unit be consumed in the record.
6004 | DS Phase 2 pH Meter (2009953) failed standardization twice | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2025-08-20 | 2025-08-20 | pH sample will be between 6.95 and 7.05, per SOP-0084641 step 6.8.23
6003 | 8NYA - GEMBA - DS Suite 2 - Failed Pre-Use Shipping Bags | danny.martinez@sanofi.com | 8NYA RM1703 | 2025-08-20 | 2025-08-20 | Per QU-SOP-0080323 Step 6.16.15, If a particulate is observed to be embedded, hold a gemba to determine if the bag is acceptable for use.
6002 | Multiple Shipping bag failures | Juan.Bonilla@sanofi.com | 8NYA/1313 | 2025-08-18 | 2025-08-18 | Shipping bag should pass pre inspection
6001 | 76 NYA - FBYCMC122 - pH probe stopped working during the EDC reaction with HA-CMC | Yamini.patcha@sanofi.com | 76-229 | 2025-08-20 | 2025-08-20 | pH during the EDC reaction should be between 4.60 to 5.00
6000 | GEMBA - 8 NYA - Suite Two - Leak on Sample Port of Zinc Eluate After Using Gator Crimper | adam.keister@sanofi.com | 8 NYA Room 1503 | 2025-08-20 | 2025-08-20 | For crimp not to leak - path forward.
5999 | P1 ATF6 B not perfusing | elaine.jansen@sanofi.com | 1333 | 2025-08-20 | 2025-08-20 | Open air valves and restart atf | Yes | No Deviation
5998 | 8 NYA Support Services- RF1N8F006 Liquid weigh out wont convert units | William.Lee@sanofi.com | 8 NYA 1350 | 2025-08-20 | 2025-08-20 | step 6.4.5.2 in MABC-QU-SOP-0083660-Solution Preparation 8 NYA. For weigh out to be converted from grams to liters to satisfy the BOM .
5997 | 8NYA-Gemba-Downstream Suite 2- Pump asset has labeling for 2 different assets | aidan.mula@sanofi.com | 8RM1503 | 2025-08-19 | 2025-08-19 | Equipment should have the appropriate labeling
5996 | 8NYA sampling curriculum did not include MABC-QU-SOP-0020206-Manual Test Request System - 5.0 - English | adam.hildebrand@sanofi.com | 8NYA rm 1334 | 2025-08-19 | 2025-08-19 | It is expected that the sampling curriculum for 8NYA would include all requisite SOP guidance. | No
5995 | 8NYA - Support Service - RF1ASF023 PCS alarm occurred for exceeding timer prematurely | bailey.gleason@sanofi.com | 8NYA RM 1350 | 2025-08-10 | 2025-08-19 | Alarms not be generated for instances when we have not gone out of specification
5994 | 8NYA-Gemba-Suite 2 Downstream MMC Eluate SUM is not mixing due to agitator error | aidan.mula@sanofi.com | 8RM/1503 | 2025-08-19 | 2025-08-19 | Agitator should function properly after setup
5993 | Missed Weekly Walkthrough Week of 10Aug25 | kara.mack@sanofi.com | 8NYA Support Services CNC, 11 Forbes Label & Pack, 8NYA Suite 1 Downstream | 2025-08-17 | 2025-08-18 | FDA 483 Response, CAPA-040603, SOP-0075222 | Yes | Deviation
5992 | 8NYA - unclear about impact of WFI Flush being skipped on loop 3B and 4 for 3 days | bailey.gleason@sanofi.com | 8NYA | 2025-08-15 | 2025-08-18 | per MABC-QU-SOP-0001515 section 6.12.1 automated flushing of the WFI drops should occur daily during the daily loop sanitization.
5991 | 8NYA-Gemba-Suite 2 Downstream Cz MMC Pooling SUM bags failed pre-use inspection | aidan.mula@sanofi.com | 8RM1503 | 2025-08-18 | 2025-08-18 | Bags should not contain particulates
5990 | Badge Access PM Not Completed as Scheduled | brenda.tremblay@sanofi.com | TBD | 2025-08-18 | 2025-08-18 | Badge access review occurs within the specified period of time
5989 | step. 6.34.17 Missing from MES MMC record. No pi pull. | genesis.torres2@sanofi.com | 8nya CNC | 2025-08-18 | 2025-08-18 | update SOP to reflect changes.
5988 | 8NYA - Support Service - RF573F122 incorrect status change of equipment | bailey.gleason@sanofi.com | 8NYA - RM1350 | 2025-08-05 | 2025-08-18 | Hold gemba to see path forward
5987 | Large amount of DP alarms related to AHU possibly going down | jesse.demers@sanofi.com | 51 NYA multiple rooms | 2025-08-18 | 2025-08-18 | Control of effected areas would be maintained.
5986 | 8NYA All Areas - HBOD Leak Experienced in Waste Airlock Room 1115 | wilfredo.lopez@sanofi.com | 8RM1115 | 2025-08-18 | 2025-08-18 | It is expected that no leaks are observed. A Gemba is being requested to discuss further containment, disposal, and cleaning operations.
5985 | OSMO 2009180 DS P1 Placed Out of Service | elias.mcdonnell@sanofi.com | 8NYA, 1323 | 2025-08-18 | 2025-08-18 | OSMO daily measurement will come in at 290mOsm/kg +/-2 of ea. other
5984 | FZ Upstream Suite 1 RF503F012/FGHCE012 pCO2 Probe 2 unable to recal | jared.watson@Sanofi.com | 8RM1333 | 2025-08-17 | 2025-08-17 | in process cal performed
5983 | 8NYA Downstream Suite 2 - CM Standalone CIP MBR BOM Discrepancy | wilfredo.lopez@sanofi.com | 8RF1503 | 2025-08-16 | 2025-08-16 | The MBR BOM needs to be updated to reflect the correct SUT required for the operation. | Yes | Deviation
5982 | 8NYA Downstream Suite 2 - Phenyl Standalone CIP Recipe Pulling Buffer from the Incorrect Port. Automation Intervention Required | wilfredo.lopez@sanofi.com | 8RF1503 | 2025-08-16 | 2025-08-16 | It is expected that the buffers are connected to the correct port of the buffer manifold in the CNC so that the recipe executes as expected. The batch record for Phenyl skid cleaning needs to be updated to instruct the proper setup and connection of Buffer 1 PN 705664 onto port 9 (BV-3322016)
5981 | 8NYA GEMBA - Support services - Autoclave aborted probe recipe | jack.laplante@sanofi.com | 8NYA 1373 | 2025-08-16 | 2025-08-16 | Path forward? keep probes
5980 | x2220 Harvest ENDO/BIO Not taken. | John.Lytle@sanofi.com | 8 NYA RM 1403 Upstream suite 2 | 2025-08-15 | 2025-08-15 | GEMBA to determine QE event .
5979 | Column inlet conductivity meter AI-3001053 not reading correct value | Robert.pino@sanofi.com | 8NYA, RM1323, Suite 1 | 2025-08-16 | 2025-08-07 | It is expected that the column inlet and outlet conductivity meters would be reading close to identically during equilibration, and not be off by a magnitude of 1000.
5978 | 8NYA GEMBA Suite 1 Upstream NRSR 25SR8131 Delivered to QC Expired with No Manual Test Request | Gary.dodakian3@sanofi.com | 8 nya suite 1 upstream room 1333 | 2025-08-15 | 2025-08-15 | Sample is delivered within expiry with a manual test request
5977 | Harvest SUMs & MCC Eluate Pool SUMs do not have Batch IDs during production. | Hiren.amin@sanofi.com | 8NYA DeltaV Production | 2025-08-14 | 2025-08-14 | Harvest SUMs & Eluate Pool SUMs and their devices should have Batch IDs | Yes | Deviation
5976 | 8NYA - GEMBA Downstream Suite 1 - Discrepancy with Fz Zinc SOP During Eluate Filtration Setup | danny.martinez@sanofi.com | 8NYA RM1323 | 2025-08-15 | 2025-08-15 | Per QU-SOP-0083859 Step 6.29.4, Expected results would be to remove the step of connecting pressure membrane to pressure transmitter from SOP.
5975 | 8NYA- Support Services Gemba-standardization of Osmo unlcear | william.lee@sanofi.com | 8NYA/1503 | 2025-08-15 | 2025-08-15 | Osmo standardization should be clear for both SOP and MES.
5974 | 76NYA-No Lots-Insect found in HA Common Hallway Room 363 | Phillip.Mank@Sanofi.com | 76 NYA Room 363 | 2025-08-15 | 2025-08-15 | No insects within the Grade D area
5973 | 51NYA CC Gemba: freezer in 140 out of range, dispase impacted | adam.rowe@sanofi.com | 51NYA, 140 | 2025-08-15 | 2025-08-15 | The freezer should maintain setpoint of -20C
5972 | Aquasyn Valve Failure During Sanitization of UFDF Skid | brian.madigan@sanofi.com | 8NYA, Suite 1, Room: 1323 | 2025-08-14 | 2025-08-14 | It is expected that closed valves do not allow solution to pass through them during processing.
5971 | 8NYA-Gemba-Downstream Suite 2 Cz Shipping Bag failed pre-inspection | aidan.mula@sanofi.com | 8NYA/RM1703 | 2025-08-14 | 2025-08-14 | All parts should be integral from vendor
5970 | datalogger alarm during vial transfer from 51NYA to 45NYA | nicole.lufkin@sanofi.com | 45 NYA room 101 | 2025-08-14 | 2025-08-14 | It is expected that the temp remains at or below 150C.
5969 | 8NYA - Support Services - Transfer filter sterile connection failed inspection | maria.palmar@sanofi.com | 8 New York Ave, room 1350. | 2025-08-14 | 2025-08-14 | The aseptic connection to be in perfect conditions | Yes | Deviation
5968 | Material GZ1476-00 not properly placed into QI status per CC-015157 | timothy.pyne@sanofi.com | 51 NYA | 2025-04-30 | 2025-08-14 | It is expected that the change action would have been successfully completed
5967 | During the receiving process on 13Aug25 at 8NYA it was identified the case quantity from the vendor for PN 757793 batch 25GN01340 was short by one piece ( case of 20ea/ 19 received) | michael.montiverdi@sanofi.com | 8NYA | 2025-08-13 | 2025-08-13 | Correct case quantity sent from vendor
5966 | X2220 high pO2 values during daily sampling | michael.yashin@sanofi.com | 8NYA, Room 1403 | 2025-08-14 | 2025-08-14 | pO2 values should be below 99.5 mmHg
5965 | Buffer 705668 Endotoxin Sample Tested at Wrong Dilution | Alyssa.Barrett@sanofi.com | 68 NYA/Room 4700/QC Micro | 2025-08-04 | 2025-08-14 | It is expected that the sample be tested at the minimum validated dilution of 1:40, as stated in QU-SOP-0010784.
5964 | 8NYA - Support Services - Incorrect Part Number request on Pompe Seed Media record | maria.palmar@sanofi.com | 8 New York Ave, room 1350. | 2025-08-14 | 2025-08-14 | The record to call the correct P/N
5963 | 8NYA-Gemba-Downstream-Cz Post Viral SUM Bag failed pre-use inspection | aidan.mula@sanofi.com | 8RM1703 | 2025-08-13 | 2025-08-13 | Bag should not contain particulates
5962 | Material 6239-01 envelopes lot 25MN00325 not varnish coated | alan.rousseau@sanofi.com | Room 1017 74 NYA | 2025-06-30 | 2025-08-14 | It was expected envelopes would have varnish coating.
5961 | Material 6239-01 envelopes lot 25MN00325 not varnish coated | alan.rousseau@sanofi.com | Room 1017 74 NYA | 2025-06-30 | 2025-08-14 | It was expected envelopes would have varnish coating.
5960 | 8NYA - MFG ALL - WFI used prior to AN 2025-I-076 being released | Bailey.gleason@sanofi.com | 8NYA | 2025-07-03 | 2025-08-14 | AN be closed out prior to WFI useage
5959 | Material 6239-01 envelopes lot 25MN00325 not varnish coated | alan.rousseau@sanofi.com | Room 1017 74 NYA | 2025-06-30 | 2025-08-14 | It was expected envelopes would have varnish coating.
5958 | Possible Incorrect Status of CC-0052535 Prior to Thaw of Vials for WCB 12TP040 Qualification | mark.lindblad@sanofi.com | 51 NYA | 2025-02-12 | 2025-08-12 | Per Change Control Management it is expected that the CC be in the correct state prior to execution. A blocking action to generate an execution protocol / plan is still pending.
5957 | 8NYA pH Meter gap with the configurations | angela.keister@sanofi.com | 8NYA Suite 2 | 2025-01-01 | 2025-08-11 | pH meter needs to be configured to PI
5956 | 8NYA - Support Services - Multiple issues with verification comments in FAC cleaning EURs | christopher.arnold@sanofi.com | 8NYA | 07/28/2025 (date of original Gemba Request #5861). | 2025-07-28 | That manual entries will be made per guiding SOPs with the correct verification signatures and format.
5955 | 8NYA - Support service - RF1NMF003 SOP not aligning for cystine tyrosine transfer | bailey.gleason@sanofi.com | 8NYA - rm1350 | 2025-08-13 | 2025-08-13 | SOP align with operation
5954 | Duplicate BioLIMs lot | phoebe.karkos@sanofi.com | Upstream Suite 1 | 2025-08-06 | 2025-08-08 | There should only be one lot for each batch in biolims
5953 | Duplicate BioLIMs lot | phoebe.karkos@sanofi.com | Upstream Suite 1 | 2025-08-06 | 2025-08-08 | There should only be one lot for each batch in biolims
5952 | Duplicate BioLIMs lot | phoebe.karkos@sanofi.com | Upstream Suite 1 | 2025-08-06 | 2025-08-08 | There should only be one lot for each batch in biolims
5951 | 8NYA -Support Service - RF573F129 - 705691 stock created for post filtration media after prefiltration 10 hour timer alarmed | Bailey.gleason@sanofi.com | 8NYA RM 1350 | 2025-08-13 | 2025-08-13 | hold gemba to see path forward
5950 | MABC-QU-SOP-0083660-Solution Preparation 8 NYA does not align with MO RF573F129 | joshua.crowley@sanofi.com | 8NYA, RM1350 | 2025-08-13 | 2025-08-13 | n/a
5949 | 51NYA CC Gemba: Refeed timer exceeded on E3 | adam.rowe@sanofi.com | 51NYA, 140 | 2025-08-12 | 2025-08-12 | Work should be within the refeed timer.
5948 | Vante Tube Sealer X-3322004 Failed Initial IQOQ in 2020 and No Follow up Testing was Done | paul.mclaughlinjr@sanofi.com | 8NYA | 2020-05-22 | 2025-06-24 | All assets should have completed validation studies prior to GMP use
5947 | 8NYA Support Services GEMBA - pre use inspection failure on air pressure line. | JACK.LAPLANTE@SANOFI.COM | 8NYA RM1350 | 2025-08-11 | 2025-08-11 | SUT failure
5946 | GEMBA - 8 NYA DSP Suite Two - Ambient Timer Inaccurate in MES Thaw Record | adam.keister@sanofi.com | 8 NYA Room 1703 | 2025-08-12 | 2025-08-12 | For MES record to correctly calculate ambient exposure time of thawed freeze bags. | Yes | Deviation
5945 | 8NYA, Downstream , Suite 1, Butyl Load Tote bag failed SUT inspection | Robert.pino@sanofi.com | 8NYA, RM1323, Suite 1 | 2025-08-11 | 2025-08-11 | It is expected that the butyl load tote bag would not fail inspection due to particulates | Yes | No Deviation
5944 | Potential TOR limit not met/ Discussion of TOR cummulative or not | vincent.palumbo@sanofi.com | 8 NYA | 2025-08-11 | 2025-08-11 | It is expected that material remain within its TOR
5943 | Incorrect filter test program used to test filter part number 649134 during Nanofiltration/Formulation Operation | Alyson.Belleville@sanofi.com | 51, 87 | 2022-05-23 | 2025-08-10 | It is expected that the correct test program would be used to integrity test a filter.
5942 | 8 NYA Support Services / Cystine Tyrosine (P/N: 805015): Executed under incorrect record version | nicholas.candela@sanofi.com | 8 NYA / Solution Prep / Rm. 1350 | 2025-08-09 | 2025-08-09 | Per MABC-QU-SOP-0041372 / Master Batch Record Change Request (MBRCR) Process at Framingham Biologics / Step 6.5.8 notification should be made to signatories to initiate step 6.5.9 to abort previous version of record. | Yes | Deviation
5941 | No Product, No Lots - 68 NYA QC Micro - SDA Lot 25GN00847 found to be unsealed along the packaging seam | amanda.lavalley@sanofi.com | 68 | 2025-08-06 | 2025-08-11 | It is expected that the packaging is fully sealed and intact
5940 | Antifoam Line on X-2210 is clogged. | todd.niziak@sanofi.com | 1333 | 2025-08-11 | 2025-08-11 | Antifoam doses
5939 | Time recorded, instead of agitator speed during sampling TSH media | nicholas.shoenfelt@sanofi.com | 51-124 | 2025-08-02 | 2025-08-11 | It is expected that the agitator speed will be recorded at the start of pH adjust and will be within range of the batch record parameters.
5938 | 4 Conductivity CZ CV samples 2040536, 2040541, 2040545, 2040549 sampled in incorrect container | Stefanie.Erickson@sanofi.com | QC Micro Conductivity Bay | 2025-08-09 | 2025-08-11 | for conductivity samples to be taken in the correct container: part number: 2002-0004.
5937 | LAL samples 2002802 and 1941925 had invalid results and was not put back on test within expiry | Stefanie.Erickson@sanofi.com | QC Micro Nexus bench | 2025-08-03 | 2025-08-10 | for each sample to be put back on test before 10Aug2025
5936 | X2250 Air EM missing setpoint | ralston.najarro@sanofi.com | 8NYA1403 | 2025-08-10 | 2025-08-10 | Recipe should have all setpoints set
5935 | 8NYA-Gemba-Downstream Suite 1-MES references SOP step that does not exist | aidan.mula@sanofi.com | 8RM1313 | 2025-08-02 | 2025-08-10 | MES and SOP should align
5934 | Persistent Antifoam Issue on X2210 | MARCELO.DOSSANTOS2@SANOFI.COM | 1333 | 2025-08-09 | 2025-08-09 | Foam should be added to bioreactor
5933 | Antifoam building up on 2210/Antifoam Build up on X2210(Upstream) | MARCELO.DOSSANTOS2@SANOFI.COM | 1333 | 2025-08-09 | 2025-08-09 | Foam should be added to bioreactor
5932 | 8NYA - Support Service - QU-SOP-0084585 does not have 805015 solutions listed for filter integrity tester | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-08-09 | 2025-08-09 | 805015 should be listed on table 6
5931 | Credo cooler samples went over the 750mL total | dennis.bazinet@sanofi.com | 8 NYA / 1410 sample pick up room | 2025-08-09 | 2025-08-09 | Total volume needs to be below 750mL
5930 | SUB X2220 in hold for High High alarm for weight during ATF changeout. | adam.hildebrand@sanofi.com | 8NYA, RM 1403 | 2025-08-09 | 2025-08-09 | It is expected that the ATF changeout would not cause a weight excursion triggering the high high alarm limit for weight. | Yes | Deviation
5929 | Handling unit discrepancy | craig.valery@sanofi.com | 51 NYA | 2025-06-15 | 2025-08-09 | It is expected that accurate inventory counts are consistent with actual process order consumption
5928 | 8NYA - Support Service - RF567F011 - e stop alarms | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-07-23 | 2025-08-08 | hold gemba to see path forward | Yes | No Deviation
5927 | 8NYA - Support Service - RF616F018 - 3000L SUM bag leak | bailey.gleason@sanofi.com | 8NYA - RM 1350 | 2025-08-08 | 2025-08-08 | Hold Gemba to see path forward.
5926 | 8 NYA GEMBA Downstream Fabrazyme - shipping bag pre-inspection failure | lukus.palzkill@sanofi.com | 8NYA downstream suite 1 postviral room 1313 | 2025-08-08 | 2025-08-08 | Bag will be discarded new bag will be used
5925 | Downstream Suite 1 Osmometer status went out of service. | Juan.Bonilla@sanofi.com | 8NYA/1324 | 2025-08-08 | 2025-08-08 | The Osmo record should have blocked operators from re-standardizing equipment.
5924 | 8NYA Downstream Fz Nano formulation sample out of spec. | Christian.Venegas@Sanofi.com | 8NYA Downstream suite 1 room 1323 | 2025-08-08 | 2025-08-08 | To get a Phath forward
5923 | ATF 10 Visual Inspection Failure for particulates | Ralston.Najarro@sanofi.com | 8NYA Room 1403 | 2025-08-08 | 2025-08-08 | ATF To pass inspection
5922 | HA - 76 NYA - No Lots - Room 352 - Gruenberg Oven (F-4507) Temperature went high to 267°C During Depyrogenation Cycle | Yamini.patcha@sanofi.com | 76-352 | 2025-08-05 | 2025-08-06 | temperature to be maintained at 254°C for 3 hours
5921 | No Product-8NYA-Refrigerator Asset#2009951 temperature excursion | john.lytle@sanofi.com | 8NYA Rm 1403 | 2025-07-31 | 2025-08-02 | Its is expected that refrigerator maintain temperature ranges
5920 | 51 NYA - TSH - Cell Culture - Steps in Refeed SOP no longer performed due to MBR revision | MARIE.VALLIERE@SANOFI.COM | 51 NYA, RM 140 | 2025-08-07 | 2025-08-07 | MBR and SOP should align, and Automation recipe for refeed should be updated to remove post-refeed timer. | Yes | Deviation
5919 | Black particle found in 692174 batch # 25GN00358 3D TOTE BAG HU #1001453769 | michael.montiverdi@sanofi.com | 8NYA | 2025-08-06 | 2025-08-06 | Materials to be free of particulates.
5918 | 8NYA- Improper gowning observed at 8NYA CNC space | Holly.rocco@sanofi.com | 8NYA room 1331, 1201 south corridor, 1240 gowning room | 2025-08-07 | 2025-08-07 | Per MABC-QU-OPE-0510073 high-top boot should be donned when entering CNC with non-plant shoes.
5917 | 2 HEPAs failed certification | glen.ciano@sanofi.com | 76-3103,3101 | 2025-08-07 | 2025-08-07 | Per QU-SOP-0074732 unit should have produced passing results
5916 | Fly Found On Cart In Room 51-124 | jeff.hill@sanofi.com | 51-124 Media Preparation Area | 2025-08-07 | 2025-08-07 | Pest will not be found in a Manufacturing Suite
5915 | 76NYA – No lots – Walk through procedure not documented correctly in Work Order 3323965 | Phillip.Mank@Sanofi.com | 76 NYA | 2025-07-18 | 2025-07-18 | MFG or QA names for who performed the walkthrough was expected to be listed on the work order.
5914 | 51NYA, CC: 2 errors found in MBR during review | adam.rowe@sanofi.com | 51/140 | 2025-08-03 | 2025-08-07 | The allocate step should be signed. And the cell count should have been 4 quadrants
5913 | Suite 1 Upstream FZ RF503F008 SUB antifoam not subsided after second bolus | jared.watson@Sanofi.com | 8RM1333 | 2025-08-05 | 2025-08-06 | antifoam was expected to enter the SUB
5912 | Manifold part number 813903 is not available for Freeze setup | zachary.rickard@sanofi.com | 8NYA/1703 | 2025-08-07 | 2025-08-07 | All SUT required for processing will be available
5911 | Column inlet conductivity meter AI-3001053 not communicating reading to DeltaV | Robert.pino@sanofi.com | 8NYA, RM1323, Suite 1 | 2025-08-07 | 2025-08-07 | Meters are supposed to reflect their appropriate readings on the DeltaV terminal.
5910 | Hold time in MES exceeds hold allowed per hold study | nathan.kallen@sanofi.com | 68 NYA | 2025-08-06 | 2025-08-06 | Hold time in MES should not exceed QU-OPE-1945641 and should align with MABC-QU-OPE-0519399 (Cz Control Strategy) |  | No Deviation
5909 | Quarterly review of Filemaker Pro user access forms not completed since Q2 2024 | kimberly.riley@sanofi.com | N/A | 2025-08-05 | 2025-08-05 | It is expected that the user access assessment is performed once per quarter as defined in QU-OPE-00121666, step 6.4.
5908 | Incorrect Approval Workflow for MES PVL Approvals | mark.smith@sanofi.com | 8NYA | 2024-12-04 | 2025-08-06 | Manufacturing approvals performed for circulating PVLs
5907 | Gemba 8NYA Suite 1 Upstream eluate SUM PBM-2547001 Pressurizing when in overlay 706930 FGHCE00806 wo 3377154 | gary.dodakian3@sanofi.com | 8nya room 1333 suite 1 upsrtream | 2025-08-08 | 2025-08-05 | SUM does not have avctive air over the eluate
5906 | 8NYA - Fz - 728285 FGHDT013- RF1CJF013 Nano was released in MES. The release should applied in Athena SAP | Daniel.bati@sanofi.com | 8NYA | 2025-08-02 | 2025-08-05 | Athena should have the Tcode ACAN to release
5905 | Fz UFDF- Load Grams of Protein out of range | nathaniel.barnes@sanofi.com | 8NYA 1323 | 2025-08-06 | 2025-08-06 | Determine path forward for UFDF processing
5904 | 8NYA gemba, transfer-cap operations require critical surface, no location specified in diagram | Jacob.Roman@sanofi.com | 8RM1343 | 2025-08-05 | 2025-08-05 | EM samples to all have a pre-determined sampling location
5903 | CZ - RF639F018 - 8NYA - DO Probe 1 Re-enable discussion | michael.knapp@sanofi.com | 8NYA 1403 | 2025-08-05 | 2025-08-05 | Controlling probe is most representative DO probe within system |  | No Deviation
5902 | Bioburden sample 1934019 missing test volume | kathryn.lentoni@sanofi.com | 68 New York Ave, QC Micro | 2025-07-30 | 2025-08-05 | Per QU-SOP-0073109 the testing volume should be 10mL
5901 | Bioburden sample 1934019 missing test volume | heidy.ruano@sanofi.com | 68 New York Ave, QC Micro | 2025-07-30 | 2025-08-05 | Per QU-SOP-0073109 the testing volume should be 10mL
5900 | Bioburden sample 1934019 missing test volume | heidy.ruano@sanofi.com | 68 New York Ave, QC Micro | 2025-07-30 | 2025-08-05 | Per QU-SOP-0073109 the testing volume should be 10mL
5899 | Bioburden sample 1934019 missing test volume | heidy.ruano@sanofi.com | 68 New York Ave, QC Micro | 2025-07-30 | 2025-08-05 | Per QU-SOP-0073109 the testing volume should be 10mL
5898 | External DeltaV System Resource Missing Admin Training | Seyda.Dumrul@sanofi.com | 8NYA, 51NYA | N/A | 2025-08-04 | DeltaV system administrators are required to complete the admin training
5897 | X2220 pH probe 1 drift | sean.gormley@sanofi.com | 8RM1403 | 2025-08-04 | 2025-08-04 | Probes should not become fouled during course of run.
5896 | 788610 and 907799 SPC Storage Conditions do not Match Athena | joshua.brooks@sanofi.com | N/A | 2025-08-04 | 2025-08-04 | SPC and Athena and MES will all align and material will articulate to MES when transferred into MES
5895 | 51NYA-Thyrogen-high pO2 QCA on E3 bioreactor | adriana.alban@sanofi.com | 51-140 | 2025-08-04 | 2025-08-04 | It is expected that pO2 process stays in range
5894 | 8NYA - GEMBA Downstream, Pulling tabs ripped off during transfer setup | alex.tello@sanofi.com | 1503 | 2025-08-04 | 2025-08-04 | for the entire tab to pull out. |  | No Deviation
5893 | Fz Zinc Eluate SUM bag failed inspection | elias.mcdonnell@sanofi.com | 8NYA, 1323 | 2025-08-03 | 2025-08-03 | Bag will have no debris, markings nor blemish within the bag itself.
5892 | 8NYA-Fz-UF/DF CIP recipe RF1CBF001 went on hold due to EM-X3071 temperature failure | Edgar.vale@sanofi.com | 8NYA, room 1323 | 2025-08-03 | 2025-08-03 | It is expected that the recipe would complete without any holds
5891 | 8NYA-Support Services-Failed inspection for SUT | william.lee@sanofi.com | 8NYA/1373 | 2025-08-01 | 2025-08-01 | Nano filter should have had correct oetiker clamp placement.
5890 | 8NYA- Support Services Gemba-10 hour timer on 705665 expired | william.lee@sanofi.com | 8NYA/1350 | 2025-08-02 | 2025-08-02 | Pump should have worked. Batch should have been completed.
5889 | High temp in cold room 147 @ 51NYA | adam.rowe@sanofi.com | 51NYA, 147 | 2025-08-02 | 2025-08-02 | Cold rooms should control to temp setpoint of 8C.
5888 | Butyl hold/load tote fail visual inspection | KELVIN.GLEAN@SANOFI.COM | 8NYA 1323 | 2025-08-01 | 2025-08-01 | Bags to free of any black spots | Yes | No Deviation
5887 | 8NYA- Support Services Gemba-tote transferred incorrectly for batch 705677 | william.lee@sanofi.com | 8NYA/CNC | 2025-08-01 | 2025-08-01 | Batch should have been transferred through the inlet line and recirculation line should be intact.
5886 | Support Services-8NYA-NovaFlex Coupling no established PM | Suman.Dhal@sanofi.com | 8NyA | 2025-07-28 | 2025-07-31 | PM should be assigned during onboarding of equipment
5885 | FZ Upstream suite 1 Eluate SUM bag seated incorrectly | Jared.Watson@Sanofi.com | 8RM1333 | 2025-08-01 | 2025-08-01 | Lip to stay seated | Yes | Deviation
5884 | Overdue Quarterly PM | Thomas.Loik@sanofi.com | 8RM 1350 | 2025-08-01 | 2025-08-01 | PM is completed
5883 | Increase pO2 value consequently 3 days for SUB 2240 / RF639F016 | bhargav.patel@sanofi.com | 8 NYA / 1403 | 2025-08-01 | 2025-08-01 | Daily BGA samples results should be below 99.1 mm/HG |  | Deviation
5882 | 8NYA-Nexviazyme Support service-FGHMD003X Bioburden and Endotoxin sample Stored at incorrect Temp | elizabeth.endsley@sanofi.com | Sample Receipt Room 4720 at 68 NYA | 2025-07-31 | 2025-08-01 | Samples to be stored in correct condition with the manual test request following the procedure and correctly identifing the storage requirement.
5881 | Support Services - Buffer 805007 Prep Automated Addition Occurring Outside Qualified Range | stephanie.aceto@sanofi.com | 8NYA, 1350 | unknown | 2025-07-31 | Automated transfer setpoints are performed within the qualified range of the associated equipment.
5880 | Power outage at 45 NYA effecting MAHU and possibly other assets | jesse.demers@sanofi.com | 45 NYA | 2025-07-31 | 2025-07-31 | No lose of power and all assets running as they are designed to. | Yes | Deviation
5879 | 8NYA- Support Services Gemba - Conductivity test result out of range | maria.palmar@sanofi.com | 8NYA / 1350. | 2025-07-31 | 2025-07-31 | Conductivity testing to not fail.
5878 | 8NYA GEMBA - CZ Zinc Load SUM bag and Eluate SUM failed pre-use inspection | alex.tello@sanofi.com | 1503 | 2025-07-31 | 2025-07-31 | Bag to pass inspection | No | Deviation
5877 | 8NYA APU-Room 1403-Flow transmitter FB-4222-012 and flow indicator FB-4222-013 found out of tolerance during a scheduled calibration. Product impact unknown | robert.bauch@sanofi.com | 8-1403 | 2025-07-30 | 2025-07-31 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5876 | Viability fell below 80% minimum on bioreactor X2210 | brendan.dickhaut@sanofi.com | 8NYA/Room 1333 | 2025-07-31 | 2025-07-31 | Results are above the 80% minimum. | Yes | No Deviation
5875 | LAL Sample 1966747 MEDIA_POU760741_ENDOTOXIN cracked sample tube | samuel.mbugua@sanofi.com | 68 NYA 4700 QCM | 2025-07-30 | 2025-07-30 | Per QU-SOP-0070558 Step 6.2.3, confirm sample containers are not cracked, leaking. Expectation is that sample tubes are not cracked.
5874 | LAL Sample 1966747 MEDIA_POU760741_ENDOTOXIN cracked sample tube | samuel.mbugua@sanofi.com | 68 NYA 4700 QCM | 2025-07-30 | 2025-07-30 | Per QU-SOP-0070558 Step 6.2.3, confirm sample containers are not cracked, leaking. Expectation is that sample tubes are not cracked.
5873 | Adjust cycle number on X-2543 from 130 to 126 | elon.elzam@sanofi.com | Phase 1 Upstream room # 1333 | 2025-07-30 | 2025-07-06 | 8 eluates sums would fill on 126 cycles and the end of run prompt would appear on the MCC
5872 | Fly / bug found in the light fixture of suite 1323 | pavel.dobrovolskiy@sanofi.com | 8nya- 1323 | 2025-07-30 | 2025-07-30 | its expected that there wont be any flies or bugs in side the light fixtures in the clean room
5871 | 8NYA-downstream-CZ zinc-multiple 200L tote bags failed inspection | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 room 1503 | 2025-07-30 | 2025-07-30 | Bag will be discarded new bag will be used
5870 | GEMBA 8NYA Upstream FZ Small Scale 10L Wave not Meeting Day 5 Criteria for Scale Up to 20L PN: 706929 Lot: FGHSD008 MO: RF500F008 | gary.dodakian3@sanofi.com | 8 NYA Upstream suite 1 room 1333 | 2025-07-30 | 2025-07-30 | the 10L FZ wave should meet scale up criteria to 20L wave by day 5
5869 | WO 3375287 investigate WFI high high pressure on wfi manufacture loop at 45 NYA. | jesus.acosta2@sanofi.com | 45-181 | 2025-07-29 | 2025-07-29 | WFI loop loop pressure should stays at 30 psi. High High pressure alarm worked as expected triggering an alarm and shutting down loop for safety. FBL-SOP-000767
5868 | GEMBA 8NYA Upstream FZ Small Scale 10L Wave not Meeting Day 5 Criteria for Scale Up to 20L PN: 706929 Lot: FGHSD008 MO: RF500F008 | gary.dodakian3@sanofi.com | 8 NYA Upstream suite 1 room 1333 | 2025-07-30 | 2025-07-30 | the 10L FZ wave should meet scale up criteria to 20L wave by day 5
5867 | Crushed box received of material 692171 from NBO delivery, box contains 4ea. | michael.montiverdi@sanofi.com | 8NYA | 2025-07-28 | 2025-07-28 | Cases of materials arriving from NBO integral.
5866 | material GZ7091-03 (disposable BOOT COVER, HIGH-TOP,X-LARGE) that was delivered to FRM site without NBO identification labels. | michael.montiverdi@sanofi.com | 80NYA | 2025-07-28 | 2025-07-28 | materials arriving from NBO should have Athena identification labels affixed
5865 | Gemba for a particulate found in 200L tote bag (P2, train 1 CZ) | joseph.gonzalez@sanofi.com | 8NYA, room 1403 | 2025-07-29 | 2025-07-29 | Obtain new tote bag with no particulates
5864 | Support Service - Missed Quarterly Cleanings Identified | Shawnique.Sudler@sanofi.com | 8NYA Phase 1: 1323, 1303, 1312, Phase 2: 1503, 1211, 1703, 1403, Support Service: 1200, 1212, 1210 | 2025-07-29 | 2025-07-29 | These missed cleanings represent a failure to adhere to the established cleaning and disinfection program, potentially impacting the facility's contamination control strategy and overall compliance with GMP requirements. |  | Deviation
5863 | Cerezyme, FGHBP023, 8 NYA, Downstream, Room 1503 Peristaltic Pump Setup and Operation Performance Without Training | eric.horten@sanofi.com | 8NYA Room 1503 | 2025-06-05 | 2025-07-29 | It is expected that operators complete OJT training prior to performance per GQCMS-QU-GOP-0000565-Personnel Qualification within GxP Functions (v7.0) | Yes | Deviation
5862 | Downstream Suite 1 Fabrazyme: Butyl : Pre-use bag failure (P/N: 692299) | William.Bradley@sanofi.com | 8NYA /RM1323 | 2025-07-28 | 2025-07-28 | It is expected that bags are in good integrity with no particulates | Yes | No Deviation
5861 | Manufacturing Cleaning Records Missing Verification | JOEL.HICKEY@SANOFI.COM | 8 NYA | 2025-07-28 | 2025-07-28 | All steps manually recorded should have verification.
5860 | Cleaning Records Missing Verification | JOEL.HICKEY@SANOFI.COM | 8 NYA | 2025-07-28 | 2025-07-28 | All steps manually recorded should have verification.
5859 | 45NYA Support APU-Room 1136- Balance FB-1243 found out of tolerance during an unscheduled calibration. | robert.bauch@sanofi.com | 45-1136 | 2025-07-25 | 2025-07-29 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5858 | Missing Supplier Audit Risk Assessment | Nashmin.Hojatti@sanofi.com | 68nya | 2025-07-28 | 2025-07-28 | it is expected that the risk assessment is completed prior to the year of the audit schedule |  | Deviation
5857 | Multi shipping bag failures | Juan.Bonilla@sanofi.com | 8NYA1313 | 2025-07-24 | 2025-07-24 | Shipping bags are expected to pass visional inspection | Yes | No Deviation
5856 | Delta V and downstream chilled totes do not match in Delta V | William.Rojcewicz@sanofi.com | 8 NYA Phase 2 Zinc / ANX line | 2025-07-28 | 2025-07-28 | All valve numbers and pumps should match to Delta V screen.
5855 | 8NYA-SS-GEMBA, Glycol leak from tote jacket. | Kelley.Gibson@sanofi.com | 8NYA-CNC1210 | 2025-07-28 | 2025-07-28 | No leaks
5854 | 76NYA Environmental Monitoring HA FBYHAH005 missing incubation entry | cleiton.lopes@sanofi.com | 76NYA/68NYA 4700 | 2025-07-23 | 2025-07-28 | EM incubation done in conjunction with receipt
5853 | 8NYA - Support service - 3 X 692174 failed pre-use inspection for particles observed | william.lee@sanofi.com | 8NYA/CNC | 2025-07-24 | 2025-07-24 | Pre inspection should pass. Particles should not appear in SUT.
5852 | 8NYA - Support Service- RF551F024 - Agitator pins found in filter while performing post transfer breakdown | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-07-27 | 2025-07-27 | Pins stay intact during operations
5851 | 8NYA - Support Service - RF564F015 recipe was manually launched under the incorrect volume | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-07-26 | 2025-07-26 | drain WFI and abort record
5850 | 8NYA - Support Service - RF564F015 SUM bag failed pre inspection | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-07-26 | 2025-07-26 | N/A | Yes | Deviation
5849 | CZ - RF639F016 - 8NYA - Automation Intervention to adjust Eluate collection QTY | michael.knapp@sanofi.com | 8NYA 1403 | 2025-07-26 | 2025-07-26 | Eluate count allows for maximum number of eluates to be collected.
5848 | 8NYA - Support Service - RF564F015 SUM bag leaked during initial WFI | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-07-26 | 2025-07-26 | SUT to no leak
5847 | 8NYA-Gemba-Downstream Fz SP Suite 1 discrepancy in SOP for eluate filtration | aidan.mula@sanofi.com | 8NYA/RM1323 | 2025-07-26 | 2025-07-26 | SOP should align with operation on the floor
5846 | 8NYA GEMBA- Suite 1 SP sampling bag not capped during sampling | alex.tello@sanofi.com | 1323 | 2025-07-26 | 2025-07-26 | For cap to remain on the sample bag | No | Deviation
5845 | 8NYA Downstream Fz Nano filter not fully bleed during equilibration | Christian.Venegas@Sanofi.com | 8NYA Downstream suite 1 room 1323 | 2025-07-26 | 2025-07-26 | To get a Phath forward
5844 | 8NYA - Support service - CZ- RF621F189_001 no second verify for manual entry | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-06-29 | 2025-07-25 | N/A
5843 | Missing access request forms | dennis.bazinet@sanofi.com | Record management room at 68NYA | 01/20/2025 and 03/20/25 | 2025-07-25 | Should have been able to show the training was completed | Yes | Deviation
5842 | 8NYA - Support Service - RF545F096 3X 3000L tote bags failed pre inspection | bailey.gleason@sanofi.com | 8NYA | 2025-07-24 | 2025-07-24 | SUT be ready to be used by MFG with no issues
5841 | 8NYA - Support service -CZ- RF621F248 - meter 2008051 incorrectly put out of service | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-07-25 | 2025-07-25 | Meter not be put out of service
5840 | 8 NYA - Support Services - 805007 - Agitation Rate - GARAMP Item | nicholas.grote@sanofi.com | 08-1350 | GARAMP | GARAMP | If speed is changed it should be done through the correct change mechanism with clear justification as to why.
5839 | 8 NYA Support Services - GARAMP Item - Pipette Additions | nicholas.grote@sanofi.com | 08-1350 | 2025-07-24 | 2025-07-24 | Each pipette addition will be documented to ensure all data is captured for the addition.
5838 | 8NYA- Support Service - CZ - RF613F028_003 filter tested under incorrect parameters | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-07-08 | 2025-07-25 | N/A
5837 | 8NYA Downstream Suite 1 - SOP Missing Instructions for probe 1 pt. standardization and Equilibration sampling | wilfredo.lopez@sanofi.com | 8RM1323 | 2025-07-24 | 2025-07-24 | It is expected that all necessary instructions for the operation are included in the SOP
5836 | 8NYA Downstream Suite 1 - Thermo Scientific P/N 692171 Failed Pre-Use Inspection | charles.salgadinho@sanofi.com | 8NYA Room 1313 | 2025-07-24 | 2025-07-24 | received bags don't have particulates and tears.
5835 | 8NYA Downstream Suite 2-Recipe for Skid 3323 went into hold during Standalone CIP | Jim.ramjattan@sanofi.com | 8NYA, Room 1503 | 2025-07-25 | 2025-07-25 | The CIP is supposed to run without any holds or alarms.
5834 | x2543 Unable to Acquire SUM, Automation to release. | ricardo.rodriguez@sanofi.com | 1333 | 2025-07-25 | 2025-07-25 | Automation to release EM.
5833 | 760741 POU Lot FGHAU177 was sampled as POU 760744 lot FGHBN014 | ricardo.rodriguez@sanofi.com | 8NYARM1403 Upstream Phase 2 | 2025-07-24 | 2025-07-24 | QE to be created to then be placed into the record. | No | Deviation
5832 | Osmometer standardization MES record was created as standalone record, not under parent record | yunbo.lou@sanofi.com | 8 NYA | 2025-07-23 | 2025-07-24 | The pH and Osmometer standardization record is required to be created under parent record
5831 | 8 NYA DS FZ Suite 1 Shipping Bag Automation Issue | katherine.reardon@sanofi.com | DS Post Viral 1 | 2025-07-24 | 2025-07-24 | scale tares weight of barrel | Yes | Deviation
5830 | P1-FZ MCC 2543x in hold for Column equilibration timer expires" needs automation intervention. | joseph.gonzalez@sanofi.com | 1333, 8NYA | 2025-07-24 | 2025-07-24 | timer to not expire. PH of buffer is in range | No | No Deviation
5829 | 8NYA- Support Services Gemba-Stainless Steel cover missing from scale | william.lee@sanofi.com | 8NYA/1350 | 2025-07-24 | 2025-07-24 | Stainless steel cover should be in place on scale
5828 | P1 Fz SUM bag failed inspection (3x) | elaine.jansen@sanofi.com | 8NYA 1333 | 2025-07-24 | 2025-07-24 | Bag good for use
5827 | HA - Powder Final Filling - FBYHAH005 - Room - 76/355 | Yamini.patcha@sanofi.com | Room 76-355 | 2025-07-23 | 2025-07-24 | QU-OPE-0039925 Step 7.4.8 - An empty overfill bottle should be tare and print the tare weight. Step 7.5.27 Over filled HA powder collected bottle should be weighted and printed.
5826 | HA - Powder Final Filling - FBYHAH005 - Room - 76/355 | Yamini.patcha@sanofi.com | Room 76-355 | 2025-07-23 | 2025-07-24 | QU-OPE-0039925 Step 7.4.8 - An empty overfill bottle should be tare and print the tare weight. Step 7.5.27 Over filled HA powder collected bottle should be weighted and printed.
5825 | All buildings - Dynamic EM Non-Viable Alert/Action levels not detected in OneLIMS | craig.stevenson@sanofi.com | MABC | 2025-07-24 | 2025-07-24 | It is expected that each of the Non-Viable results be assessed as a single sample in a collection batch and have the alert/action levels associated with reach individual sample.
5824 | All buildings - Dynamic EM Non-Viable Alert/Action levels not detected in OneLIMS | craig.stevenson@sanofi.com | MABC | 2025-07-24 | 2025-07-24 | It is expected that each of the Non-Viable results be assessed as a single sample in a collection batch and have the alert/action levels associated with reach individual sample.
5823 | Incorrect Cartridge Lot selected in the Nexus Software | Amy.Roy@sanofi.com | 4700 68 NYA QC Micro | 2025-07-14 | 2025-07-22 | Per QU-SOP-0070395 Step 6.3.22, the desired cartridge Lot number should be selected prior to starting assay run.
5822 | Incorrect Cartridge Lot selected in the Nexus Software | Pushti.Shah@sanofi.com | 4700 68 NYA QC Micro | 2025-07-14 | 2025-07-22 | Per QU-SOP-0070395 Step 6.3.22, the desired cartridge Lot number should be selected prior to starting assay run.
5821 | Insect found in grade D corridor at 51 NYA | Alyson.Belleville@sanofi.com | 51 NYA corridor 83 | 2025-07-23 | 2025-07-23 | there should be no pests in the controlled area
5820 | Incorrect Cartridge Lot selected in the Nexus Software | Pushti.Shah@sanofi.com | 4700 68 NYA QC Micro | 2025-07-14 | 2025-07-22 | Per QU-SOP-0070395 Step 6.3.22, the desired cartridge Lot number should be selected prior to starting assay run.
5819 | 8 NYA GEMBA - Support Services - Leak noticed during initial WFI fill | william.lee@sanofi.com | 8NYA/1350 | 2025-07-23 | 2025-07-23 | Bag should not leak.
5818 | Computerized Systems have not undergone Periodic Review & User Access Review | stephen.walker@sanofi.com | Massachusetts BioCampus | 2025-07-23 | 2025-07-22 | Per QU-SOP-0000910 (Steps 2.1 & 6.1.1), GxP-critical or major local and multi-site computerized systems must undergo periodic review for a computerized system every three years. Per QU-SOP-0051867 (Step 6.5.2), the user access review must be performed annually & every 3 years for GxP critical and major systems, respectively.
5817 | 8NYA-CZ-RF1BGF016X-non-GMP record version used for GMP operations | john.lytle@sanofi.com | 8NYA Rm 1403 | 2025-07-22 | 2025-07-23 | it is expected that cGMP operations be performed using a cGMP lot record
5816 | 8NYA - Support Service- FZ - RF573F114 Leak post filtration from Oetker clamp | William.lee@sanofi.com | 8NYA RM 1350 | 2025-07-22 | 2025-07-22 | SUT be ready to be used by MFG with no issues
5815 | Mens Gowning Room Accidental Booty Removal | Thomas.Loik@sanofi.com | 8NYA Mens Gowning | 2025-07-22 | 2025-07-22 | No event generated, routine cleaning and room monitoring to ensure integrity
5814 | 8NYA GEMBA - 2 Flies found in mens locker room- Rm1125 | carl.pierre@sanofi.com | 8 NYA -1125 | 2025-07-22 | 2025-07-22 | No bugs to be around
5813 | 8NYA - Support Service - FZ - RF525F016 SUT not allocated to record | bailey.gleason@sanofi.com | 8NYA RM 1373 | 2025-07-22 | 2025-07-22 | SUT be allocated to record
5812 | 8NYA- Support Service - Stock discrepancy RF621F183 | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-06-19 | 2025-07-22 | Stock be created for the correct amount | Yes | No Deviation
5811 | 8NYA - Support Services - Pests found in CNC/8RM1350 | bailey.gleason@sanofi.com | 8RM1350 | 2025-07-22 | 2025-07-22 | That no pest will be found in the controlled spaces of 8NYA.
5810 | 8NYA - Support Service - FZ - RF563F005 Condo sample OOS | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-07-22 | 2025-07-22 | solution be within specification range
5809 | Fz Zinc bag inspection failed due to particles found in bag | PAUL.DEVANEY@SANOFI.COM | 8NYA 1323 | 2025-07-22 | 2025-07-22 | No particulates should be found in the bag. | Yes | Deviation
5808 | 8NYA- Support Services- Condo out of range | Bailey.gleason@gmail.com | 8NYA/ 1350 | 2025-07-21 | 2025-07-21 | Result between range
5807 | 8NYA - Support Service - FZ - RF525F016 - nanofilter failed pre inspection | bailey.gleason@sanofi.com | 8NYA RM1373 | 2025-07-21 | 2025-07-21 | Hold gemba to see path forward |  | Deviation
5806 | Phase 1 UF\DF retentate vessel not empty after transfer | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1323 | 2025-07-21 | 2025-07-21 | To get a Phath forward for engineering
5805 | EMS alarm that occurred 28Jun2025 in rooms 1340 and 1343 with no comment. Without knowing the cause of the alarm, impact cannot be assessed. | dana.hopkins@sanofi.com | 8NYA ROOM 1340 AND 1343 | 2025-06-28 | 2025-07-17 | Alarms should have correct comments entered
5804 | 51 NYA Small Scale-Non Viable alarm during open processing FGFSB004 | nicole.lufkin@sanofi.com | 51NYA 152 B | 2025-07-21 | 2025-07-21 | climet should not alarm during open processing
5803 | 68NYA-Support APU-Metrology-Room 1835 Conductivity Meter FB-1853 overdue for calibration. | robert.bauch@sanofi.com | 68-1835 | 2025-04-16 | 2025-07-21 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
5802 | UFDF Rrecord referencing incorrect SOP | pavel.dobrovolskiy@sanofi.com | 8nya 1323 | 2025-07-21 | 2025-07-21 | Its expected that the record will call-out the correct SOP for operators to follow
5801 | 8NYA Downstream phase 1 Butyl tote SUT failed inspection 2x | Robert.pino@sanofi.com | 8NYA, RM1323, Suite 1 | 2025-07-21 | 2025-07-21 | It is expected that the butyl eluate tote bag would not fail inspection due to particulates
5800 | Phase 1 Downstream Butyl Eluate SUM went into alarm due to HIHI temperature | Robert.pino@sanofi.com | 8NYA, RM1323, Suite 1 | 2025-07-21 | 2025-07-21 | It is expected that the butyl eluate SUM would not go into hold and remain within 16-24C as per CST
5799 | Automation intervention needed for SUM V-2570001 | ricardo.rodriguez@sanofi.com | rm 1403 | 2025-07-20 | 2025-07-20 | Automation to release EM.
5798 | Suite 1 Upstream FZ RF503F008 Eluate SUM diaphragm lip unseated during mixing | jared.Watson@Sanofi.com | 8RM1333 | 2025-07-20 | 2025-07-20 | eluate SUM mix
5797 | 8 NYA GEMBA - Support Services - Leak noticed during initial WFI fill | william.lee@sanofi.com | 8NYA/1350 | 2025-07-20 | 2025-07-20 | SUM bag should not have leaked
5796 | 3x Failed 250L sum bags failed inspection due to multiple particles. | joseph.gonzalez@sanofi.com | 1403, 8NYA | 2025-07-18 | 2025-07-18 | No particulates in the bag | Yes | No Deviation
5795 | Butyl Filtration/Sampling record stuck on "Room Status" whilst also having FIT prematurely open | elias.mcdonnell@sanofi.com | 8NYA, 1324 | 2025-07-19 | 2025-07-19 | Step will automatically transition, and open up "area clearance". Also, FIT step will not be prematurley started
5794 | Butyl Filtration/Sampling record stuck on "Room Status" whilst also having FIT prematurely open | elias.mcdonnell@sanofi.com | 8NYA, 1324 | 2025-07-19 | 2025-07-19 | Step will automatically transition, and open up "area clearance". Also, FIT step will not be prematurley started
5793 | 8NYA-CZ-FGHBN016-MCC X2831 Load Inlet Pump leak identified | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2025-07-19 | 2025-07-19 | It is expected the system remain integral/ free of leaks
5792 | Production recipe stuck at "disabling failure monitor" automation needed to intervene. | joseph.gonzalez@sanofi.com | 1403, 8NYA | 2025-07-19 | 2025-07-19 | Recipe should not be stuck on ems.
5791 | GEMBA: Suite 1 Fabrazyme: Butyl: Buffer Consumption not possible during Night Ops. | William.Bradley@sanofi.com | 8NYA 1323 | 2025-07-19 | 2025-07-19 | MES works properly and as intended.
5790 | 8 NYA DS CZ GEMBA REQUEST- RoSS Issue at Zero Shelves Step | katherine.reardon@sanofi.com | 8 NYA 1703 | 2025-07-19 | 2025-07-19 | Arrows on RoSS should point to the shelves being used to reflect proper recipe parameters
5789 | Two butyl load bags failed visual inspection | KELVIN.GLEAN@SANOFI.COM | 8NYA 1323 | 2025-07-19 | 2025-07-19 | Bags to free of any black spots | Yes | No Deviation
5788 | Butyl Filtration/Sampling record stuck on "Room Status" while also not allowing manual progression | elias.mcdonnell@sanofi.com | 8NYA, 1324 | 2025-07-18 | 2025-07-18 | Step will automatically transition, and open up "area clearance". Also, FIT step will not be prematurley started
5787 | GEMBA: Fabrazyme Suite 1: SOP Discrepency; Removal of Devices | william.bradley@sanofi.com | 8NYA /RM1323 | 2025-07-18 | 2025-07-18 | To have a procedure with accurte and clear step by step instructions.
5786 | DS Phase 1 SP Samples Left out of sample fridge | katherine.reardon@sanofi.com | 8 NYA 1321 | 2025-07-18 | 2025-07-18 | Put samples in fridge
5785 | vial shipment to geel over 120 hours | nicole.lufkin@sanofi.com | 45 NYA GCB room 101 | 2025-06-30 | 2025-07-17 | loggers/shipment should be under 120 hours per SOP
5784 | Media lot 760741 FGHAU235 in Q status but discarded | joshua.brooks@sanofi.com | CNC space 8NYA | 2025-07-16 | 7172025 | Material disposition changes will occur in Athena prior to discard in MES.
5783 | 8NYA - Support Service - FZ- RF557F045/705663 e stop alarm was initiated. | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-07-14 | 2025-07-16 | N/A
5782 | Cerezyme - RF639F018 - 8NYA - DO Probe Divergence on SUB X2220 | michael.knapp@sanofi.com | 8NYA 1403 | 2025-07-16 | 2025-07-16 | Probes are representative of the culture and do not display evidence of fouling.
5781 | Downstream Suite 1/2 -DS Shipping bag fill volume below minimum fill volume for E/L | May.chin@sanofi.com | 8NYA Suite 2 Post Viral | 2023-05-09 | 2025-07-16 | DS fill volumes meet the minimum fill volume requirement per MABC-QU-OPE-2553347
5780 | 2560 High buffer column Inlet pressure during blank run 3 | brett.simmons@sanofi.com | 8NYA, RM1403 | 2025-07-15 | 2025-07-15 | Buffer column inlet pressure should stay below the HiHi limit of 20.0 psi.
5779 | Two Bags Failed Pre-Use Inspection | JOEL.HICKEY@SANOFI.COM | 8 NYA/Room: 1503. | 2025-07-13 | 2025-07-13 | Per QU-SOP-0084572 - SUM bag's port seal needs to be complete (step 6.2.19) / Determination needs to be made as to whether barrel bag can be used (step 6.2.13).
5778 | 8NYA SUT leak detected on 760741/FGHAU235. | Kelley.Gibson@sanofi.com | 8NYA-CNC AREA 1210 | 2025-07-15 | 2025-07-15 | No leaks
5777 | SD filled out incorrectly | tanni.arif@sanofi.com | CMS | 2025-07-14 | 2025-07-14 | L&D and QA needed on SD form approval
5776 | Missed LAL Nexus Daily Maintenance for FB-2079 and FB-3359 | Pushti.Shah@sanofi.com | 4700 68 NYA QC Micro | 2025-07-14 | 2025-07-15 | Per QU-SOP-0070395 Step 6.1.1, Daily maintenance must be performed only on days the Nexus is in use.
5775 | Missing "Measurements" Steps in Q- Membrane Records | katherine.reardon@sanofi.com | 8NYA 1503 | 2025-07-15 | 2025-07-14 | Records should have all parts of the "Measurements" steps in order to perform measurements of drive mixers
5774 | GEMBA: 8NYA Suite 1: Delta V Issues and Zinc eluate warm | William.Bradley@sanofi.com | 8NYA 1323 | 2025-07-15 | 2025-07-15 | This process is to run off of an automated recipe.
5773 | GEMBA; 8NYA Downstream: Fabrazyme Suite 1: Pre-Use Bag Inspection failure | William.Bradley@sanofi.com | 8NYA 1323 | 2025-07-14 | 2025-07-14 | It is expected that bags are in good integrity with no particulates
5772 | FOLLOW-UP GEMBA: Phenyl eluate SUM; agitator not mixing after launching mix recipe and weight on DeltaV reads 0 | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2025-07-14 | 2025-07-14 | Agitator will mix and the weight in deltaV will match that of the SUM faceplate
5771 | Phenyl eluate SUM; agitator not mixing after launching mix recipe and weight on DeltaV reads 0 | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2025-07-14 | 2025-07-14 | Agitator will mix and the weight in deltaV will match that of the SUM faceplate
5770 | SUM bag failed inspection (4x) | elaine.jansen@sanofi.com | 8NYA 1333 | 2025-07-14 | 2025-07-14 | n/a
5769 | Suite 1 FZ SUT GEMBA Request- RF1BRF033 Bag Failure | katherine.reardon@sanofi.com | 8NYA 1303 | 2025-07-11 | 2025-07-11 | No particulates in SUT
5768 | 8NYA- Upstream - Suite 2- Cerezyme 500L Bioreactor End of Run Lot FGHBN013 Samples not submitted to QC Micro | amy.roy@sanofi.com | 8NYA | 2025-06-07 | 2025-07-14 | It is expected that end of run sample anaerobic bioburden samples to be delivered to QC Micro within the next drop off and before 24 hour sample expiry.
5767 | 8 NYA GEMBA - Support Services - Leak noticed at transfer pump line | william.lee@sanofi.com | 8NYA/1350 | 2025-07-14 | 2025-07-14 | There should be no leak during transfer
5766 | 68NYA QCM support APU-Room 4700-Dew point Hygrometer F-17408 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-4700 | 2025-07-01 | 2025-07-14 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
5765 | TSH - 14TP010 - 45 NYA - Data loggers not stopped within 120 hours for vial shipment | MARIE.VALLIERE@SANOFI.COM | 45 NYA, room 101 | 2024-06-21 | 2025-07-13 | It is expected that the shipment occurs within the validated time requirement of 120 hours by reviewing the "Start at" and "Stop at" date and times in the logging results section of the data logger reports per step 6.5.3 of MABC-QU-SOP-0070997-Receipt and Shipment of Banks and Bank Vials at Framingham Biologics (v17.0).
5764 | 8NYA-SS-GEMBA, Leak discovered on 705683 | Kelley.Gibson@sanofi.com | 8NYA-CNC1210 | 2025-07-14 | 2025-07-14 | No leaks
5763 | 51NYA - Non-Viable Failures in room 51-134 | craig.stevenson@sanofi.com | 51NYA Room 134 | 2025-07-13 | 2025-07-13 | All Non-Viable results are expected to generate passing data.
5762 | Possible crack on Column 3 and Regen buffer port | joseph.gonzalez@sanofi.com | 1403, 8NYA | 2025-07-14 | 2025-07-14 | Possible re-SIP
5761 | 2X 250L Sum bag failures | Phoebe.karkos@sanofi.com | Upstream Suite 1 | 2025-07-12 | 2025-07-12 | SUT should pass inspection criteria
5760 | 8NYA - Support service - MABC-QU-SOP-0083660 enzyme thaw instructions are out of order | bailey.gleason@sanofi.com | 8NYA | 2025-07-12 | 2025-07-12 | Instructions be in order
5759 | 8NYA - Support service - Functionality available but no instructions on how to correct stock creation | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-07-11 | 2025-07-11 | Stock be created for the correct amount
5758 | 8NYA - Support Service - Incorrect value inserted into SAP for bottle volume of P/N 844452 | bailey.gleason@sanofi.com | 8NYA | 2025-07-12 | 2025-07-12 | Values be correct in SAP
5757 | 8NYA - Support Service - WFI drop from loop 3B was used on SUM 1505 without closure of AN 2025-I-047 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-07-11 | 2025-07-11 | N/A
5756 | Gemba need: Non-uniformities found during visual inspection of UFDF skid. | kenny.leland@sanofi.com | 8 NYA 1323 | 2025-07-11 | 2025-07-11 | No Non-uniformities should be found.
5755 | Insect discovered in enclosed ceiling light fixture in room 229 of the HA-CMC area | Phillip.Mank@Sanofi.com | 76 NYA | 2025-07-11 | 2025-07-11 | It is expected that no insect would be discovered within the ceiling light.
5754 | Insect discovered in enclosed ceiling light fixture in room 229 of the HA-CMC area | Phillip.Mank@Sanofi.com | 76 NYA | 2025-06-27 | 2025-06-27 | It is expected that no insect would be discovered within the ceiling light.
5753 | Media 760741 batch records for 3 lots from 2019 cannot be located | elizabeth.sweeney@sanofi.com | 8 NYA Solution prep | 2025-07-11 | 2025-07-11 | Batch records would be able to be located
5752 | Missing performed by for loading vial shipper on 11JUN24 | nicole.lufkin@sanofi.com | 45 NYA GCB room 101 | 2024-06-11 | 2025-07-11 | performed by should have been signed for at time of vial shipment
5751 | Discrepancy between MES (760738 ) and MABC-QU-OPE-0044689-Cz-ICB Sampling Plan for 8 NYA | julie.auguste@sanofi.com | Solution prep, 1350 | 2025-07-09 | 2025-07-10 | It is expected that the EBR on MES matches the SOP to provide clear instructions to operators.
5750 | 8NYA-CZ-X2831 atypical observation in Glycol waste line following high pressure event | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2025-07-10 | 2025-07-10 | it is expected that the glycol waste filters perform as intended
5749 | GEMBA - 8 NYA DSP Suite Two - Slurry Percentage Average Out Of Range for MMC Column Pack | adam.keister@sanofi.com | 8 NYA 1303 | 2025-07-10 | 2025-07-10 | For slurry percentage average to be withing acceptable range and a path forward. | Yes | Deviation
5748 | CZ-MCC X2831 leak at Column 3 Colder Connection | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2025-07-10 | 2025-07-10 | It is expected that MCC remain integral
5747 | 51NYA - Upstream Dynamic EM - No reportable results for Zero Filter | craig.stevenson@sanofi.com | 51 NYA | 2025-06-08 | 2025-07-10 | Per QU-SOP-0074750-Use of the Climet CI-1054 (v12.0) Step 6.4, a zero filter should be performed prior to Non-Viable sampling.
5746 | CZ SUT ATF 10 Visual inspection failure | jeffrey.boldizar@sanofi.com | 8NYA1403 | 2025-07-09 | 2025-07-09 | no particulates present
5745 | Leak on Colder Steam-thru Connector | elaine.jansen@sanofi.com | 8NYA 1333 | 2025-07-09 | 2025-07-09 | n/a
5744 | Particulate found during ATF10 inspection. | joseph.gonzalez@sanofi.com | 1403, 8NYA | 2025-07-09 | 2025-07-09 | Receive new ATF with no particulates.
5743 | 8NYA - Support service - RF621F235 leak discovered on pump head tubing during transfer | william.lee@sanofi.com | 8NYA/1350 | 2025-07-09 | 2025-07-09 | There should be no leak during transfer.
5742 | 51 NYA - Cell Culture - WO 3342509 not completed | MARIE.VALLIERE@SANOFI.COM | 74 NYA, room 166, cryo storage room | 2025-06-30 | 2025-07-09 | It is expected that a monthly self-inspection work order is executed for 74 YA room 166 per MABC-QU-SOP-0075222-Walk-Through Procedure for MA BioCampus (v21.0)
5741 | CZ-X2230 FGHBN014 elevated O2 levels discussion | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2025-07-09 | 2025-07-09 | it is expected that O2 levels be monitored and trend analysis performed to ensure SUB is performing as expected |  | Deviation
5740 | Digital Systems Misalignments | ilene.robinson@sanofi.com | 68 NYA | 2025-07-08 | 2025-07-08 | The Control strategy is aligned with the Dossier and digital system
5739 | Gemba – 8NYA – Insect captured in room 1410 | william.lee@sanofi.com | 8NYA/1410 | 2025-07-09 | 2025-07-09 | Insect should not be in 1410 MAL.
5738 | SUB 2210 unstable perfusion due to Media tote touching front rail | elon.elzam@sanofi.com | Phase 1 Upstream room # 1333 | 2025-07-06 | 2025-07-06 | perfusion needs to run at a rate of 2 reactor volumes per day (200L per day). Due to not having a flow meter the combined harvest pump is used an an indicator for weight control and it needs to stay below 54RPM. At about 1:30am on Sunday 07/06 the RPM went above 54RPM and was inconsistent till tote replacement at 6:45pm that same day. |  | Deviation
5737 | Room 1363 not enrolled in RQ Program | vanessa.cardona2@sanofi.com | 08-1363 | 2020-11-04 | 2025-07-08 | Room 1363 should be part of the RQ program.
5736 | 8NYA - Support Services - SUM bag leaked during initial WFI fill | maria.palmar@sanofi.com | 8 New York Ave, room 1350. | 2025-07-08 | 2025-07-08 | The bag to not leak
5735 | Seprafilm FBYSEP046 room 232 occupancy limit exceeded | Kenneth.morgan@sanofi.com | 76-232 | 2025-07-08 | 2025-07-08 | occupancy limit should be followed
5734 | No Product – No Lots – 68 NYA - No reportable results for Sample Sites 8-PreT-1 and 8-PreT-4 week of 22JUN25 | Alyssa.Barrett@sanofi.com | 8 NYA/ room 1170 | 2025-06-22 | 2025-07-08 | It is expected that the samples listed in the Action Notice are sampled accordingly and at the frequency specified within the Action Notice per QU-SOP-0022403 | Yes | Deviation
5733 | 8NYA-Gemba-Downstream- Suite 2 Column Packing- Cleaning EUR asks to generate manual labels for CV sampling instead of LIMS | aidan.mula@sanofi.com | 8RM/1303 | 2025-07-08 | 2025-07-08 | Record should align with the SOP requirements
5732 | Sample transport Credo Cooler total volume over 750mL | dennis.bazinet@sanofi.com | 8 NYA / 1410 sample pick up room | 2025-04-09 | 2025-07-03 | The transport of samples per credo cooler not to exceed <750mL
5731 | X-2543 MCC column 3 pH probe not responding as expected | dana.hopkins@sanofi.com | 8NYA room 1333 | 2025-07-08 | 2025-07-08 | probe should respond as expected
5730 | 8NYA - Support Service - Osmometer was incorrectly put out of service | bailey.gleason@sanofi.com | DSP Phase 2 | 2025-07-07 | 2025-07-07 | hold gemba to see path forward
5729 | 8NYA- Support service - Incorrect deallocation of FIT 2009073 | bailey.gleason@sanofi.com | 8NYA 1350 | 2025-07-08 | 2025-07-08 | Hold gemba to see path forward
5728 | 8NYA - Support service - RF621F232 SUT failure on pump head during transfer | Bailey.gleason@sanofi.com | 1350 | 2025-07-07 | 2025-07-07 | SUT be manufactured correctly for operations
5727 | CZ - RF639F014 - 8NYA - DO Trend Review | michael.knapp@sanofi.com | 8NYA 1403 | 2025-07-07 | 2025-07-07 | Accurate DO probe readings lead to proper O2 control within bioreactor
5726 | 68NYA Support APU-QCC68-Digital Thermometer FB-4615 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-2700 | 2025-07-02 | 2025-07-07 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5725 | Automation Intervention needed for Production start x2831. | Dana.Hopkins@sanofi.com | 8 NYA phase 2 rm1403 | 2025-07-06 | 2025-07-06 | Prompts appear as they are written in the SOP
5724 | Harvest overflow Pump(2816) running While MCC was collecting & harvest SUM weight under required amount for pump to start. | Dana.Hopkins@sanofi.com | 8 NYA Phase 2 US | 2025-07-07 | 2025-07-07 | Gemba being held to determine path forward. |  | No Deviation
5723 | Particulates found in SUM Bag | Dana.Hopkins@sanofi.com | 8 NYA Rm 1403 phase 2 upstream | 2025-07-06 | 2025-07-06 | Gemba being held to determine path forward.
5722 | GEMBA: 8NYA Fabrazyme Suite 1 LOW LOW bubble trap Alarm | william.bradley@sanofi.com | 8NYA /RM1503 | 2025-07-06 | 2025-07-06 | Process should run without issues or alarms | Yes | No Deviation
5721 | Particulates found in SUM | elaine.jansen@sanofi.com | 1333 | 2025-07-06 | 2025-07-06 | n/a
5720 | 8NYA GEMBA - Support services - Weigh and dispense discrepancy | JACK.LAPLANTE@SANOFI.COM | 8NYA | 2025-06-14 | 2025-07-05 | Looking for path forward.
5719 | Butyl Regen/Sanitization buffer 705662 outside of butyl production temperature operating parameters | Robert.pino@sanofi.com | 8NYA, Butyl CNC Space | 2025-07-06 | 2025-07-06 | Per MABC-QU-OPE-0524531 table 3.5 on page 19, the temperature for Butyl operations must be within 16-24C | Yes | Deviation
5718 | Skid Setup for Cleaning Step Missing From Skid Cleaning EUR | zachary.rickard@sanofi.com | 8NYA/1303 | 2025-07-04 | 2025-07-04 | An MES step would exist to prompt the operator to go to the SOP and perform the steps associated with setting up the qual skid for a cleaning. | No | Deviation
5717 | 8NYA Downstream Suite 1 - SOP Missing Instructions for DeltaV Prompt | wilfredo.lopez@sanofi.com | 8RM1323 | 2025-07-05 | 2025-07-05 | Expected that the SOP contains all necessary instructions for the operation. Document needs to be revised to
5716 | GEMBA - 8 NYA DeltaV/PI Data Loss Since 29JUN2025 for 46 Tags | devyanshu.raj@sanofi.com | 68 NYA, 4th Floor. In 8NYA DeltaV Production System | 2025-06-29 | 2025-06-30 | PI Tags are expected to collect historical data
5715 | Incorrect record used for inflation on X2220 | Ralston.najarro@sanofi.com | 8NYA1403 | 2025-07-02 | 2025-07-03 | Record should not be available to be used in MES
5714 | 8NYA possible all building EMS is not working giving white screen ticket opened UR0237498 | Daniel.bati@sanofi.com | 8NYA | 2025-06-30 | 2025-07-03 | should read all buildings alarms
5713 | X-2210 FZ Train 2 weight controller output nuisance alarms | dana.hopkins@sanofi.com | 8NYA room 1333 | 2025-07-03 | 2025-07-03 | No nuisance alarms
5712 | GEMBA; 8NYA: Downstream: Pre Use bag Failure ( 650L P/N 813910) | William.Bradley@sanofi.com | 8NYA /RM1323 | 2025-07-02 | 2025-07-03 | SUT is to be in good standing condition without deformities | Yes | No Deviation
5711 | Thyrogen LAL sample not performed in OneLIMS FGFMB003 | kevin.ornell@sanofi.com | 51-124 | 2025-07-02 | 2025-07-03 | Perform Step 4.9.2, remove samples per QU-SOP-0057386 | Yes | Deviation
5710 | 8NYA- Support Services- Pump Tubing Leak | Maria.Palmar@sanofi.com | 8NYA/ 1350 | 2025-07-02 | 2025-07-02 | tubing expected to not leak
5709 | (2x) 250L SUM bag failed visual inspection part#692168 | Dana.Hopkins@sanofi.com | 8NYA, Room 1333 | 2025-07-02 | 2025-07-02 | We expected the bag to pass inspection
5708 | Wrong time documented on BIRF | Luis.dominguez@sanofi.com | 51NYA room 87 | 2025-06-27 | 2025-07-02 | per QU-SOP-0074092 times should be documented during storage and removal of product.
5707 | 8NYA butyl adjust load bag-- black particulate observed embedded on OUTSIDE of bag | elias.mcdonnell@sanofi.com | 8NYA, 1324 | 2025-07-01 | 2025-07-01 | Per SOP-0084572, step 6.2.13, particulates should not be found. If particulates are embedded within the exterior, the bag can be cleared or failed by QA. GEMBA must be called.
5706 | Process waste tote door unsecured resulting in bag bulging. | Kelley.Gibson@sanofi.com | CNC ROOM 1210 8 NYA | 2025-07-01 | 2025-07-01 | Tote is used to contain bag to collect process waste, and doors to tote are locked
5705 | Heat Sealer test pouch from lot EBYSEP058B p/n 789446 was found during labeling/pkg at 74 NYA | alan.rousseau@sanofi.com | Room 1017 74 NYA | 2025-06-30 | 2025-07-01 | It was expected operator would find 10 pouches not 11(which includes test pouch)
5704 | 8NYA- Support Services Gemba-leak on 760723 transfer line | william.lee@sanofi.com | 8NYA/CNC | 2025-07-01 | 2025-07-01 | Transfer line should not have ruptured and sealer should not have been placed on transfer line.
5703 | 8 NYA GEMBA Suite 1 Upstream Sub 2210 Over Perfusing Train 2 PN: 706930 Lot: FGHCE008 MO: RF503F008 | GARY.DODAKIAN3@SANOFI.COM | 8 NYA room 1333 | 2025-06-30 | 2025-07-01 | Perfusion rate remains stable and media coming out of tote matches the rate at which the load cells dissipate
5702 | 8NYA- Support Services Gemba-weekly clean expired | william.lee@sanofi.com | 8NYA/1350 | 2025-06-30 | 2025-06-30 | Cleaning on 29JUN25 should have been completed and verified.
5701 | 76 NYA WFI pumps faulted | john.mendenhall@sanofi.com | 76-347 | 2025-07-01 | 2025-07-01 | WFI pumps would stay running
5700 | FZ, SUT waste bag inspection failure PN: 692181 | jeffrey.boldizar@sanofi.com | 8NYA 1210 | 2025-06-30 | 2025-06-30 | no particulate present | Yes | Deviation
5699 | TOC deviation investigational sample handling pretreatment side (No impact) | shiful.kohli@sanofi.coml | 1st floor 68NYA | 2025-06-30 | 2025-06-30 | Samples are not handled per Procedure XX
5698 | Gemba needed for Suite 1 Fz Unplanned differential pressure alarms in room 1313. | kenny.leland@sanofi.com | 8 NYA 1313 | 2025-06-30 | 2025-06-30 | Unplanned DP alarms should not occur. | Yes | No Deviation
5697 | 51NYA - Action Notice procedure not followed | Craig.stevenson@sanofi.com | 51 NYA Room 125 | 2025-06-24 | 2025-06-30 | It is expected that the process owner responsible for scope change amends the A/N and re-route for approval.
5696 | GEMBA: 8NYA Downstream Suite 1 : Pre-Use Bag failures: ( P/N 692299) | William.Bradley@sanofi.com | 8NYA /RM1323 | 2025-06-29 | 2025-06-30 | It is expected that bags are in good integrity with no particulates
5695 | SUB X-2210 multiple holds for SUB perfusion EM | dana.hopkins@sanofi.com | 8NYA room 1333 | 2025-06-29 | 2025-06-29 | Recipe should not hold | No | No Deviation
5694 | 8 NYA Support Services - Post formulation filter - Bleed valve broken prior to Fit | maria.palmar@sanofi.com | 1350 | 2025-06-28 | 2025-06-29 | Filter integral prior to FIT
5693 | 8 NYA Support Services - 500L Recirculation tote bag installation missing steps on the SOP | maria.palmar@sanofi.com | 8 New York Ave | 2025-06-28 | 2025-06-28 | To have clear instructions on the SOP for tote bag installation | Yes | Deviation
5692 | 8NYA GEMBA- FZ Nano leak at inlet tri-clamp connection. | Alex.tello@sanofi.com | 1323 | 2025-06-28 | 2025-06-28 | for no leaks to be observed | Yes
5691 | 8 NYA Support Services - Cytiva clamp part number referenced in SOP | bailey.gleason@sanofi.com | 1373 | 2025-06-28 | 2025-06-28 | SOP and part number should match
5690 | Igloo coolers used in Sample Transport process for submissions at room temperature | cleiton.lopes@sanofi.com | 68NYA, 51NYA, 45NYA | 2025-06-27 | 2025-06-27 | maintain sample integrity during transport
5689 | 8NYA GEMBA FZ UFDF concentration QC results out of range on final yield. | alex.tello@sanofi.com | 1323 | 2025-06-27 | 2025-06-27 | Final yield to be above 85%
5688 | 76 NYA Product Lines- Multiple Lots - 76 NYA - 76NYA building wide power outage | Phillip.Mank@Sanofi.com | 76 NYA | 2025-06-27 | 2025-06-27 | UPS system would be engaged after power failure
5687 | GEMBA needed for repeated missing disposition types for entries in 45RM190 Sample Retain Log | jesse.demers@sanofi.com | 45RM190 | 2025-06-27 | 2025-06-27 | Log should be filled out correctly
5686 | Missing logbook for 51 NYA Q2 Equipment prep logbook audit | jesse.demers@sanofi.com | 51RM143 | 2025-06-27 | 2025-06-27 | Logbook should be accounted for | Yes | Deviation
5685 | 705682 Tote leaking | Bradd.Charles@sanofi.com | 8 NY AVE Room 1202 | 2025-06-26 | 2025-06-26 | MABC-QU-SOP-0084572, MABC-QU-SOP-0083660
5684 | GEMBA: 8NYA: Downstream Suite 1 Fabrazyme: Pre-use Bag Inspection failure | William.Bradley@sanofi.com | 8NYA /RM1323 | 2025-06-26 | 2025-06-26 | It is expected that bags are in good integrity with no particulates | Yes | Deviation
5683 | ATF 10 Particulates | brian.o'rourke@sanofi.com | 8NYA Room 1403 | 2025-06-26 | 2025-06-26 | ATF housing should be free of particulates, debris, voids, or liquid upon initial inspection.
5682 | Suite 1 FZ SUT GEMBA Request- RF1BRF029 failed eluate SUM bag | katherine.reardon@sanofi.com | 8NYA 1303 | 2025-06-26 | 2025-06-26 | No particulates in SOP
5681 | GEMBA - 8 NYA DSP Suite One - Cannot Acquire Collection SUM in UFDF Recipe | adam.keister@sanofi.com | 8 NYA Room 1323 | 2025-06-26 | 2025-06-26 | Clarity on how to proceed with recipe and steps.
5680 | Multiple instances: missing entries and missing verification for steps in TSK packing record FGFCP001. | Alyson.Belleville@sanofi.com | 51-134 | 2025-06-09 | 2025-06-25 | It is expected that all steps would be filled out and verified contemporaneously. | Yes | Deviation
5679 | Wrong clamp removed from media tote when setting to pull POU | dennis.bazinet@sanofi.com | 8NYA CNC / 1210 | 2025-06-25 | 2025-06-25 | Clamp remove and POU sample to be taken
5678 | 8NYA Fz UF/DF RF1CBF008 collection SUM bag failed pre-use visual inspection | edgar.vale@sanofi.com | 8NYA, Room 1323 | 2025-06-25 | 2025-06-25 | SUM collection bags would be free of any particulates or abnormality | Yes | Deviation
5677 | 8NYA Downstream Suite 1 - Sartorius SUM Bag PN 813907 Failed Pre-use Inspection | wilfredo.lopez@sanofi.com | 8RM1323 | 2025-06-25 | 2025-06-25 | Determine if bag is acceptable for use as SUT cannot contain particulates. | Yes | Deviation
5676 | LAL Sample Fabrazyme Sepharose Capto Butyl Pool Vial found cracked upon testing | Pushti.Shah@sanofi.com | 4700 68 NYA QC Micro | 2025-06-25 | 2025-06-25 | Per QU-SOP-0070558 Step 6.2.3, sample containers should not be cracked, leaking. Expectation is that sample tubes are not cracked.
5675 | UFDF Membrane Failed Deallocation from Previous Record | JOEL.HICKEY@SANOFI.COM | 8 NYA/Room: 1323 | 2025-06-25 | 2025-06-25 | Membranes should be released for operations.
5674 | Action Notice 2024-F-074 Had Incorrect Sampling # (51-WFI-11 Instead of 51-WFI-9) | tanni.arif@sanofi.com | 51NYA Mech Room, WFI | 2025-06-22 | 2025-06-24 | 51-WFI-11 is the correct sampling site
5673 | HA pre purification FBYHAS002 agitator coupling for V-3511 broke during mixing | Kenneth.morgan@sanofi.com | 76-3101 | 2025-06-25 | 2025-06-25 | Mixing continues until stopped by manufacturing
5672 | GEMBA 8NYA Upstream Suite 1 High levels of condensation on eluate SUMS and chilled water lines | GARY.DODAKIAN3@SANOFI.COM | 8 NYA Upstream room 1333 | 2025-06-24 | 2025-06-24 | Lower levels of condensation on chilled water lines and eluate sums
5671 | Trying to set up Eluate sum. Tote is OOS. | joseph.gonzalez@sanofi.com | 1403, 8NYA | 2025-06-25 | 2025-06-25 | Tote to be put back in service after verifying the WO is complete in INFOR.
5670 | 8NYA-FZ-RF503F006 Failed weld during Daily Sampling activities | john.lytle@sanofi.com | Rm1333, 8NYA | 2025-06-25 | 2025-06-25 | successful weld
5669 | 8NYA - Unknow Lots - Unknown products - freezer Alarms, for 1421 8NYA | Daniel.bati@sanofi.com | 8NYA/ CNC | 2025-06-23 | 2025-06-25 | Should not have too many alarms
5668 | Leak on pendotech connection ATF B pump tubing SUB 2210 | elaine.jansen@sanofi.com | 8NYA 1333 | 2025-06-25 | 2025-06-25 | QU-SOP-0083848-ATF6 Operations at 8 NYA step 6.3 Guard Filter Set Up | Yes | Deviation
5667 | 8NYA Clean steam loss of Pressure from 24JUN25 1027 to 25JUN25 0035 | andrew.apel@sanofi.com | 8 NYA / 1170 | 2025-06-24 | 2025-06-24 | Clean steam to not lose pressure.
5666 | 8NYA- Support Services Gemba-Incorrect filter test on post viral filter | william.lee@sanofi.com | 8NYA/1350 | 2025-06-24 | 2025-06-24 | Correct recipe should have been used and filter tested.
5665 | X2230 DO Probe Deviation Alarm | sean.gormley@sanofi.com | 8RM1403 | 2025-06-25 | 2025-06-25 | Probes should not become fouled during course of run. | Yes | No Deviation
5664 | Zinc load SUM bottom outlet line was not clamped fully - Fluid passed the clamp. | paul.devaney@sanofi.com | 8NYA/1323 | 2025-06-24 | 2025-06-24 | Clamp should have been fully engaged and kept product from entering one of the lines on the load SUM
5663 | X2240 ABER Probe 1 & 2 potential offset prior to inoculation | Ralston.najarro@sanofi.com | 8NYA1403 | 2025-06-24 | 2025-06-24 | ABER probes should not be offset
5662 | UF/DF membrane installation recipe deallocation issues | Christian.Venegas@Sanofi.com | 8NYA Downstream suite 1 room 1323 | 2025-06-24 | 2025-06-24 | Looking for Path forward to deallocate UF/DF skid
5661 | Four (4) Daily WFI Work Orders Not Completed Contemporaneously | paul.mclaughlinjr@sanofi.com | 8NYA | 2025-06-24 | 2025-06-24 | All work orders would be completed at time of action
5660 | missing sample delivery start and end times | craig.valery@sanofi.com | 45 NYA | 2025-05-27 | 2025-06-23 | it is expected that times are documented
5659 | 8NYA-Gemba-Downstream Suite 2 Cz-Multiple shipping bags with embedded exterior particulate found pre-use | aidan.mula@sanofi.com | 8NYA/RM1703 | 2025-06-23 | 2025-06-23 | Bags should not contain particulates
5658 | Automation intervention needed to release Fz UFDF skid to perform WFI sanitization | paul.devaney@sanofi.com | 8NYA/1323 | 2025-06-18 | 2025-06-18 | WFI valves needed to be released from X3071 to perform Sani/flush | Yes | No Deviation
5657 | Analyst missed documentation of the BSC Usage Log prior to testing Biological Indicators. | Pushti.Shah@sanofi.com | 4700 68 NYA QC Micro | 2025-06-16 | 2025-06-17 | Per QU-SOP-0075304 STEP 6.5.3 (Document pre- and post-use cleaning/disinfection on FBL-LBK-000202) and STEP 6.3.4 (Ensure magnehelic gauge is within the acceptable range for the BSC and document on FBL-LBK-000202).
5656 | X2731 Eluate SUM Chilled water supply turned off | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-06-23 | 2025-06-23 | Chilled water supply should be on | Yes | Deviation
5655 | Fz ATF 6 SOP missing instruction to clamp A2B tee vent filter | marcelo.dossantos2@sanofi.com | 8NYA Room 1333 | 2025-06-23 | 2025-06-23 | We expect the SOP to have all the instructions need it to perform the operation correctly
5654 | Post Viral Sum temp probe sleeve is folded on itself and ripped | charles.salgadinho@sanofi.com | 8NYA, 1703 | 2025-06-23 | 2025-06-23 | The port for the temp probe should not be folded or blocked and the probe's sleeve should not be ripped
5653 | pH Standardization step SUM bag pH probe leak | Robert.pino@sanofi.com | 8NYA, RM1323 | 2025-06-22 | 2025-06-22 | The pH probe housing is not expected to leak | Yes | Deviation
5652 | 8NYA - GEMBA Downstream Suite 1 - Particulates Found In Multiple SUM Bags During Bag Inspections | danny.martinez@sanofi.com | 8NYA RM1323 | 2025-06-22 | 2025-06-22 | Expected result is for the SUM bag to not have any particulates. | Yes | Deviation
5651 | Particulates found during pre use inspection of ATF6 | jared.Watson@Sanofi.com | 8RM1333 | 2025-06-22 | 2025-06-22 | no particulates within the ATF
5650 | Futura Temperature < 50 C error message on transmitter after plugging in ABER probe during standardization. | joseph.gonzalez@sanofi.com | 1403 | 2025-06-22 | 2025-06-22 | Probe to show no error message when plugged in.
5649 | 8NYA- Support Services Gemba-leak on 705691 recirculation line leak | william.lee@sanofi.com | 8NYA/1350 | 2025-06-22 | 2025-06-22 | Agitator should not have fallen out of place and the SUT at the recirculation line should not have leaked.
5648 | 8NYA Downstream Suite 1 - Bags PN 692299 Failed Pre-use Inspection | wilfredo.lopez@sanofi.com | 8RM1323 | 2025-06-22 | 2025-06-22 | It is expected that SUT be free from contaminants and pass acceptable criteria for use | Yes | Deviation
5647 | Large Particulate found Mixing Inside x2220 | John.Lytle@sanofi.com | 8NYARM1403 Upstream Phase 2 | 2025-06-22 | 2025-06-22 | Gemba being held to determine path forward.
5646 | Decent sized particulate found Mixing Inside X220 | ricardo.rodriguez@sanofi.com | 8NYARM1403 Upstream Phase 2 | 2025-06-22 | 2025-06-22 | Gemba being held to determine path forward.
5645 | Whille doing Regin on skid 3322 during the prime the PH did come in on Delta v. but on the ph meter the was 13.3. we did a single point and re start but did not come in on prime | KELVIN.GLEAN@SANOFI.COM | 8NYA 1503 | 2025-06-20 | 2025-06-20 | Ph reading would be within acceptable range
5644 | Incorrect luer lock type on P/N 845774 Lot 25GN00452 | david.loik@sanofi.com | 8nya | 2025-06-20 | 2025-06-20 | luer lock is female
5643 | 813908 Bag Failed Pre-Use Inspection | JOEL.HICKEY@SANOFI.COM | 8 NYA/Room: 1503. | 2025-06-20 | 2025-06-20 | Per QU-SOP-0084572 step 6.2.13 it must be determined whether bag is acceptable for use.
5642 | 8NYA Upstream GEMBA Suite 1 SUB 2210 Bottom Bag pn 689204 Tabs Suggested to be to Utilized in Hooks by Repligen | GARY.DODAKIAN3@SANOFI.COM | 8 NYA room 1333 suite 1 upstream | 2025-06-20 | 2025-06-20 | N/A
5641 | Pump tubing welded onto left transfer line instead of center inoculation line | Jacob.Roman@sanofi.com | 8RM1333 | 2025-06-20 | 2025-06-20 | per QU-SOP-0080320 step 6.30.6, 50L wave: ¼ “ pump tubing to ¼ “ center inoculation line
5640 | 8 NYA Support Services SUM bag failed inspection | William.Lee@sanofi.com | 8 NYA 1350 | 2025-06-20 | 2025-06-20 | Passing inspection with no particulates | Yes | Deviation
5639 | 51NYA: Nano Pump X-2610 - DeltaV validation | ines.bica@sanofi.com | 51NYA | 2025-06-20 | 2025-06-20 | DeltaV system on pump validated
5638 | 8NYA - CZ - RF639F015 - MCC Skid 2560 High Pressure Identified on Column 1 during Blank Run (follow up to QE-1265964) | michael.knapp@sanofi.com | 8NYA 1403 | 2025-06-19 | 2025-06-19 | Column is able to pass all buffers without alarming for high pressure | Yes
5637 | 8NYA-Support Services- 3k Sum bag leak MO: RF621F183 | william.lee@sanofi.com | 8NYA Room 1350 | 2025-06-19 | 2025-06-19 | SUT expected not to leak
5636 | SUB 2220 pressure indicator reading -0.2psi | Michael.Knapp@sanofi.com | 8NYA, Room 1403 | 2025-06-19 | 2025-06-19 | We expect positive pressure on SUB X-2220 pressure indicator PI-222002 | Yes | Deviation
5635 | The aseptic G-connection was accidentally pulled off during the setup for product filtration | Pavel.dobrovolskiy@sanofi.com | 8nya 1503 | 2025-06-19 | 2025-06-19 | The SUT is expected to remain intact throughout the setup process
5634 | 8NYA - Support Service - RF604F025_002 no second verify | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-05-28 | 2025-06-19 | Hold Gemba to see path forward.
5633 | (2X) Cz SUB Bag part# 753288 failed Visual Inspection | Michael.Knapp@sanofi.com | 8NYA, Room 1403 | 2025-06-19 | 2025-06-19 | No expected the bag to pass visual inspection
5632 | CZ phenyl 3D tote bag failed visual inspection | Pavel.dobrovolskiy@sanofi.com | 8NYA 1503 | 2025-06-19 | 2025-06-19 | its expected that every bag that comes to the manufacturing floor will be free or any defects
5631 | Hi pressure on X-2560 During Blank Run 1 | Michael.Knapp@sanofi.com | 8NYA, Room 1403 | 2025-06-18 | 2025-06-20 | We expect Blank Run 1 recipe to run smooth without pressure issues.
5630 | CC-015369 76NYA Action Notice (Daily QC Sampling) PreMaturely Closed Prior to Release for Use # CA-137054 | mabc_EandM_compliance@sanofi.com | 76NYA |  | 2025-06-16 | daily QC grab samples to stop AFTER CCR for phase 1 76NYA is released for use
5629 | X2731 Equilibration buffer check timer not set properly | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-06-18 | 2025-06-18 | Buffer check timer should run for a shorter time
5628 | 8NYA-LAL Samples 1834923 and 1834924- FZ2_CA_706930_ELUATE_ENDO Batch FGHCE006 submitted to QCM in polypropylene tubes and tested by QCM. 8NYA - LAL Sample 1834898- FZ2_HARV_706930_AE_ENDO Batch FGHCE006 and 8NYA - LAL Sample 1834858 - FZ2_SUB_06930_AE_ENDO Batch FGHCE006 submitted to QCM in polypropylene tubes. | samuel.mbugua@sanofi.com | 68NYA/4700/QCM | 2025-06-17 | 2025-06-18 | Per QU-SOP-0070395 Step 6.3.3 - store samples at 2-8C in polystyrene containers. Step 6.3.4 - ensure all samples are received and diluted in polystyrene or dehydrogenated glass tubes prior to sample analysis. Per QU-SOP-0014245 - Step 6.1.2 store samples at 2-8C in polystyrene containers. Step 6.1.3 - ensure all samples are received and diluted in polystyrene or dehydrogenated glass tubes prior to sample analysis.
5627 | 51NYA: Nano Pump Failed Calibration on deltaV readout | ines.bica@sanofi.com | 51NYA | 2024-08-08 | 2025-06-18 | Nano pump to be in service | Yes | Deviation
5626 | FZ butyl SOP is not aligned with Delta V prompts | pavel.dobrovolskiy@sanofi.com | 8nya 1323 | 2025-06-18 | 2025-06-18 | its expected that SOP will align with actaul DeltaV prompts |  | No Deviation
5625 | Failed weld during column install on X2560 | Ralston.najarro@sanofi.com | 8NYA1403 | 2025-06-18 | 2025-06-18 | Welder run should not fail and the tubing should not pinch during the welding operation
5624 | Incomplete Instructions for draining Harvest and overflow bags/ SOP steps missing | MARCELO.DOSSANTOS2@SANOFI.COM | 1333 | 2025-06-18 | 2025-06-18 | Instructions for draining both bags should be detailed on QU-SOP-0083846 |  | No Deviation
5623 | 8NYA - Support Services - Stock creation needed Sum weight trend showing "no data" on Delta V | joshua.crowley@sanofi.com | 8 NYA Rm. 1350/CNC space | 2025-06-18 | 2025-06-18 | weight trend expected to show process history
5622 | GEMBA - 8 NYA DeltaV/PI Data Loss Since 09MAY2025 for 882 Tags | hiren.amin@skellig.com | 8NYA | 2025-05-09 | 2025-06-17 | DeltaV and PI Tags are expected collect historical data.
5621 | Eluate SUM Embedded Particulate | brian.o'rourke@sanofi.com | Room 1403 8 NYA | 2025-06-17 | 2025-06-17 | Per QU-SOP-0084572: If particulates are observed to be embedded, hold a GEMBA to determine if the bag is acceptable for use.
5620 | DS-Suite 1 CZ- Record needed for UFDF breakdown | katherine.reardon@sanofi.com | 8 NYA 1323 | 2025-06-17 | 2025-06-17 | Path forward to breakdown skid compliantly
5619 | Cz Mod Enzyme COA Range versus Singularity | jonathan.lucas@sanofi.com | 8NYA | 2025-06-16 | 2025-06-16 | To document weight within computer systems.
5618 | Version number is not aligned in iLearn and Veeva for MABC-QU-FOR-0043551 | jonathan.lowe@sanofi.com | 68 NYA | 2025-05-23 | 2025-06-17 | Document version numbers in iLearn and Veeva are the same
5617 | Upstream CZ X2560 SIP pressure holds results in a Pass after instant loss of pressure | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2025-06-16 | 2025-06-16 | the pressure to remain around 25PSI with minimum loss. If substantial loss results to a failed result
5616 | Cerezyme Eluates Falling Outside of Scale Weight Check Brackets | brian.o'rourke@sanofi.com | Room 1403 8 NYA | 2025-06-12 | 2025-06-12 | Drops check weight brackets may need to be adjusted for eluate transfers. |  | No Deviation
5615 | Low Temp on Zinc Eluate SUM | JOEL.HICKEY@SANOFI.COM | 8 NYA/Room: 1503. | 2025-06-16 | 2025-06-16 | Temp should be within range of 2-8C | Yes | Deviation
5614 | GEMBA - 8 NYA Suite Two - Discrepancy in Instructions for Column Flush Line Work Orders | adam.keister@sanofi.com | 8 NYA 1503 | 2025-06-15 | 2025-06-15 | For alignment on procedure - what percent ethanol and what part number to use. | No | Deviation
5613 | 68NYA Support APU-Dept QCM-Room 4700-Particle Counter FB-1660 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-4700 | 2025-05-21 | 2025-06-16 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
5612 | 8NYA, manufacturing personnel entered CNC area without high-top shoe covers | david.loik@sanofi.com | 8NYA | 2025-06-16 | 2025-06-16 | Employee would be properly gowned. | Yes | Deviation
5611 | ATF 10 Visual Inspection Failure for particulates | jonathan.ahiaba@sanofi.com | 8NYA | 2025-06-14 | 2025-06-14 | ATF To pass inspection
5610 | Conductivity result from PI Failed to meet expected passing criteria | lauren.ruffing@sanofi.com | 8 NYA | 2025-06-12 | 2025-06-15 | Conductivity result from PI will be less than or equal to 1.3
5609 | 8NYA-Gemba-Downstream Suite 2 Cz Zinc- Multiple pre-use SUT failures | aidan.mula@sanofi.com | 8NYA/RM1503 | 2025-06-14 | 2025-06-14 | Bags should meet specifications required by procedure
5608 | 8 NYA - Support Services - 692169 3K SUM Bag - particulate found pre use inspection | bailey.gleason@sanofi.com | 1350 - Solution prep | 2025-06-15 | 2025-06-15 | no particulate expected inside SUT greater than ?? mm
5607 | 8 NYA - Support Services - Agitator retaining Pins broken - found when draaining SUM | william.lee@sanofi.com | 1350 - Solution prep | 2025-06-15 | 2025-06-15 | SUM agitrator retainer pins expected not to breaak
5606 | 8 NYA - Support Services - SUM Bag leak after initial WFI | william.lee@sanofi.com | 1350 - Solution prep | 2025-06-15 | 2025-06-15 | SUM SUT expected NOT to leak
5605 | 8NYA - Support Service - FZ RF573F109 - 1000L tote bag failed pre inspection. | Bailey.gleason@sanofi.com | 8NYA CNC | 2025-06-14 | 2025-06-14 | SUT be ready to be used by MFG with no issues
5604 | 8NYA GEMBA, CZ Zinc pump tubing pressurized during load transfer | alex.tello@sanofi.com | 08-1503 | 2025-06-14 | 2025-06-14 | for line and pump not to be pressurized
5603 | 8NYA - Support Service - CZ - RF645F020_001 filter tested under incorrect parameters | Bailey.gleason@sanofi.com | 8NYA - RM 1290 | 2025-06-05 | 2025-06-14 | Filter be tested under correct formula
5602 | 8NYA - Support Service - CZ P/N 613319 Bgal mod enzyme physically here but with no load label | Bailey.gleason@sanofi.com | 8NYA USP ballroom phase 2 | 2025-06-14 | 2025-06-14 | Hold Gemba to see path forward.
5601 | CZ - RF639F014 - 8NYA - Liquid observed post-amber clamp on unused ATF spare permeate port | michael.knapp@sanofi.com | 08-1403 | 2025-06-13 | 2025-06-13 | There is no liquid after the amber clamp on that spare line.
5600 | bubble trap filter was not replaced pre skid cip | Pavel.dobrovolskiy@sanofi.com | 8nya 1323 | 2025-06-12 | 2025-06-13 | its expected that operator will change out the bubble trap vent filter prior to cip and storage of the skid
5599 | 8NYA-Gemba-Downstream Suite 2 Cz MMC no log entries for A280 equipment allocation | aidan.mula@sanofi.com | 8NYA/RM1503 | 2025-05-25 | 2025-06-13 | Allocation and deallocation of equipment should be captured in the equipment log | Yes | Deviation
5598 | 68NYA - QCM - Incorrect result entered and authorized in oneLIMS | amy.roy@sanofi.com | 68 NYA 4700 QCM | 2024-10-16 | 2025-06-12 | For the correct results to be entered into oneLIMS and confirmed upon review and authorization.
5597 | 8NYA - Support Service - Certificate of analysis bgal batch 1001-24 stating volume 40 +/- 0.5.g, exact volume may be needed | Bailey.gleason@sanofi.com | 8NYA | 2025-06-12 | 2025-06-12 | Hold Gemba to see path forward.
5596 | Particulate found in 10L wave bag | dana.hopkins@sanofi.com | 8RM1333 | 2025-06-12 | 2025-06-12 | Expected that there would be no particles in the wave bag per QU-SOP-0084572, step 6.2.13.
5595 | DS Suite 2- SUT Gemba Request- 200L Tote Bag Failed Inspection | katherine.reardon@sanofi.com | 8NYA 1503 | 2025-06-12 | 2025-06-12 | bag free from particulates
5594 | Re-GEMBA of 5587 | rudolf.boampong@sanofi.com | N/A | 2025-06-12 | 2025-06-12 | Calibration program should align with Processing criteria | Yes | 556 Event
5593 | Sample 1883093 MEDIA_POU760741_ENDOTOXIN cracked sample tube | johny.lisitu@sanofi.com | 68 NYA 4700 QCM | 2025-06-12 | 2025-06-12 | Per QU-SOP-0070558 Step 6.2.3, confirm sample containers are not cracked, leaking. Expectation is that sample tubes are not cracked.
5592 | Information Missing in Eluate Filtration Transfer | JOEL.HICKEY@SANOFI.COM | 8 NYA/Room: 1323 | 2025-06-12 | 2025-06-12 | QU-SOP-0083844 step 6.49 does not specify pump's specifications during transfer.
5591 | No Aseptic connection in place on air assembly filter for CIP. | joseph.gonzalez@sanofi.com | 1403, 8NYA. | 2025-06-12 | 2025-06-12 | Scan in new SUT for use. | Yes | Deviation
5590 | HA fermentation room 364 76 NYA lot FBYFRM022 did not meet the required viscosity of 1000 cps as required in step 7.9.7.1 of QU-OPE-0309293 | Kenneth.morgan@sanofi.com | 76 NYA room 3103 | 2025-06-11 | 2025-06-11 | viscosity must be greter than or equal to 1000 cps
5589 | HA-CMC Particulates found in V-100 during production | relja.nedeljkovic@sanofi.com | 76 NYA room 229 | 2025-06-12 | 2025-06-12 | Per QU-OPE-0520874 step 7.3.1: Confirm that V-100 has been cleaned according to QU-SOP-0042221.
5588 | Upstream phase 2 CIP tote breakdown performed not documented. Tote unable to trend drain. LvL indicator MODBAD | jeffrey.boldizar@sanofi.com | 1210 | 2025-06-12 | 2025-06-12 | Tote Drain to be documented at time of completion
5587 | 8NYA-CZ-MSAT GEMBA-All CZ Lots- Current Metrology Load Cell Bracket for SUBs do not align with Process Criteria | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2025-06-12 | 2025-06-12 | Calibration program should align with Processing criteria | No | 556 Event
5586 | 45NYA-Thyrogen APU-Room 169A-Conductivity indicator F-15836-020 FINAL RINSE CONDUCTIVITY MEASUREMENT on parts washer #1 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 45-169A | 2025-06-11 | 2025-06-12 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
5585 | Hex Glycol offload paused. Tanker moved to 76NYA and resealed. | Robert.petersjr@sanofi.com | 8NYA Tank Farm | 2025-06-11 | 2025-06-11 | The offload of Hex Glycol 805952 will not be paused, the delivery will be completed without delay.
5584 | FZ MCC 2543 leaking air from 713983(skid hook up) | jared.watson@sanofi.com | 8RM1333 | 2025-06-11 | 2025-06-11 | no leaking air
5583 | Particulates found during execution of media challenge QU-OPE-2721421 | kevin.sorge@sanofi.com | 8 NYA, Suite 2, QC Lab | 2025-04-15 | 2025-04-15 | It is expected per QU-OPE-2721421, section 6.5, that samples are negative for turbidity. |  | No Deviation
5582 | 8NYA-Cz-RF639F014- Bleed Pump not turned on per SOP | john.lytle@sanofi.com | 08-1403 | 2025-05-30 | 2025-06-11 | It is expected that the local pump switch be turned on and off according to SOP instructions
5581 | No goods receipt or goods issuance in Athena since 0900 10JUN2025 | Joshua.Brooks@sanofi.com | 8NYA MFG | 2025-06-10 | 2025-06-11 | Tibco communicates movements in MES to Athena | Yes | No Deviation
5580 | Incorrect batch record used in Sepra Labeling and packaging | alan.rousseau@sanofi.com | room 1017 74 NYA | 2025-06-09 | 2025-06-09 | Expected that the labels/pkg materials would match the batch record
5579 | CZ Media 760741 has expiry of 60 days set in Athena and in previous versions of MES PMBRs when it should be 14 days | elizabeth.sweeney@sanofi.com | 8 NYA Solution prep | 2020-06-05 | 2025-06-10 | It is expected that media 760741 have a 14 day expiry in MES and Athena as it is always stored at room temperature
5578 | 250L Eluate SUM bag failed visual inspection part#692168 Batch:23GN02370 | marcelo.dossantos2@sanofi.com | 8NYA, Room 1333 | 2025-06-10 | 2025-06-10 | We expect the SUT otb
5577 | 8NYA - Support Services - Incorrect pH ranges - Buffers 805015 & 705667 | alejandro.martinez2@sanofi.com | Pompe review session | 2025-06-10 | 2025-06-10 | pH ranges in MES should align to buffer SPC docs.
5576 | 8NYA-Gemba-Downstream Suite 2- SUT failed pre-use inspection due to embedded particulates | aidan.mula@sanofi.com | 8NYA/RM1703 (Bags 1 and 2), RM1503 (Bag 3) | 2025-06-09 | 2025-06-09 | Bags should not contain particulates
5575 | 8NYA-Gemba-Downstream Suite 2 Cz Pre-freeze hold time exceeded | aidan.mula@sanofi.com | 8NYA/RM1703 | 2025-06-09 | 2025-06-09 | The time to freeze should be clearly aligned in the CST and the EBR.
5574 | 1421 freezer going down and alarming issues | dennis.bazinet@sanofi.com | 8 NYA room 1421 (freezer) | 2025-06-09 | 2025-06-09 | Contacted Tim Biglow, facilities who works on the freezer to inspect the freezer | Yes | No Deviation
5573 | EMS alarm, 76-234 to 233, dry room door ajar | john.beaule@sanofi.com | 76-233 | 2025-06-09 | 2025-06-09 | TBS
5572 | 76NYA - QC Micro - No reportable result for EM site 425-6 RODAC B in May 2025 | craig.stevenson@sanofi.com | 76-425 | 2025-05-10 | 2025-06-09 | It is expected that all activities are documented and accounted for per QU-SOP-0019490 (V6.0) step 6.3.16
5571 | Waste bag on Nanofiltration skid unable to complete operation | timothy.ryley@sanofi.com | 8RM 1503 | 2025-06-08 | 2025-06-09 | The waste bag should have enough volume to complete equilibration
5570 | Hold x2230(train 2) LoLo media flow during automated production media change/Kinked line | bhargav.Patel@sanofi.com | 8NYARM1403 Upstream Phase 2 | 2025-06-09 | 2025-06-09 | GEMBA in the A.M
5569 | No Product, No Lots - 68 NYA QC Micro - Single Wrap TSA Lot 25GN01029 found to be unsealed along the packaging seam | amanda.lavalley@sanofi.com | 68 NYA/ QCM 4700/ cold room 4745 | 2025-06-08 | 2025-06-08 | It is expected that the packaging for these places are fully sealed and intact.
5568 | GEMBA: Downstream Suite 2 cerezyme: Bag particulate | William.Bradley@sanofi.com | 8NYA 1703 | 2025-06-08 | 2025-06-08 | It is expected for the SUT to be in good standing condition without deformities or particulates. | Yes | Deviation
5567 | Media tote had missing clamps prior to media install ¦ FZ 8NYA RM1210 | christian.galvez-tello@sanofi.com | 8NYA RM1210 CNC | 2025-06-08 | 2025-06-08 | 6.17 of MABC-QU-WIN-0019199-Single Use Containers (v28.0) | Yes | Deviation
5566 | GEMBA: Suite 2 Cerezyme: Need to Align with Nano operation | William.Bradley@sanofi.com | 8NYA 1503 | 2025-06-08 | 2025-06-08 | There's a lot of issues to align and I can't figure out on my own. We have limited resources and I don't want to risk the product anymore than it has gone through this weekend.
5565 | 8NYA Downstream Phase 2 - Cerezyme FGHDS028 - Expired Sampling Manifold Installed, but Not Used on Post-Viral SUM | wilfredo.lopez@sanofi.com | 08RM1703 | 2025-06-08 | 2025-06-08 | It's expected that the SUT is not expired when used.
5564 | 8NYA Downstream Phase 2 - Cerezyme FGHDS028 - Particulate found during Nano Equilibration Setup | wilfredo.lopez@sanofi.com | 8RM1503 | 2025-06-07 | 2025-06-07 | It's expected that SUT is free of foreign particles.
5563 | While doing the Cerezyme Zinc pH adjustment, we ran out of the 705661 buffer | KELVIN.GLEAN@SANOFI.COM | 8 NYA room1503 | 2025-06-07 | 2025-06-07 | That the pH probe is in the correct position to read solution pH during adjustment operations
5562 | SUT failed Inspection | elaine.jansen@sanofi.com | 8NYA 1333 | 2025-06-07 | 2025-06-07 | MABC-QU-SOP-0084572-Single Use Containers Step 6.2.14 Inspect the smooth surface of the entire bag for unevenness, damage, and abnormality.The presence of small folds and shallow folds are permitted. Visible tears, holes or cuts are never permitted. | Yes | Deviation
5561 | 8NYA-Cz- FGHBN013- ATF10 Failure | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2025-06-07 | 2025-06-07 | It is expected that the ATF10 system is functional
5560 | Downstream gemba request- damaged nanofiltration temp monitor (772680) | kreardon13@gmail.com | 8NYA 1503 | 2025-06-07 | 2025-06-07 | sut in tact
5559 | GEMBA: Cerezyme Suite 2 DS: Alarm on Delta V when starting the Formulation Equilibration | William.Bradley@sanofi.com | 8NYA /RM1503 | 2025-06-07 | 2025-06-07 | System should run as it is designed. | Yes | No Deviation
5558 | GEMBA: Downstream Suite 2 Cerezyme: Omsmo sample high out of spec | William.Bradley@sanofi.com | 8NYA /RM1503 | 2025-06-07 | 2025-06-07 | Sample is to be within range
5557 | GEMBA: Downstream Suite 2 Cerezyme: Overlay not acquired from CM Filtration Setup | William.Bradley@sanofi.com | 8NYA /RM1503 | 2025-06-06 | 2025-06-07 | Equipment Module is to be set to "overlay"
5556 | Pump Tubing Leak During Anti-Foam Prime X-2230 | brian.o'rourke@sanofi.com | 8NYA Room 1403 | 2025-06-07 | 2025-06-07 | Antifoam should prime the line leading to the SUB via DeltaV prompted sequence
5555 | 8NYA Gemba Support Services 760741 SUT discarded after setup | william.lee@sanofi.com | 8NYA/1350 | 2025-06-06 | 2025-06-06 | SUT should have still been setup for batch to occur. | No | Deviation
5554 | Harvest overflow pump not starting | Jared.Watson@Sanofi.com | 8RM1333 | 2025-06-06 | 2025-06-06 | the pump turn on
5553 | Gemba 8NYA Solution Prep - Leak discovered during 705691 transfer | william.lee@sanofi.com | 8NYA/1350 | 2025-06-06 | 2025-06-06 | There should not have been any leaks and 1k tote should have transfered with no problems. | Yes | Deviation
5552 | HA - FBYMFL002 - 76NYA Room 355 - QA Observed that operator did not wait for 5 minutes to start filling after bottle breakage per QU-OPE-0345655 | Yamini.patcha@sanofi.com | 76 NYA, Room# 355 | 2025-05-22 | 2025-06-06 | Per QU-OPE-0345655 Attachment 2, after the components in the torques area of the fill table have been changes out with the new sterile components. Operator should wait for five minutes settling period to be completed prior to starting the filling activities.
5551 | HA-CMC- Multiple Lots - 76 NYA - Vessel V-200 Failed TOC Reading During Cleaning Requalification | Phillip.Mank@Sanofi.com | 76 NYA Room 229 | 2025-06-06 | 2025-06-06 | It is expected that the V-200 requalification TOC reading be below the Action level of >0.5ppm
5550 | Upstream CZ agitator shaft residue discovered in SUB breakdown | jeffrey.boldizar@sanofi.com | 8NYA1403 | 2025-06-06 | 2025-06-06 | minor amounts of residue |  | No Deviation
5549 | Cz Phenyl- 200L Barrel Bag failed inspection | nathaniel.barnes@sanofi.com | 8 NYA 1503 | 2025-06-05 | 2025-06-05 | Bag seams to not be breeched
5548 | 51NYA: Cold Room 132 - Temperature QCA | ines.bica@sanofi.com | 51NYA | 2025-06-05 | 2025-06-05 | Temperature was above limit between 1522 - 1548 reaching the highest temperate of 10.17
5547 | pH EUR misaligned with Cz and Fz MCC record | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-06-01 | 2025-06-01 | Record would advance as normal without looking for parent Cx values to populate
5546 | BOM Discrepancy | Thomas.Loik@sanofi.com | 8YNA | 2025-06-05 | 2025-05-05 | There are no discrepancies between MES and Athena
5545 | WFI spill from SUT while filling a Media tote with WFI for calibration | dennis.bazinet@sanofi.com | 8NYA CNC / 1210 | 2025-06-04 | 2025-06-04 | 6.18.10 connect the air line
5544 | 8NYA - Support Service - CZ - HiHigh DP alarm for RM 1350 during FAC WO 3349968 | bailey.gleason@sanofi.com | 8NYA - RM 1350 | 2025-06-04 | 2025-06-04 | N/A
5543 | Cryopod allows vial transport per MES for 4 hours while it was qualified for only 3 | elon.elzam@sanofi.com | 68NYA | 2025-05-14 | 2025-05-30 | OQ - F008-LN2-102-8NYA-CQ-0001(1.0)Post-Approved, the OQ document qualifies the cryopod transport of a vial for only 3 hours | No | 556 Event
5542 | 691058 (FILTERS ASSEMBLY,(0.2UM)PRODUCT 10"), 754042 (XCELL ATF SINGLE USE DEVICE 0.2M PES), 690234 (FILTERS ASSEMBLY,(.5/0.1UM)30"), 692327(UTILITY WELDABLE, 10ML SYRINGE SAMPLER) , 777971 (TUBING ASSEMBLY, 30" FILTER LINE) were received on 03 Jun25 and 04Jun25 from 11 Forbes with damaged boxes | michael.montiverdi@sanofi.com | 8 NYA JIT | 2025-06-03 | 2025-06-03 | Boxes from daily delivery are integral upon transit and delivery | Yes | Deviation
5541 | 8NYA- No Product- Gaps identified in MCC probe operations-SOPs/ MES misaligned | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1333, 1403 | 2025-06-02 | 2025-06-03 | Continuous improvement opportunities/ GDP compliance | Yes | Deviation
5540 | 8NYA-CZ-RF639F013/ FGHBN013 X2811 MCC hold for high pressure on column 1 | john.lytle@sanofi.com | 8NYA Rm 1403 | 2025-06-03 | 2025-06-03 | expected to remain within allowable/ alarm limits for pressure while forward processing | Yes | Deviation
5539 | Autoclave/Parts Washer Service Room Drain Overflow | jeff.hill@sanofi.com | 45-169 and 45-178 | 2025-06-03 | 2025-06-03 | n/a | Yes
5538 | Autoclave/Parts Washer Service Room Drain Overflow | jeff.hill@sanofi.com | 45-169 and 45-178 | 2025-06-03 | 2025-06-03 | n/a
5537 | GEMBA - 8 NYA DSP Suite Two - Need Path Forward for Storing CM Skid | adam.keister@sanofi.com | 8 NYA Room 1503 | 2025-06-03 | 2025-06-03 | Path forward with CM skid. Can it remain empty until equil is run tomorrow? If it needs to be stored, do we abort the setup recipe and manually launch a CIP recipe?
5536 | PH Probe 2 Disabled Per MSAT x2230, Aber probes diverging impacting Top Off | John.Lytle@sanofi.com | 8 NYA Phase 2 Upstream RM 1403 | 2025-06-03 | 2025-06-03 | Probe Disabled Per MST
5535 | expired alcohol foam cans found expired at 8NYA room 1240 | Adeline.martin@sanofi.com | 8NYA room 1240 | 2025-06-02 | 2025-06-02 | Alcohol foam cans should not be expired
5534 | NRSR Paperwork for Cz lot FGHDT020 Not Retained | nathan.kallen@sanofi.com | 8 NYA - Downstream Suite 2 Room1503 | 2025-05-22 | 2025-05-30 | It is expected the original NRSR paperwork is retained and archived.
5533 | 8NYA GEMBA - Support Services - IMPORTANT MOD ENZYME | jack.laplante@sanofi.com | 8NYA 1350 | 2025-06-02 | 2025-06-02 | Looking for path foward. | Yes | Deviation
5532 | Version 5.0 of QU-SOP-0034242 approved multiple discrepancies | Amy.McIntyre@sanofi.com | QualiPSO CMS | 2024-08-26 | 2025-06-02 | Expected results would be that the workflow owner added all required signatories, added required TIA statement, updated versions correctly in revision history table to align with versioning in QualiPSO, and updated margins correctly in order to not have overlays impede content, per steps 6.2.3, 6.2.5, and 6.2.6 of QU-WIN-0055911
5531 | GEMBA 8NYA Suite 1 Upstream SUB 2200 Train 1 Lot FGHCE006 Antifoam Pump Tubing Off Pump Roller | GARY.DODAKIAN3@SANOFI.COM | 8 nya | 2025-06-03 | 2025-06-02 | Antifoam doses when the sub automation sends the signal to the pump
5530 | GEMBA 8NYA Suite 1 Upstream SUB 2200 Train 1 Lot FGHCE006 | GARY.DODAKIAN3@SANOFI.COM | 8 nya | 2025-06-03 | 2025-06-02 | Antifoam doses when the sub automation sends the signal to the pump
5529 | Upstream FZ BGA values discussions | jeffrey.boldizar@sanofi.com | 8RM 1333 | 2025-06-02 | 2025-06-02 | BGA values to be at an acceptable range for SUB adjustments
5528 | 51NYA-Thyrogen-No lots- Sterile tube welder PM extended overdue | adriana.alban@sanofi.com | 51NYA | 2025-05-22 | 2025-06-02 | It is expected to complete an equipment PM when due | Yes | Deviation
5527 | 68NYA-Support APU-Room 1830-Metrology Humidity Standard FBS-1568 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2025-05-22 | 2025-06-02 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5526 | GEMBA: 8NYA: Downstream : Suite 2 (RM1503) 100L Bag pre-use failure | William.Bradley@sanofi.com | 8NYA /RM1503 | 2025-06-01 | 2025-06-02 | It is expected that bags are in good integrity with no particulates
5525 | Samples Found Outside of Temp | JOEL.HICKEY@SANOFI.COM | 8 NYA/Room: 1503. | 2025-06-01 | 2025-06-02 | Per QU-OPE-0044689 table 13, pg. 26 temperature much be between 2-10C
5524 | Flooring materials found in production areas at 51 NYA | jesse.demers@sanofi.com | 51 NYA multiple rooms | 2025-06-01 | 2025-06-01 | There should be no loose flooring material in GMP production areas
5523 | 8NYA - Support services -3 k tote bag 692174 - particulate inside bag discovered at pre use inspection | william.lee@sanofi.com | 1202 CNC East Corridor | 2025-05-31 | 2025-05-31 | oversized particulate not expected
5522 | 250L SUM Bag Failed Visual Inspection; No batch associated | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-05-31 | 2025-05-31 | Bag is free of damage/abnormalities
5521 | 8NYA Gemba -Support Services - 3K bag leak | maria.palmar@sanofi.com | 8NYA / 1350. | 2025-05-31 | 2025-05-31 | 3K bag to not leak
5520 | CZ - RF639F012 - 8NYA - DO Probe divergence on SUB X2240 | michael.knapp@sanofi.com | 8NYA 1403 | 2025-05-30 | 2025-05-30 | Accurate probe reading for DO.
5519 | Executed hypothesis testing protocol cannot be located | nathan.kallen@sanofi.com | 8 NYA - Downstream Suite 2 Room1503 | 2025-05-22 | 2025-05-30 | It is expected the executed protocol is retained and archived.
5518 | MES takeout discrepancies | Joshua.brooks@sanofi.com | 8NYA | 2025-05-30 | 2025-05-30 | When a storage unit is scanned, the correct amount is being entered.
5517 | HA -FBYHAH002 76 NYA room 3103 HA transferred to SWECO without closing product plug | Kenneth.morgan@sanofi.com | 76 NYA room 3103 | 2025-05-30 | 2025-05-30 | It is expected that the product plug is closed and no product is in the collection can prior to drying completion
5516 | 8NYA - SUT Visual Inspection Failure - ATF 10 Visual Inspection Failure for particulates and improper sealing along bottom of ATF | michael.knapp@sanofi.com | 8NYA 1403 | 2025-05-26 | 2025-05-26 | Passing ATF 10 for use
5515 | 8NYA- Support Services Expired conductivity standard used for standardization | joshua.crowley@sanofi.com | 8 NYA RM1350 | 2025-05-28 | 2025-05-29 | MABC-QU-SOP-0084641-Operation of pH and Conductivity Meters at 8NYA, step 6.10.10
5514 | 50L wave bag failed visual inspection | Jacob.Roman@sanofi.com | 8RM1403 | 2025-05-29 | 2025-05-29 | Expected that no particles are found inside single use bags.
5513 | 8NYA-Fz-FGHCE004- Antifoam EM did not switch to MED dose rate on Day 12 | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2025-05-28 | 2025-05-29 | it is expected that Antifoam dosing rate increase to MED on Day 12
5512 | 8NYA Gemba - Support Services - Operations performed while not trained on current SOP version | christopher.arnold@sanofi.com | 8RM1364, 8RM1350 | 2025-04-20 | 2025-05-29 | That all employees are trained on the most recent SOP version before performing tasks. | Yes | Deviation
5511 | Flooring material found in 51 in suite air returns | john.mendenhall@sanofi.com | 51 NYA | 2025-03-10 | 2025-03-10 | N/A
5510 | 51NYA: WO 3283359: SELF INSPECTION Walk through checklist not fully completed | INES.BICA@SANOFI.COM | 51 NYA | 2025-02-22 | 2025-05-28 | Full checklist of the WALK-THROUGH to be completed
5509 | upstream FZ buffer prime waste line connection, Additional 5 foot accidently connected to waste line | jeffrey.boldizar@sanofi.com | 8NYA 1333 | 2025-05-28 | 2025-05-28 | The wye to connect the to the 30ft
5508 | X2230 SUB Integrity Test Held for EM-X2230-AIR-FAILURE | brian.o'rourke@sanofi.com | 8NYA Room 1403 | 2025-05-28 | 2025-05-28 | Recipe should launch from the MBR and begin inflation of the SUB bag
5507 | Upstream FZ MCC record missing CX value placement for MO | jeffrey.boldizar@sanofi.com | 8NYA 1333 | 2025-05-28 | 2025-05-28 | the CX value has a position to receive the command in MES
5506 | 8NYA GEMBA - Downstream Suite 1 - Path Forward Needed for Closing Out UFDF Membrane Record | danny.martinez@sanofi.com | 8NYA RM1323 | 2025-05-28 | 2025-05-28 | Get a path forward
5505 | Ceiling Breach 45 NYA 190 | jeff.hill@sanofi.com | 45 NYA RM 190 | 2025-05-28 | 2025-05-28 | QU-SOP-00773883, Step 6.1.6
5504 | Ineffective CAPA-052780 | mark.gwozdzik@sanofi.com | 8NYA | 2025-05-23 | 2025-05-23 | GQCMS-QU-OPE-2513742 to be referenced in GQCMS-QU-GOP-0000514 and appropriate users trained to GQCMS-QU-GOP-0000514.
5503 | Failed SUT Gemba | ryan.devanna@sanofi.com | 8RM1403 | 2025-05-27 | 2025-05-27 | Expected SUT to pass
5502 | Phath forward for UF\DF membrane | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1323 | 2025-05-27 | 2025-05-27 | To get a Phath forward
5501 | Incorrect step identified in EWI during change over to SOP | phoebe.karkos@sanofi.com | 1333, 1403 | 2025-05-27 | 2025-05-27 | need to evaluate impact of incorrect parameters for Centrifuge
5500 | 51-NYA: Use of Obsolete Logbook for 106 Cold Room | INES.BICA@SANOFI.COM | 51 NYA | 2025-04-28 | 2025-05-26 | Logbook and SOP to be revised at the same time and operators not using obsolete logbook
5499 | TSH - 45NYA- Insect found in gowning airlock 177A | cindy.guerra@sanofi.com | 45NYA 177A | 2025-05-27 | 2025-05-27 | It is expected MFG area would not have any insects present
5498 | 8NYA Gemba -Support Services - MO not able to close | mariapalmar1810@gmail.com | 8NYA / 1350. | 2025-05-27 | 2025-05-27 | Record to be close with no integrity issues.
5497 | Personnel observed entering the clean room Grade 4 (room1404) without the proper gowning | Yves.Antoine@sanofi.com | 8NYA Room 1403 | 2025-05-27 | 2025-05-27 | We expect everyone entering the clean room to gown in properly
5496 | 8NYA-Gemba-Downstream-Phase 2- Mod Enzyme bag had small leak observed after transfer completed | aidan.mula@sanofi.com | 8NYA/RM1503 | 2025-05-26 | 2025-05-26 | SUT should remain integral
5495 | SUB 2200 Probe Co2 Divergence | Michael.Knapp@sanofi.com | 8 NYA Phase 1 Upstream RM 1333 | 2025-05-26 | 2025-05-26 | N/A
5494 | Fz SUB X2210 DO Control Not Stabilizing | jonathan.lucas@sanofi.com | 8NYA | 2025-05-24 | 2025-05-25 | DO to stabilize within a range of 40%-60% and High O2 MFC take sole control.
5493 | Low concentration on MMC Eluate | charles.salgadinho@sanofi.com | 8NYA | 2025-05-25 | 2025-05-25 | the mmc recovery percentage would be in range
5492 | 10L wave bag failed due to particulate found during bag inspection. | joseph.gonzalez@sanofi.com | 1403 | 2025-05-24 | 2025-05-24 | Bag set aside, new bag ordered | Yes | Deviation
5491 | FZ upstream Skid X2543 failed visual inspection for CV during CIP | jeffrey.boldizar@sanofi.com | 8NYA 1333 | 2025-05-24 | 2025-05-24 | the Bubble trap to be free/ clean of caustic and to pass visual inspection
5490 | 8NYA Gemba -Support Services -Standard in EBR does not match with standard in SOP for 760721 | julie.auguste@sanofi.com | Downstream, phase 2, 1503 | 2025-05-24 | 2025-05-24 | The EBR needs to reflect MABC-QU-SOP-0084642 and be set to the correct standard.
5489 | Rust on BSC in 51-152B (Inoculation Room) | kimberly.rice@sanofi.com | 51-152B | 2025-05-23 | 2025-05-23 | No rust on equipment in suite
5488 | 8 NYA - Support Services - Verifiable source of information not noted for manual entry - Identified during Internal Audit | nicholas.grote@sanofi.com | 08-1202 | 2025-04-19 | 2025-05-22 | Operator sites a verifiable source that the information has come from.
5487 | Upstream, CZ, SUB X2220 DO probes discussion | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2025-05-23 | 2025-05-23 | DO probes to be closer in value | Yes | No Deviation
5486 | 8NYA - No Product - Training - OJT pre-requisite and Curriculum section filled out two days after check list complete. | michael.knapp@sanofi.com | 8NYA | 2025-05-18 | 2025-05-22 | Training is documented in the correct order. verification of prerequisities will be completed prior to training being performed.
5485 | Post viral Shipping bag, embedded particulate was found | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1703 | 2025-05-22 | 2025-05-22 | To get a Phath forward
5484 | Cleaning Logbooks for 74NYA and 76 NYA from 2022 have not been reconcile by cleaners and QA OPS | Michael.Corrigan@sanofi.com | 68NYA ROOM 1825 | 2025-05-22 | 2025-05-22 | The Logbooks that needed QA review would be processed and that the cleaning staff would process logbooks and give to RM to File | No | Deviation
5483 | Fz X2200 SUB media charge pumped Base instead of Media | Kurt.marsocci@sanofi.com | 8NYA1333 | 2025-05-22 | 2025-05-22 | MABC-QU-SOP-0083846 6.6.4
5482 | QE-176526: HACMC OOS for FBYCMC055 | Aleen.Merzoian@sanofi.com | 76NYA-107 | 2025-05-22 | 2025-05-22 | 22-45% CMC
5481 | 8NYA-Gemba-Downstream Phase 2 Cz Nano- Formulation concentration high out of range | aidan.mula@sanofi.com | 8NYA/RM1503 | 2025-05-22 | 2025-05-22 | Sample should be within expected range
5480 | Routine HA Fill EM - 76 NYA MFG - zero filter check not re performed between sites and static or dynamic not selected on paperwork for all sites. | megan.botti-marino@sanofi.com | 76 NYA | 2025-05-09 | 2025-05-22 | Zero filter was not re-performed per step 6.4.1 of MABC-QU-SOP-0074238-Use of the Climet CI-450t (v12.0)
5479 | 76NYA and 51NYA-Clean Facilities-4 TOC Analyzers found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76NYA room 347, 76NYA room 140, 51NYA room 146, 76NYA room 347 | 2025-05-19 | 2025-05-22 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5478 | Seprafilm - 76 NYA - Non-conformance of Steris Gamma radiation dosage | karthik.sriram@sanofi.com | N/A | 2025-05-14 | 2025-05-16 | There is misalignment in the specification provided in the OCS (25 - 40 kGy) and batch record steps (step 7.17.1.3 and 7.17.1.4) ≥ 25.0 kGy and ≤ 40.0 kGy
5477 | During Internal Audit Review on 20May25, 2 WO's for Self-Inspection had individuals execute without completion of MABC-QU-SOP-0075222-Walk-Through Procedure for MA BioCampus | nicholas.candela@sanofi.com | 8 NYA | 2024-01-21 | 2025-05-20 | Individuals would have been trained on MABC-QU-SOP-0075222-Walk-Through Procedure for MA BioCampus
5476 | 8NYA - Support Service - Zumerror discovered incorrect allocation of HU 1001413001 | bailey.gleason@sanofi.com | 8NYA | 2025-05-21 | 2025-05-21 | N/A | Yes | Deviation
5475 | 8NYA Badge audit work order missing information | sarah.francis@sanofi.com | 8 NYA | 2024-06-20 | 2025-05-20 | All areas reviewed are uploaded into the work order and that attachments contain references to the work order and are paginated.
5474 | Pre use FIT's being performed while Compressed Air AN 2025-F-023 is in place | jesse.demers@sanofi.com | 51RM143 | 2025-05-10 | 2025-05-20 | No Pre Use FIT's should have been performed.
5473 | 8NYA - Support Service - FZ Autoclave Probe recipe aborted for max time under temp executed | bailey.gleason@sanofi.com | 8NYA 1373 | 2025-05-21 | 2025-05-21 | Recipe run through completion
5472 | ATF Sampling/ Phase 1&2 BGA inoperable | bhargav.Patel@sanofi.com | 8 NYA Phase 2/1. RM 1403 and 1333 | 2025-05-20 | 2025-05-20 | Contact Vendor to troubleshoot and repair.
5471 | Lack of instructions for swapping out spent buffer totes | edgar.vale@sanofi.com | 8NYA/Room1503 | 2025-05-20 | 2025-05-20 | Sufficient buffer would be available to complete wash phase and SOP would contain instructions for swapping out empty tote | Yes | Deviation
5470 | 51NYA-Thyrogen-no products-Sterile Tube Welder unable to be located | adriana.alban@sanofi.com | 51NYA | 2025-03-05 | 2025-05-20 | It is expected to receive equipment from a third party vendor
5469 | pH Meter 2009946 Failed Standardization Twice. Placed Out of Service | brian.o'rourke@sanofi.com | 8 NYA Room 1403 | 2025-05-19 | 2025-05-19 | pH should fall into expected range of 6.95 - 7.05 per MABC-QU-SOP-0084641 step 6.4.22
5468 | Sartorius Mixer 2012474/FB-4291 has multiple alarms running. They read: "Mixing Motor Error: Controller type" and "Coupling Warning: Impeller". Asset needs to be sent to vendor for repairs. | paul.mclaughlinjr@sanofi.com | 8NYA Rm 1503 | 2025-05-09 | 2025-05-09 | Mixer would not be displaying error codes | Yes | No Deviation
5467 | Lots assigned to investigation or experiment used in commercial production. | timothy.ryley@sanofi.com | 8RM 1403 | 2025-03-27 | 2025-05-19 | X Lots of materials should not be used in commercial production.
5466 | HA LOT FBYFRM016 76 NYA room 364 NaOH pump 76-pump10 stopped working during overnight pH feed into lot FBYFRM016. | Kenneth.morgan@sanofi.com | 76 NYA room 364 | 2025-05-20 | 2025-05-20 | It is expected that the pH remains at the setpoint of 7.73 +/- 0.3
5465 | 8NYA Downstream - A280 Test Method File Location Unclear in SOP | wilfredo.lopez@sanofi.com | 8RM1503 | 2025-05-19 | 2025-05-19 | That the SOP contains clear and detailed instructions on the specific location of the test method files to be selected in the process.
5464 | Equil Timer expired on Skids 2811/2560/2831 Same issue as QE-165668 | bhargav.Patel@sanofi.com | 8 NYA Phase 2 upstream RM 1403. | 2025-05-19 | 2025-05-19 | N/A
5463 | Nano hold bag failed pre-use visual inspection | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2025-05-20 | 2025-05-20 | No hair/microfibers, embedded particulate, debris, nor surface scratches are permitted for bags | Yes | Deviation
5462 | 8NYA-DOWNSTREAM-CZ- MMC Pre-use bag inspection failure | katherine.reardon@sanofi.com | 8 NYA 15303 | 2025-05-19 | 2025-05-19 | Per QU-SOP-0084572 step 6.2.13, bags should be free from particulates
5461 | Skid 2560 Equilibration buffer recipe check pH value 60 minutes vs 60 sec | bhargav.patel@sanofi.com | 8 NYA / 1403 | 2025-05-16 | 2025-05-16 | recipe goes in hold quicker if pH criteria does not match
5460 | 8NYA Gemba Support services 760741 3K SUM bag did not pass inspection | william.lee@sanofi.com | 8NYA/1350 | 2025-05-19 | 2025-05-19 | 3K SUM bag should have the correct pressure line configuration
5459 | Vesi virus retain missing | adam.rowe@sanofi.com | 51NYA, 108C (-80 freezer) | 2025-05-15 | 2025-05-15 | All retains should be held until scheduled disposal
5458 | Capacitance probe AI2 diverged towards an alarm limit and triggered an alarm | brendan.dickhaut@sanofi.com | 8NYA | 2025-05-19 | 2025-05-19 | Probes operate as expected throughout entire campaign without failing.
5457 | X-2453 CIP needs CV post PM. Status and Manufacturing Order are calling for non-CV | dana.hopkins@sanofi.com | 8NYA Upstream Suite 1 room 1333 | 2025-05-18 | 2025-05-18 | MCC skid status should be correct for CIP |  | No Deviation
5456 | 8NYA - Support Service - CZ RF618F019 - Automatic steps not pulling for mod enzyme record | bailey.gleason@sanofi.com | 1350 | 2025-05-18 | 2025-05-18 | N/A
5455 | Differential Pressure QCAs for room 51-87 to 51-88 during TSH unformulated bulk processing | Craig.valery@sanofi.com | 51-87 to 51-88 | 2025-05-17 | 2025-05-18 | there are no QCA occurrences during processing
5454 | 8NYA - MFG ALL - Freezer 1421 alarming for HI HI temp | bailey.gleason@sanofi.com | 8NYA - RM1421 | 2025-05-17 | 2025-05-17 | Hold Gemba to see path forward.
5453 | 8NYA - Support Service - CZ RF621F154 - 3000L SUM bag failed pre inspection | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-05-17 | 2025-05-17 | SUT be ready to be used by MFG with no issues
5452 | 8 NYA Support Services - 760744 Ppost filtration Osmo result out of range | leo.gillespie@sanofi.com | 8 NYA Room 1350 (Solution Prep) | 2025-05-16 | 2025-05-17 | In range result was expected - prefiltration result was 277
5451 | Gapless amber clamp use during ATF change out for SUB 2240 | bhargav.patel@sanofi.com | 8 NYA / 1403 | 2025-05-16 | 2025-05-16 | Amber clamp without gap should not be use due to CCR and PQ has not been completed
5450 | pH Offest Miss during Daily Sampling X-2230 Cerezyme | mike.niland@sanofi.com | 8 NYA | 2025-05-10 | 2025-05-11 | Required pH probe offsets must be applied for Cerezyme bioreactors during their standardization per MABC-QU-SOP-0084571.Skid In-Line Analytical Instruments
5449 | CZ - RF639F008 and RF639F011 - 8NYA - Potential impact to media supply, awaiting new media batch | michael.knapp@sanofi.com | 8NYA 1403 | 2025-05-16 | 2025-05-16 | Media is prepared and ready for consumption
5448 | GEMBA 8NYA Upstream Suite 1 TCU 2213 for SUB X-2210 going into alarm for Pressure Alarm | GARY.DODAKIAN3@SANOFI.COM | 8 NYA room 1333 suite 1 upstream / penthouse | 2025-05-15 | 2025-05-15 | TCU should not be in alarm for pressure
5447 | 8 NYA Drain Cleaning | andrew.doheny@sanofi.com | 8 NYA | 2025-05-14 | 2025-05-14 | Step 6.12 for drain cleaning in the SOP procedure
5446 | DP alarm (51-82 CNC corridor to 51-127 MAL) found during QCA review without comment | nicholas.shoenfelt@sanofi.com | 51NYA | 2025-03-08 | 2025-05-14 | It is expected that DP is maintained between grade D and CNC areas.
5445 | 8NYA GEMBA - Support services - Failed 3k SUM bags due to particles inside bag | JACK.LAPLANTE@SANOFI.COM | 8NYA RM1350 | 2025-05-14 | 2025-05-14 | Path forward | Yes | Deviation
5444 | Embedded particulate in ATF6 tubing | elaine.jansen@sanofi.com | 8NYA 1333 | 2025-05-15 | 2025-05-15 | MABC-QU-SOP-0084522-Single Use Tubing (v5.0) section 6.2.12
5443 | Ineffective PM/Not being performed properly. ~2,300 personnel need access revoked. Process of assigning access needs improvement. | paul.mclaughlinjr@sanofi.com | Campus Wide | 2025-05-14 | 2025-05-14 | PM Code 3178393 would have been executed properly therby removing unecessary access
5442 | Drain Cleaning not performed in accordance with procedure | andrew.doheny@sanofi.com | 8 NYA | 2025-05-14 | 2025-05-14 | Step 6.12 for drain cleaning in the SOP procedure
5441 | Eluate SUM bag failed inspection | brett.simmons@sanofi.com | 8NYA, RM1403 | 2025-05-15 | 2025-05-15 | SUM bag/SUT are free of particulates. MABC-QU-SOP-0084572 v4.0 step 6.2.13
5440 | 8NYA GEMBA - Downstream PV filter failed FIT twice | JACK.LAPLANTE@SANOFI.COM | 8nya | 2025-05-14 | 2025-05-14 | Path forward
5439 | Cz-FGHBN008- X2230 DO Probe Divergence | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2025-05-14 | 2025-05-14 | When probes are approaching or at alarm limits, It is expected that a GEMBA is held and probe performance discussed to decide if disabling of probes are necessary
5438 | MES Scanners not qualified prior to use | paul.mclaughlinjr@sanofi.com | 8NYA | 2025-05-09 | 2025-05-09 | All assets involved in MFG should be qualified
5437 | 8NYA GEMBA IMPORTANT - Support services - Probe load aborted due to low temperature | JACK.LAPLANTE@SANOFI.COM | 8nya 8rm1373 | 2025-05-14 | 2025-05-14 | Path forward
5436 | 250L Eluate SUM bag failed visual inspection part#692168 Batch:23GN02369 | Yves.Antoine@sanofi.com | 8NYA, Room 1403 | 2025-05-13 | 2025-05-13 | Expect the bag to be free of particulates
5435 | Probe bellow missing ziptie, white port showing different sizes and unusual port placement. | kurt.marsocci@sanofi.com | 8 NYA room 1333 | 2025-05-13 | 2025-05-13 | Replace questionable bellows with new ones.
5434 | High quantity of QCA callouts following maintenance work. Not under action notice. Lack of control. | jesse.demers@sanofi.com | 45 NYA | 2025-05-13 | 2025-05-13 | *
5433 | CZ - RF639F008 - 8NYA - Increased pO2 identified during offline daily sampling | michael.knapp@sanofi.com | 8NYA 1403 | 2025-05-13 | 2025-05-13 | No O2 increase during campaign
5432 | 8NYA GEMBA - Downstream Phase 2, SUM allocated to Freeze record. | alex.tello@sanofi.com | 8NYA RM1703 | 2025-05-13 | 2025-05-13 | We will need to deallocate the SUM manually
5431 | 8NYA- Support service - CZ RF606F006/ 760735 - 3000L SUM bag with 2 male connections | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-05-13 | 2025-05-13 | replace bag, hold gemba for defect
5430 | CZ - RF639F013 - 8NYA - DO Probe divergence requiring MSAT trend review and potential disabling | michael.knapp@sanofi.com | 8NYA 1403 | 2025-05-13 | 2025-05-13 | Probes are representative of the culture and do not display evidence of fouling.
5429 | Gemba-8NYA-Downstream Suite 2-Nano post form concentration too high | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 room 1503 | 2025-05-12 | 2025-05-12 | NA |  | Deviation
5428 | 250L Eluate SUM bag failed visual inspection part#692168 Batch:23GN02369 | Yves.Antoine@sanofi.com | 8NYA, Room 1403 | 2025-05-12 | 2025-05-12 | Expect the bag to be free of particulates and abnormality in the bag film
5427 | Gemba 8NYA Upstream Suite 1 2 x Failed SUB Bag Visual Inspection 689204 | GARY.DODAKIAN3@SANOFI.COM | 8 NYA room 1333 suite 1 upstream | 2025-05-09 | 2025-05-09 | SUB bag should be deemed acceptable for use from the visual inspection
5426 | Gemba 8NYA Upstream Suite 1 2 x Failed SUB Bag Visual Inspection 689204 | GARY.DODAKIAN3@SANOFI.COM | 8 NYA room 1333 suite 1 upstream | 2025-05-09 | 2025-05-09 | SUB bag should be deemed acceptable for use from the visual inspection
5425 | Gemba 8NYA Upstream Suite 1 2 x Failed SUB Bag Visual Inspection 689204 | GARY.DODAKIAN3@SANOFI.COM | 8 NYA room 1333 suite 1 upstream | 2025-05-09 | 2025-05-09 | SUB bag should be deemed acceptable for use from the visual inspection
5424 | 8NYA - Support Services - Failed SUM bag inspection - Two female connections in both bag and pressure line | julie.auguste@sanofi.com | Solution prep, 1350 | 2025-05-12 | 2025-05-12 | Per QU-SOP-0084572 step 6.6.27 the pressure sensing line needs to connect to pressure sensing port in bag. That connection is rendered impossible due to having the same quick connect port on both ends.
5423 | SUM bag failed inspection | elaine.jansen@sanofi.com | 1403 8NYA | 2025-05-11 | 2025-05-11 | Per MABC-QU-SOP-0084572 6.2.13 particulates within the product contact surface require GEMBA
5422 | 8NYA-Cz-Automation intervention of the Cz SUBs (x4) | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2025-05-11 | 2025-05-11 | It is expected that media changeout occur when supply tote levels drop below SOP criteria (50L) | Yes | No Deviation
5421 | 8NYA All - Fire alarm. Second event | wilfredo.lopez@sanofi.com | 8NYA | 2025-05-10 | 2025-05-10 | Expected that clean space integrity is maintained. Gemba to determine any impact to ongoing CM Chromatography operations. Area cleanings coordination required. Path forward for resuming other ongoing operations. | Yes | Deviation
5420 | Inadequate control of nKPP Pompe ICB Vial Thaw Duration within MES | Timothy.houck@sanofi.com | 8NYA Upstream | 2023 | 2025-05-07 | QU-OPE-0522556 Table 8-3 Vial Thaw duration in waterbath should be less than or equal to 3 minutes
5419 | Introduction of Pompe ICB Production Media GZM 2.2 in 3K SUM never assessed or qualified | Timothy.houck@sanofi.com | 8NYA Solution Preparation | 2020 | 2025-05-07 | C&Q, QAV and MSAT would include in CCR assessments the need for Mixing PQ for new media
5418 | Monthly (April) Thyrogen QCA Review PM Not Completed | nicholas.shoenfelt@sanofi.com | 45, 51, 80 NYA | 2025-05-01 | 2025-05-01 | It is expected the monthly PMs are completed within the month they are due. | Yes | Deviation
5417 | 8NYA GEMBA- Support services - sample ID correction with no verifier | JACK.LAPLANTE@SANOFI.COM | 8NYA 8RM1350 | 2025-04-24 | 2025-05-08 | Path forward | Yes | No Deviation
5416 | No Product, No Lots - 68 NYA QC Sample Management- Single Wrap TSA Lot 25GN00915 found to be unsealed along the packaging seam | elizabeth.endsley@sanofi.com | Sample Receipt Room 4720 at 68 NYA | 2025-05-07 | 2025-05-09 | It is expected that the packaging for these places are fully sealed and intact.
5415 | BDT GZ6874-01 labelled incorrectly | joshua.brooks@sanofi.com | 74NYA UT1-02. However the deviation to procedure occured at 11 Forbes | 2025-02-13 | 2025-05-08 | Bowie dick tests packs will have labels packaged inside the box in a sealable pouch containing an approved label for each packet inside the box
5414 | 76NYA-Biosurgery APU-Room 3103-HABP-Pressure Gauge F-6938-036 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-3103 | 2025-05-08 | 2025-05-09 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5413 | MCC 2560 Hold for EM not acquired by Column, | sean.gormley@sanofi.com | 8RM1403 | 2025-05-09 | 2025-05-09 | Columns should be able to acquire EM's as necessary during production run. |  | No Deviation
5412 | 8NYA-DOWNSTREAM Q-MEMBRANE PROBE TEMP and pH Probes | Marc.Latta@sanofi.com | 1503 | 2025-05-08 | 2025-05-08 | FBL-CST-000029
5411 | 250L Eluate SUM bag failed visual inspection part#692168 Batch:23GN02369 | Yves.Antoine@sanofi.com | 8NYA, Room 1333 | 2025-05-08 | 2025-05-08 | Expect the bag to be free of particulates and abnormality
5410 | Multiple - Training of MABC-QU-REC-0767552-SD-MABC-QU-SOP-0084572-version 4.0 unable to be documented | michael.knapp@sanofi.com | 8NYA | 2025-05-08 | 2025-05-08 | Team able to submit training
5409 | Use of Spotfire Dev Data in GxP Assessments for 51 NYA (TSH) | mark.lindblad@sanofi.com | 51 NYA (impacted building) Identified at 68 NYA RM 1160 | 2025-05-08 | 2025-05-08 | Data queries in Spotfire (dev) for TSH visualizations are not validated for GxP support. No Spotfire (prod) exists for TSH visualizations. Per Data Integrity standard (QU-STD-0000244), Section 3.1.1, GxP data values must be recorded, calculated, analyzed, and reported correctly. | Yes | Deviation
5408 | QA kickbacks for RF503F002 | kurt.marsocci@sanofi.com | 8 NYA Phase 1 | 2025-04-13 | 2025-05-06 | Get clarification on how to fix these comments considering the operators involved are not in today to explain, but this record needs to be closed today.
5407 | 8NYA - Support Service - CZ - RF616F014 / 760737 - 3k SUM bag agitator did not fit in SUM | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-05-08 | 2025-05-08 | Hold Gemba to see path forward.
5406 | Third Party Complaints opened in Non Significant Quality Events | sheila.pandolfo@sanofi.com | MA Biocampus | 2024-03-28 | 2025-04-17 | The extra steps to notify the vendor of the event or request investigation from the vendor are not required per GQCMS-QU-GOP-0000974-Management of Third Party Complaint (v1.0) or GQCMS-QU-GOP-0000518-Management of Deviations in QualiPSO (v1.0 or 2.0). It is expected a TP complaint child record would not be created for non significant quality events. It is expected the non significant events would be closed by the due date and not held open for vendor investigations.
5405 | Compressed Gas Cylinders Left outside 51NYA Gas Shed | kimberly.rice@sanofi.com | 51NYA Exterior Gas Shed | 2025-05-07 | 2025-05-07 | Tag on Gas Cylinders State "Storage Condition +15C to + 30C"
5404 | 8NYA GEMBA Phase 2 Downstream. 200L Barrel bag failed Pre-use inspection | alex.tello@sanofi.com | 1503 | 2025-05-08 | 2025-05-08 | Pass visual inspection
5403 | 51NYA-Thyrogen APU-Room 143-Dept SS-Filter Integrity Tester F-15260 found non-functional for the pressure drop functions during a scheduled calibation | robert.bauch@sanofi.com | 51-143 | 2025-05-01 | 2025-05-08 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5402 | Cz Phenyl Eluate- Transfer not able to be completed due to open action notice. | nathaniel.barnes@sanofi.com | 8 NYA 1503 | 2025-05-08 | 2025-05-08 | determine potential impact of process holding |  | Deviation
5401 | No Product - No Lot - 68NYA QC Micro - No reportable results for Personnel Monitoring and Settling Plates | franchesca.heine@sanofi.com | 68 NYA / 4700 / QCM | 2025-05-01 | 2025-05-07 | Per QU-SOP-0019490 step 6.4.1, data entry for OneLIMS samples must occur concurrently with counting activities. The expectation is to enter results concurrently while counting plates. | Yes | Deviation
5400 | 8NYA-Gemba-Phase 2 Downstream Shipping Bags multiple pre-use inspection failures | aidan.mula@sanofi.com | 8NYA/RM1703 | 2025-05-06 | 2025-05-06 | Bags should not contain particulates within the product contact area.
5399 | QU-REC-0686863-SD Retired with out L&D and QA ops signature | bhargav.patel@sanofi.com | 1403 | 2025-05-06 | 2025-05-06 | per SOP, L&D and QA approval for SD retirement,
5398 | "68NYA-QC Microbiology-media lot not logged for ID testing" | Peter.Gagne2@sanofi.com | QC Microbiology, 68 NYA Lab 4700 | 2025-04-02 | 2025-05-06 | per SOP QU-SOP-007045 step 6.1.3 " Log all media, reagents, lot numbers and expiration dates on page 1 of QU-FOR-0116063 Media/Reagent Log. Record sample ID numbers and corresponding item log numbers used for testing on page 2".
5397 | QC Microbiology media log event | Peter.Gagne2@sanofi.com | Lab 4700 | 2025-04-02 | 2025-05-06 | per SOP QU-SOP-007045 step 6.2.7 "Prepare and complete Attachment 1A, QC Microbiology ID Lab Isolate Identification worksheet for each isolate"
5396 | Wrong Crimper Used for NovaSeal Pinch Pipe | Dang.tran@sanofi.com | 8NYA, 1703 | 2025-05-06 | 2025-05-06 | Using right equipment for the right SUT
5395 | 76NYA-Biosurgery APU-Room 3103-HABP-Pressure Gauge F-6938-034 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-3103 | 2025-05-06 | 2025-05-07 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5394 | 8NYA All - Fire alarm and loss of differential pressure in several areas | wilfredo.lopez@sanofi.com | 8NYA | 2025-05-07 | 2025-05-07 | Gemba to determine any potential impact to ongoing operations. Area cleanings coordination required. Path forward for resuming operations. | Yes | Deviation
5393 | GEMBA: MLC Warehouse 8NYA: SUT in question | william.bradley@sanofi.com | 8NYA RM1130 | 2025-05-06 | 2025-05-06 | Packaging should be in good condition without any kind of deformities or abnormalities
5392 | 2 Logbook errors @ 51NYA | adam.rowe@sanofi.com | 51NYA, rm129 | 2025-02-13 | 2025-05-05 | The date should be correct | Yes | Deviation
5391 | Motorized crimper failed during Cz freeze bag operations | zachary.rickard@sanofi.com | 8NYA 1703 | 2025-05-06 | 2025-05-06 | Crimpers will perform as intended and properly crimp the freeze bag to allow for disconnection from the ROSS manifold. |  | Deviation
5390 | Pre-Use Shipping Bag inspection performed by QA not documented contemporaneously | timothy.ryley@sanofi.com | 8RM 1704 | 2025-04-04 | 2025-05-06 | Signature for Pre-Use Bag Inspection should be executed when inspection is completed.
5389 | 8 NYA Gemba - Support Services - Failed 1K Tote Bag, and 1k SUM bag Inspection | nicholas.grote@sanofi.com | 08-1202 | 2025-05-06 | 2025-05-06 | No particulates on tote bag.
5388 | Cycle count recount sheet inadvertently discarded. | Joshua.brooks@sanofi.com | JIT warehouse | 2025-05-05 | 2025-05-05 | cycle count recount version 1 of cycle count 1219 count sheet should be retained.
5387 | 8NYA-No Product- 200L SD Tote bag P/N 691804 (x3) inspection failures | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2025-05-05 | 2025-05-05 | It is expected that SUT meet passing criteria outlined in SOP
5386 | Equipment (76-PUMP-10) turned off during production of HA between steps 7.8.13.4 and 7.9.8.4 in QU-OPE-0309296 | relja.nedeljkovic@sanofi.com | 76 NYA room 364 | 2025-05-06 | 2025-05-06 | Per QU-OPE-0309296 7.8.13.4 should remain in the "ON" state.
5385 | GEMBA - 8 NYA Suite Two - Recipe Did Not Launch Correctly to ROSS HMI | adam.keister@sanofi.com | 8 NYA Room 1703 | 2025-05-06 | 2025-05-06 | For recipe to run correctly. |  | No Deviation
5384 | WFI Flush WO not completed | Thomas.Loik@sanofi.com | 8YNA | 2025-05-03 | 2025-05-05 | Infor WO 3337064 would be fully executed
5383 | 8NYA Upstream P2 - FGHBN011 - Cz - MCC eluate SUM 2 product leak". | brett.simmons@sanofi.com | 8NYA, RM1403 | 2025-05-06 | 2025-05-06 | Material to be contained with no leaks, GEMBA required to determine corrective action.
5382 | No Product, No Lots - 68 NYA QC Micro - Single Wrap TSA Lot 25GN00650 found to be unsealed along the packaging seam | amanda.lavalley@sanofi.com | 68 NYA/ QCM 4700/ cold room 4755 | 2025-04-30 | 2025-05-05 | It is expected that the packaging for these places are fully sealed and intact.
5381 | GEMBA: 8NYA: Downstream: Missing Witness verification on BIRF (Formulated Bulk Cerezyme Intermediate) | william.bradley@sanofi.com | 8NYA . Originally, this intermediate would be in RM1220, but it was moved and stored in the JIT. Email of the issue was sent out at 1451 | 2025-05-05 | 2025-05-05 | Entries are to be complete real-time
5380 | 8NYA Gemba - WFI drop in Solution Prep workspace on loop 3B 8-WFI-3B-3 | JACK.LAPLANTE@SANOFI.COM | 8rm1350 | 2025-05-04 | 2025-05-04 | Path forward | Yes | No Deviation
5379 | GEMBA; 8NYA: Downstream: Pre Use bag Failure ( 650L P/N 813910) | william.bradley@sanofi.com | 8NYA 1503 | 2025-05-05 | 2025-05-05 | Single Use technology is to be in good standing condition without deformities
5378 | 8NYA Gemba - Support Services - Gemba ID 5372, SUT Agitator Shaft failed | JACK.LAPLANTE@SANOFI.COM | 8RM1350 | 2025-05-03 | 2025-05-03 | Path forward | Yes | No Deviation
5377 | 68NYA-Support APU-Dept QCM-Room 4700-Glass thermometer FB-5294 no as found data taken, dropped by Metrology. | robert.bauch@sanofi.com | 68-1830 | 2025-05-05 | 2025-05-05 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5376 | 250L SUM bag failed visual inspection | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-05-04 | 2025-05-04 | The bag should be free of particulates
5375 | GEMBA for Insect Discovered in Bioburden Filter Cup - Associated Sample ID 1821088 ( CZ2_Q_761221_LOAD) | jessica.kychler@sanofi.com | QC Micro 68 New York Avenue, floor 4, room 4700 | 2025-05-03 | 2025-05-03 | It is expected that materials used for testing are free from debris/insects/etc | Yes | Deviation
5374 | 8NYA - Support Service - CZ- RF605F007/760722 - No second verify for sample ID correction | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-26 | 2025-05-03 | Hold Gemba to see path forward.
5373 | Modbad on bleed pressure sensor X2250 placing recipe into hold | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-05-03 | 2025-05-03 | There should not be any modbads on pressure sensors
5372 | 8NYA - Support service - FZ - RF543F004/705676 - 3000L Sum bag failed inspection - agitator shaft issue | bailey.gleason@sanofi.com | 1350 | 2025-05-03 | 2025-05-03 | SUT be ready to be used by MFG with no issues | No | No Deviation
5371 | 8NYA - Support service- FZ RF568F004 - PH standardization expired prior to testing of sample | bailey.gleason@sanofi.com | 1350 | 2025-04-17 | 2025-05-03 | hold gemba to see path forward
5370 | Large amount of QCA alarms following work under AN 2025-T-012 | jesse.demers@sanofi.com | 45 NYA | 2025-05-02 | 2025-05-02 | *
5369 | CAPA-046447 was to access low flow alarms for CZ Down Stream Operations 1 low flow alarm identified. | james.hubbard@sanofi.com | 8 NYA | 2025-04-02 | 2025-05-02 | Expected result is that the operation would not have a low flow alarm.
5368 | Gemba- CZ Eluate Storage one-step transport Fridge and cold rooms not in MES | JONATHAN.AHIABA@SANOFI.COM | 8 NYA Upstream 1403 | 2025-04-24 | 2025-04-24 | Cold rooms and Fridges assigned in MES
5367 | 8NYA-Gemba-Downstream Phase 2 Cz Osmometer Standardization passed in MES but fails per SOP | aidan.mula@sanofi.com | 8NYA/RM1503 | 2025-05-02 | 2025-05-02 | SOP and MES should align
5366 | 8NYA - Support Service - CZ RF621F072 / 760741 - Incomplete comment for manual weigh out | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-14 | 2025-05-02 | Hold gemba to see path forward
5365 | 8NYA Downstream P2 - SUM bag failed pre-use inspection | wilfredo.lopez@sanofi.com | 8RM1503 | 2025-05-02 | 2025-05-02 | Expected that no contaminants are observed present on the new SUT.
5364 | 8NYA Downstream P2 - SUM bag failed pre-use inspection | wilfredo.lopez@sanofi.com | 8RM1503 | 2025-05-02 | 2025-05-02 | Expected that no contaminants are observed present on the new SUT.
5363 | 8NYA Upstream P2 - FGHBN011 - Cz - MCC eluate SUM 1 product leak | jarod.mcanern@sanofi.com | 1403 8NYA | 2025-05-02 | 2025-05-02 | Material to be contained with no leaks, a following GEMBA to determine corrective action.
5362 | WE NEED A GEMBA TO BE HELD IN ORDER TO CLOSE OUT THIS RECORD; Particulate found in Q load SUM bag | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2025-04-27 | 2025-04-27 | Per SOP-0084572, step 6.2.13, particulates should not be found. If particulates are within the interior are not acceptable and bag cannot be used. GEMBA must be called.
5361 | Cerezyme - Muliple Lots - 8NYA - Suite 2 Downstream - Multiple Alarm Setpoints Outside CST Acceptable Ranges | nicholas.shoenfelt@sanofi.com | 8NYA Mulitple | Unknown | 2025-04-30 | It is expected that all process alarms and set points align with MABC-QU-OPE-0524531.
5360 | 8NYA-Cz-FGHBN008- SUB hold on X2230, no media input during hold | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2025-04-26 | 2025-05-01 | it is expected that there are no restrictions in the Media supply line and that alarm response procedures are followed if there is a recipe hold
5359 | 8NYA - Downstream CZ - 200L SUM Bag Failed Pre-use Inspection | katherine.reardon@sanofi.com | 8 NYA 1530 | 2025-05-01 | 2025-05-01 | Per QU-SOP-0084572 step 6.2.13, bags should be free from particulates
5358 | Multiple issues with outstanding cold room logbooks for 51-108C/109 | Alyson.Belleville@sanofi.com | 51 NYA, rooms 108C and 109 | 2025-04-30 | 2025-04-30 | It is expected that all entries in a logbook would be complete and legible.
5357 | MCC Column connected is not the column issued to the BOM | brian.o'rourke@sanofi.com | 8NYA Room 1403 | 2025-05-01 | 2025-05-01 | Ensuring that the SUT attached to equipment is the SUT scanned.
5356 | 8 NYA - Support Services - Lot numbering discrepancy for 760741 media in Athena/MES | MARIE.VALLIERE@SANOFI.COM | 8 NYA | N/A | 2025-05-01 | It is expected that lot numbers are appropriately assigned per MABC-QU-SOP-0068781-Lot Numbering System (v12.0), section 6.6 and MABC-QU-FOR-0010819-Lot Numbering System for 8 NYA Operations (v9.0)
5355 | GEMBA 8NYA Autoclave, Late RQ | david.zhong@sanofi.com | 8NYA Autoclave | 2025-05-01 | 2025-05-01 | RQ to be completed by due date; QU-OPE-0134974 — Sub-VMP: Requalification, MA BioCampus, Step 6.3.10
5354 | Freezer 80-141 failure | andy.le@sanofi.com | 80NYA | 2025-04-30 | 2025-04-30 | n/a
5353 | 76NYA-Support APU-Dept QCBP-Room 107-Weight Set F-7477 overdue for calibration | robert.bauch@sanofi.com | 76-107 | 2025-05-01 | 2025-05-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
5352 | CCR work for CCR 2023FRACC0004 was closed in decommissioned Phenix system | michael.montiverdi@sanofi.com | 74 NYA (COVID) | 2025-04-15 | 2025-04-30 | CCR work completed in a commissioned quality system
5351 | 45NYA- Support APU-Room 1139-Dept QCC-EMS readout FB-1465-300 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 45-1139 | 2025-04-30 | 2025-04-30 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5350 | GEMBA - 8 NYA Downstream Suite Two - No available Sample Lines for CM EOH Sampling | adam.keister@sanofi.com | 8 NYA, room 1503 | 2025-04-30 | 2025-04-30 | Path forward on how to sample.
5349 | 68NYA-Support APU-Room 1830-Dept Metrology-Ultrasonic flow meter S-1174 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2025-04-14 | 2025-04-30 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5348 | 8NYA - Support Service - CZ RF618F013 760739 - Freezer containing mod enzyme not in MES for 760739 record | Jack.Laplante@sanofi.com | 8NYA RM1333 | 2025-04-29 | 2025-04-29 | equipment be selected in MES
5347 | MCC X-2831 Eluate SUM (PBM2845001)Temperature excursion | Yves.Antoine@sanofi.com | Rm 1402, 8NYA | 2025-04-26 | 2025-04-29 | Expect to stay within range of the CST
5346 | bags Ph probe failed standardization | Pavel.dobrovolskiy@sanofi.com | 8NYA 1503 | 2025-04-28 | 2025-04-28 | its expected that
5345 | Sartorius Biosealer TC Asset: 2012081 had Control Board replaced by vendor during repairs with no CCR | paul.mclaughlinjr@sanofi.com | 68NYA | 2025-04-29 | 2025-04-29 | per GQCMS-QU-GOP-0000514-Change Control Management (v3.0), all changes that affect validated assets require a formal Change Control Request (CCR) | No | Deviation
5344 | REGEMBA for FZ 500mL Flask Sample | ryan.devanna@sanofi.com | 8 NYA, innoc prep | 2025-04-29 | 2025-04-29 | n/a
5343 | Defective bag discovered during bag check | Luis.dominguez@sanofi.com | 51NYA room 131 | 2025-04-28 | 2025-04-28 | it is expected per QU-SOP-0041773 that the bags be integral.
5342 | 8NYA - Support Service - Spider found in CNC | bailey.gleason@sanofi.com | CNC | 2025-04-29 | 2025-04-29 | IPA area, collect bug, hold gemba
5341 | 8NYA - Support Service - 692169 2 x 3000L SUM bags failed pre inspection | bailey.gleason@sanofi.com | 1350 - Solution prep | 2025-04-28 | 2025-04-28 | expect to observe no particulate
5340 | 8NYA - Support Service - RF1AUF019 1000L SUM bag failed pre inspection | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-28 | 2025-04-28 | N/A
5339 | AN stop signs were not posted onto impacted areas entrances | Pavel.dobrovolskiy@sanofi.com | 8nya 1323 | 2025-04-28 | 2025-04-28 | its expected that every room thats impacted by the AN will have a "stop" sign attached
5338 | Fz-ICB DS Lots EGHDT011, Month 4 & BGHDT020 Month 4: Activity out of window testing. | yekaterina.astakhov@sanofi.com | QC Chem | 2025-04-25 | 2025-04-25 | It is expected that testing occurs in window.
5337 | Fz-ICB DS Lot EGHDT011 PN 728285, 1 Month Time Point: No Result (NR) generated for Aggregation | yekaterina.astakhov@sanofi.com | QC Chem | 2025-04-25 | 2025-04-25 | It is expected that Aggregation result is generated for the 1 month time point.
5336 | Cm eluate bag failed visual for broken Ph probe connection | Pavel.dobrovolskiy@sanofi.com | 8nya 1503 | 2025-04-28 | 2025-04-28 | its expected that all bags coming onto the manufacturing floor will pass visual inspection
5335 | Freezer FB-1244 | craig.valery@sanofi.com | 80-150 | 2025-02-19 | 2025-04-24 | it is expected that as founds be obtained before shipping to vendor for repair
5334 | CZ - SOPs are missing Profibus connection for ph adjustments / dilutions | Pavel.dobrovolskiy@sanofi.com | 8nya 1503 | 2025-04-28 | 2025-04-28 | its expected that manufacturing SOPs would specify which profibus connections are designated for each operation.
5333 | GXP Pm Late | andy.le@sanofi.com | 8NYA | 2025-04-28 | 2025-04-28 | MABC-QU-SOP-0081784 | Yes | Deviation
5332 | FZ 500mL Vicell counts dropped from 27Apr25 to 28Apr25 | ryan.devanna@sanofi.com | 8 NYA | 2025-04-28 | 2025-04-28 | n/a
5331 | 51NYA-Thyrogen APU-Room 132-Dept Purification-UV Indicator F-4236-009 UV Indicator had no as found readings during an unscheduled calibration due to a burned-out UV Bulb | robert.bauch@sanofi.com | 51-132 | 2025-04-25 | 2025-04-28 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5330 | Particulate found in Q load SUM bag | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2025-04-27 | 2025-04-27 | Per SOP-0084572, step 6.2.13, particulates should not be found. If particulates are within the interior are not acceptable and bag cannot be used. GEMBA must be called.
5329 | Particulates Found During ATF 10 Pre-Use Inspection | brian.o'rourke@sanofi.com | 8 NYA Room1403 | 2025-04-27 | 2025-04-27 | ATF housing should be free of particulates, debris, voids, or liquid upon initial inspection.
5328 | GEMBA: 8NYA: Downstream: 650Bag 813909 found white hair like particulate inside | william.bradley@sanofi.com | 8NYA /RM1503 | 2025-04-27 | 2025-04-27 | It is expected that bags are in good integrity with no particulates inside
5327 | 8NYA Gemba Support Services 760725 RF607F023 defective 1K SUM bag | william.lee@sanofi.com | 8NYA/1350 | 2025-04-26 | 2025-04-26 | 1K SUM bag should have the correct pressure line configuration
5326 | Upstream CZ, CIP UV measured auto selects both 280nm and 302nm when switched to auto | jeffrey.boldizar@sanofi.com | 8NYA1403 | 2025-04-26 | 2025-04-26 | the UV selection would remain the selected UV measurement
5325 | Door gasket in 1503 was found to be falling out of door frame. | kenny.leland@sanofi.com | 8 NYA room 08-1503 | 2025-04-26 | 2025-04-26 | Door gaskets should remain inside door frame.
5324 | 8NYA-Cz-FGHBN011- Eluate SUM not chilling as required during Eluate Collection | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2025-04-25 | 2025-04-25 | It is expected for the Eluate SUMs to be chilling via the recipe automation | No | Deviation
5323 | Gemba 8NYA Support Services - Freezer for Mod Enzyme Batch used without MES label | william.lee@sanofi.com | 8NYA/1333 | 2025-04-25 | 2025-04-25 | Freezer should have MES label with asset number.
5322 | 8NYA - support services - 3k bag leak on initial WFI fill | JACK.LAPLANTE@SANOFI.COM | 8RM1350 | 2025-04-26 | 2025-04-26 | Path forward
5321 | Store Recipe for zinc column held up | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2025-04-26 | 2025-04-26 | SOP-0080419-- it is expected that the recipe will finish without an issue
5320 | 8NYA Gemba Solution Prep-Transfer hose split inside of Peristaltic Pump | Bradd.Charles@sanofi.com | 8 NY AVE Room 1350 | 2025-04-25 | 2025-04-26 | it was observed that transfer line split on inside of Peristaltic pump. Consulted with QA and Batch was Discarded. | Yes | Deviation
5319 | 8NYA- Cz- FGHBN012/ FGHBN011- X2220 and X2240 pCO2 probes diverging | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2025-04-25 | 2025-04-25 | pC02 probes remain within process allowable limits |  | No Deviation
5318 | 8NYA- X2240- FGHBN012- DO Probes diverging | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2025-04-25 | 2025-04-25 | It is expected that probe 1 and probe 2 remain within 25% delta while enabled |  | No Deviation
5317 | TSK Loading Capacity Issue/Memo for Lots EGF2B11A & EGF2B11B | jacob.kallon@sanofi.com | 51 NYA Room 132 | 2025-03-03 | 2025-04-21 | Per MSAT memo, excess material should have been discarded
5316 | Periodic mixed gas PMs at 51 NYA | Alex.Biskaduros@sanofi.com | 51 Nya gas shed | 2025-04-24 | 2025-04-24 | Periodic PMs to be completed per SOP | Yes | Deviation
5315 | Regemba QE-157555 | mike.niland@sanofi.com | 8 NYA | 2025-03-21 | 2025-04-24 | None
5314 | Badge reader ID discrepancies QU-SOP-0071975 | nicholas.shoenfelt@sanofi.com | 80-101 (J.I.T. Room), 80-176 (Cryo Back-up Storage) | 2025-04-24 | 2025-04-24 | It is expected that the badge IDs for controlled area within QU-SOP-0071975 align with the ID with the secuirty system. | Yes | No Deviation
5313 | QU-FOR-0027702 Was final authorized without Doc control approval | Michael.Corrigan@sanofi.com | 76 NYA | 2025-03-28 | 2025-04-23 | the owner added all required signatories per step 6.2.5 of QU-WIN-0055911
5312 | 8NYA - Support Service - RF1AUF018 - 1k SUM bag pressure sensor line has 2 male connections | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-20 | 2025-04-20 | SUT be manufactured correctly for operations
5311 | Zinc load pH test is out of range. Samples were taken (Load Bio, Load Endo). According to record we need to run another pH adjust however all sample ports are used up.p | PAUL.DEVANEY@SANOFI.COM | downstream 1503 8NYA | 2025-04-25 | 2025-04-25 | need more sample ports on sum bag,
5310 | Aber probe divergence for SUB 2240 : Train 3 | bhargav.patel@sanofi.com | 8 NYA / 1403 | 2025-04-24 | 2025-04-24 | Probe 1 and 2 should not reading same value and control the SUB based on two probes reading average
5309 | QU-FOR-0027702 Not Final Authorized By Doc Control | Michael.Corrigan@sanofi.com | 76 NYA | 2025-03-28 | 2025-04-23 | the owner added all required signatories per step 6.2.5 of QU-WIN-0055911
5308 | 8NYA - Support Service - RF1AUF018 - 1k SUM bag pressure sensor line has 2 male connections | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-20 | 2025-04-20 | SUT be manufactured correctly for operations
5307 | GEMBA -8NYA -CZ(RF639E015)- Issues discovered during batch record review  | Yves.Antoine@sanofi.com | 8NYA Room 1403 | 2025-01-20 | 2025-04-23 | we expect ATF waste tote to be drained per SOP after the completion of ATF wetting operation
5306 | 8NYA - Support Services - Aseptiquik X membrane not integral at inspection | bailey.gleason@sanofi.com | 8RM1202 | 2025-04-24 | 2025-04-24 | The aseptic membranes will be found integral and ready for use.
5305 | 51NYA - Compressed Air 51-CA-12 failed Dewpoint | Craig.Stevenson@sanofi.com | 51 NYA Room 139 | 2025-04-23 | 2025-04-23 | It is expected that all dewpoints generate passing results.
5304 | Gemba-8NYA-Downstream suite 2- phenyl skid aquasyn valves did not open | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 room 1503 | 2025-04-23 | 2025-04-23 | N/A
5303 | ENDO/BIOB FHD samples. Not accepting storage location. | Dana.Hopkins@sanofi.com | 8NYARM1403 Upstream Phase 2 | 2025-04-23 | 2025-04-23 | Correctly log samples and perform sample in onelims.
5302 | 8NYA - Support Services/Upstream - Missing NRSR verification in record | JACK.LAPLANTE@SANOFI.COM | 8nya | 2025-04-04 | 2025-04-22 | Path forward
5301 | Missing 4 copies of Action Notice 2025-I-0040 | lauren.ruffing@sanofi.com | 8 NYA | 2025-04-18 | 2025-04-22 | 5 copies of action notice retrieved from manufacturing space
5300 | Material Not Stored Per Material Specification | Robert.petersjr@sanofi.com | 51NYA, room 108. | 2025-04-21 | 2025-04-22 | Material is expected to be stored in proper storage condition. MABC-QU-SOP-0072800-Storage and Transportation of Indirect Packages and Controlled Materials at Fram Section 6.2.1 "Store material in accordance with material’s required storage condition (temperature, humidity and lighting) as indicated in product Material Specification (MS) or Product Specification (PS) or supplier information (e.g. production record)." | Yes | Deviation
5299 | GEMBA: 8NYA Downstream Post Viral Suite 2(RM1701) Pre-Use Bag Failure | william.bradley@sanofi.com | 8NYA /RM1701 | 2025-04-22 | 2025-04-22 | It is expected that bags are in good integrity with no particulates
5298 | No incubation component assigned to bioburden samples 1792606 and 1792627 lot FGHDS020 P/N 757791 and 757792 | Stefanie.Erickson@sanofi.com | QC Micro | 2025-04-14 | 2025-04-22 | for there to be an incubation component built in onelims
5297 | 2560 Eluate SUM PBM-2570001 Not Communicating with DeltaV | brian.o'rourke@sanofi.com | 8 NYA Room 1403 | 2025-04-22 | 2025-04-22 | SUM should be reading and responsive to commands in DeltaV | Yes | No Deviation
5296 | 8NYA - Support services - Compromised 760744 seed media bottles | JACK.LAPLANTE@SANOFI.COM | 8NYA RM1350 | 2025-03-27 | 2025-04-22 | Path forward | Yes | Deviation
5295 | 8NYA Gemba - Support Services Gemba - Particulate found in transfer pump during inspection | maria.palmar@sanofi.com | 8NYA / 1350. | 2025-04-22 | 2025-04-22 | Tubing to be free of particulates/debris |  | Deviation
5294 | Incorrect asset on PM schedule | jody.teevens@sanofi.com | 8 NYA | 2025-04-21 | 2025-04-21 | Weekly PM is performed on WFI LOOP 2
5293 | 76NYA-Biosurgery APU-Room 364-Dept HAFM-Torque Wrench F-18837 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-364 | 2025-04-08 | 2025-04-22 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5292 | 8NYA-Cz-RF639F008- SUB X2230 DO Probe Divergence | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2025-04-22 | 2025-04-22 | It is expected that active probes remain within allowable limits of each other | Yes | No Deviation
5291 | "Nanofiltration System Disassembly" step is missing nano load SUM disassembly step | Robert.pino@sanofi.com | 8NYA, RM1503 | 2025-04-21 | 2025-04-21 | Expected that the nanofiltration load SUM will have an SOP step for disassembly per MABC-QU-SOP-0084572.
5290 | Punctures found on Fabrazyme media bottle tubing. | Jacob.roman@sanofi.com | 8RM1343 | 2025-04-21 | 2025-04-21 | All LFH materials are expected to be aseptically sealed prior to use. | Yes | Deviation
5289 | 51NYA - 51-CA-11 Compressed Air failed Dewpoint | Craig.Stevenson@sanofi.com | 51 NYA Room 134 | 2025-04-21 | 2025-04-21 | It is expected that all dewpoints generate passing results.
5288 | 8NYA - Support Services - OOS 705691 Media re-gemba | JACK.LAPLANTE@SANOFI.COM | 8NYA RM1350 | 2025-04-20 | 2025-04-20 | Path forward
5287 | No Products - No Lots - QC Micro 68 NYA - Missing TOC and Cond Samples for A/N 2024-F-074 | Megan.botti-marino@sanofi.com | 51 NYA | 2025-04-06 | 2025-04-21 | Per 2024-F-074, 51NYA WFI sample site 51-WFI-11 to be sampled by QC daily for TOC and Conductivity | Yes | Deviation
5286 | 8NYA GEMBA, Cleaning records executed without operator signature | alex.tello@sanofi.com | 8NYA | 2025-04-20 | 2025-04-20 | Operator actually performing the step should enter a comment, so a signature is captured in the record.
5285 | 51 NYA - AHU 7 outage/blip | nicole.lufkin@sanofi.com | 51 NYA 153, 149, 131, 132, 138 | 2025-04-20 | 2025-04-20 | cold room alarms should remain within temp
5284 | 8NYA - Support service - FZ - RF573F106/705691 OOS PH result | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-20 | 2025-04-20 | solution be within specification range
5283 | 8NYA-Gemba-Downstream Phase 2 Cerezyme-Zinc and Shipping Bag records have missing equipment log entries for allocation. | aidan.mula@sanofi.com | 8NYA/RM's 1503 and 1703 | 2025-03-30 | 2025-04-20 | Record should deallocate equipment at the appropriate time when it is no longer in process. Gemba should be called to capture allocation issues
5282 | 8NYA GEMBA Downstream Phase 2. Shipping bag failed pre-use inspection | alex.tello@sanofi.com | 1703 | 2025-04-15 | 2025-04-15 | for shipping bag to be free of any particulates and pass inspection
5281 | Leak detected during SUB decon procedure. | brendan.dickhaut@sanofi.com | 8NYA | 2025-04-20 | 2025-04-20 | clamps are attached prior to the check valve to avoid contact with NaOH
5280 | 8NYA GEMBA Phase 2 Downstream. Post Viral SUM allocated to Freeze record unable to allocate to shipping barrel record. | alex.tello@sanofi.com | 1703 | 2025-04-19 | 2025-04-19 | SUM to be deallocated so we can perform shipping barrel transfer. |  | No Deviation
5279 | Suite 1 FZ Typographical Error during EOR Sampling in BioLIMs | phoebe.karkos@sanofi.com | 1333 | 2025-04-19 | 2025-04-19 | sample time in biolims must reflect sample time in mes |  | No Deviation
5278 | 8NYA - Support Services - 705674 leak during transfer at Oetiker clamp on tubing extension (post filter) | bailey.gleason@sanofi.com | 8RM1350 | 2025-04-19 | 2025-04-19 | That SUT used in the process will not leak during use.
5277 | 8NYA-Gemba-Phase 2 Column Packing Resin Slurry Percentage high out of range | aidan.mula@sanofi.com | 8NYA/RM1303 | 2025-04-18 | 2025-04-18 | Resin concentration should remain within 40-60%
5276 | 8NYA - Support service - RF557F034 705663 - Phantom bag SU 255453 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-18 | 2025-04-18 | hold gemba to see path forward | Yes | Deviation
5275 | 8NYA SUB x2220(Train 1) Temp LOLO | bhargav.Patel@sanofi.com | 8NYARM1403 Upstream Phase 2 | 2025-04-17 | 2025-04-17 | SUB temp stays within operating perameters of 31.5 - 33.5.
5274 | 8NYA- Cz-FGHBN011- Automation intervention required to start Bleed of X2220 | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2025-04-17 | 2025-04-17 | It is expected that Bleed start based on capacitance criteria, and that the system is set up correctly to initiate bleed
5273 | Particulates Found During Inspection of 2 SUM bags | brian.o'rourke@sanofi.com | 8 New York Avenue Suite 2 Room 1403 | 2025-04-17 | 2025-04-17 | SUM bags should be free of any interior particulates.
5272 | 8NYA Gemba - Facilities/Cleanings - Cleaning records with exceptions without second verifiers. | maria.palmar@sanofi.com | 8NYA / 1350. | 2025-03-10 | 2025-03-17 | Cleaners to use a second operator verifier for manual entrys
5271 | 45NYA: Buffer sampled not initiated in onelims. No verifiable time source. | adam.rowe@sanofi.com | 45NYA, 190 | 2025-04-16 | 2025-04-17 | When a product is sampled it is expected to be initiated in onelims at that time.
5270 | GEMBA - 8NYA - Phase 1 - Shipping Bags Failed Pre-Inspection/Particulates in Bag | danny.martinez@sanofi.com | 8NYA RM 1313 | 2025-04-17 | 2025-04-17 | Expected results would be for the shipping bag to be clear of any particulates and pass.
5269 | FZ Nano SOP Discrepancy for Shipping bag Ops | Juan.Bonilla@sanofi.com | 8NYA/1313 | 2025-04-17 | 2025-04-17 | FZ SOP should not instruct Operators to turn TCU on for FZ PV Ops | Yes | No Deviation
5268 | MCC X-2831 Buffer Pinch Valves Communication Failure | brian.o'rourke@sanofi.com | 8 New York Avenue Suite 2 CNC | 2025-04-17 | 2025-04-17 | Operators should be able to set the automated valves into "auto" mode and open or close the valve.
5267 | Incorrect mop head part number referenced in MABC-QU-SOP-0084636-Facilities Disinfection at 8NYA. Support Services | robert.morawski@sanofi.com | 8NYA MFG | 2025-04-17 | 2025-04-17 | The part number of the mop head that is used at 8NYA matched the part number in SOP MABC-QU-SOP-0084636-Facilities Disinfection at 8NYA. | Yes | Deviation
5266 | 68 NYA-Support APU-No product impact-Metrology Standard S-1708 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2025-03-27 | 2025-04-17 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5265 | 8NYA Gemba - Support Services - Manual Weigh outs Documented Incorrectly 760741 | JACK.LAPLANTE@SANOFI.COM | 8NYA 8RM1364 | 2025-04-11 | 2025-04-17 | Path forward |  | Deviation
5264 | 8NYA Gemba Support Services - 705665 batched with manual agitation | william.lee@sanofi.com | 8NYA/1350 | 2025-04-15 | 2025-04-16 | Agitator should have been working and not put recipe into hold.
5263 | Login and Password information recorded on paper found in gowning area | timothy.ryley@sanofi.com | 8RM1127 | 2025-04-16 | 2025-04-16 | Passwords should not be shared/recorded on document were unauthorized personnel could utilize. |  | Deviation
5262 | Cz, upstream, X2831 regulator continually exhausting during SIP | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2025-04-16 | 2025-04-16 | regulator to maintain in correct state
5261 | Approved not Effective versions of documents used in GMP processes | Janet.langon@sanofi.com | MABC | 2025-04-15 | 2025-04-15 | The effective not approved version of VEEVA documents is used
5260 | Discuss 712 Athena Transactions identified during investigaton of QE-123390 | elizabeth.sweeney@sanofi.com | 8 NYA Solution prep | 2024-11-03 | 2025-04-16 | N/A
5259 | Nano filter inlet was leaking during priming phase of Equil recipe. | Juan.Bonilla@sanofi.com | 8NYA/1323 | 2025-04-16 | 2025-04-16 | Nano filter should not leak during priming phase
5258 | Differential Pressure sensor alarmed during Scheduled door maintenance | DEVIN.VIEL@SANOFI.COM | 76 NYA room 356/357 | 2025-04-16 | 2025-04-16 | Pressure must be 0.035-0.09 inH2O per QU-SOP-0069064
5257 | 8NYA Gemba Support Services - 705665 batched with WFI when valve was not ready to use | william.lee@sanofi.com | 8NYA/1350 | 2025-04-15 | 2025-04-16 | Flow parameters should have been downloaded to delta V prior to use of loop 2 WFI valve FV-9257064.
5256 | 8NYA Gemba Support Services - 705661 exceeded 10hr mix timer | william.lee@sanofi.com | 8NYA/1350 | 2025-04-16 | 2025-04-16 | Transfer pump should have been functioning properly. Transfer should have occurred within 10 hour mix timer limit. | Yes | No Deviation
5255 | Fraction Table for Q EGF2B11A is instructing to discard Fraction 5 | Alyson.Belleville@sanofi.com | 51-131 | 2025-04-16 | 2025-04-16 | It is expected that most fractions would be pooled, and that there wouldn't be an outlier in the middle of the fraction collection.
5254 | 8 NYA - Support Services - Issues with FAC cleaning records and ALCOA | William.Lee@sanofi.com | 8NYA | 2025-04-03 | 2025-04-15 | Per ALCOA, operator should sign for work their performed under their own name.
5253 | 8NYA-Downstream Phase 2 Gemba-Material found in CM Eluate SUM after elution | aidan.mula@sanofi.com | 8NYA/RM1503 | 2025-04-15 | 2025-04-15 | No particulates within product eluate
5252 | 8NYA - Support Service - CZ RF1QJF011 - 832327 - Agitator / Pump issues | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-15 | 2025-04-15 | Hold Gemba to see path forward.
5251 | UF\DF CIP doesn't Alaine with SOP for Remove installed device step. | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1323 | 2025-04-15 | 2025-04-15 | To get a Phath forward | Yes | Deviation
5250 | 45NYA - Possible Room breach (189B) | Craig.Stevenson@sanofi.com | 45NYA Rom 189B | 2025-04-14 | 2025-04-14 | It is expected that all electrical outlets have protective covers and have appropriate sealant around them.
5249 | Cold room 174 had temp excursion with buffers stored within. | jesse.demers@sanofi.com | 45RM174 | 2025-04-13 | 2025-04-14 | QU-SOP-0070733, Section 6.3 | Yes | Deviation
5248 | Butyl Skid Load line Aquasyn valve not opening | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1323 | 2025-04-14 | 2025-04-14 | To get a Phath forward
5247 | 8NYA Downstream Phase 1 - AQ-X connection sterile cover detached during manipulation | wilfredo.lopez@sanofi.com | 8RM1212 | 2025-04-13 | 2025-04-13 | It is expected that the sterile barriers on the SUT remain integral. The unused end of the adapter should remain covered when not in use. Need Gemba to determine potential impact to the operation from the open end of the SUT potential impact to the operation. | Yes | No Deviation
5246 | GEMBA - 8 NYA DOWNSTREAM Suite One - Retentate Flow Controller Low Low Alarm | adam.keister@sanofi.com | 8 NYA 1324 | 2025-04-13 | 2025-04-13 | For alarm to clear and recipe to resume
5245 | 8 NYA - Support Services - Failed Weld | Bradd.Charles@sanofi.com | 8 NY AVE Room 1202 | 2025-04-13 | 2025-04-13 | Seal failed. it was observed that the blade was cold upon opening machine after weld. Per MABC-QU-SOP-0084522-Single Use Tubing 6.11
5244 | 8NYA Downstream Phase 2 Gemba Request- No location found to store Phenyl Eluate in 1220 | katherine.reardon@sanofi.com | 8 NYA 1220 | 2025-04-11 | 2025-04-11 | QU-OPE-0519399 v 18 (storage conditions bottom of table 3.8) storage temp between 2 and 10 degrees celcius
5243 | after sending the mix recipe the mix timer did not come on, on delta v | KELVIN.GLEAN@SANOFI.COM | 8NYA 1323 | 2025-04-11 | 2025-04-11 | Countdown off the mix timer on delt V
5242 | Missing verbiage in SOP-0083859 Fz Zinc Chrom (Skid Set-up) | elias.mcdonnell@sanofi.com | 8NYA, 1323 | 2025-04-11 | 2025-04-11 | SOP-0083859 will have "ready for buffer manifold SUT install" rather than "Ready for SUT install" instruction.
5241 | GEMBA - 8 NYA DOWNSTREAM Suite One - Discrepancies in FZ SP SOP | adam.keister@sanofi.com | 8 NYA Room 1323 | 2025-04-11 | 2025-04-11 | For SOP to be updated.
5240 | Cleanings not performed for 4 rooms under Action Notice 2024-F-016 | craig.valery@sanofi.com | 51-130, 130A, 130B and 131 | 2025-04-09 | 2025-04-11 | it is expected that all rooms listed on action notice are cleaned following work completion | Yes | No Deviation
5239 | Rupture disk blew out of the SWECO (media sim) | MUNISH.KUMAR2@SANOFI.COM | 76-3103 | 2025-04-11 | 2025-04-11 | There should not be any loss of pressure on SWECO, and it should be maintained within 5-20 psig
5238 | 8NYA - Support Service - CZ 760741 - 3000L SUM bag leaked during initial WFI | William.lee@sanofi.com | 8NYA RM1350 | 2025-04-10 | 2025-04-10 | drain WFI, replace SUT
5237 | Pompe Vial shipment not clarified in the SOP | Daniel.bati@sanofi.com | 8NYA | 2025-03-28 | 2025-04-08 | To be able to release pompe after confirming temp trend
5236 | Modbad on bleed pressure sensor X2230 | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-04-10 | 2025-04-10 | There should not be any modbads on pressure sensors | Yes | No Deviation
5235 | Empower validated report copied to another project but the report settings are different in the new project | jacqueline.rapko@sanofi.com | N/A | 2025-03-03 | 2025-04-09 | Copied Report is the same as the source.
5234 | Magnahelic Gauge in 189B is reading zero | jesse.demers@sanofi.com | 45RM189B | 2025-04-10 | 2025-04-10 | Per QU-FOR-0091683, gauge F7019 reading should be greater than 0.05 and gauge F7021 reading should be greater than 0.01 | Yes | Deviation
5233 | 8NYA - Support Service - 705689 material added to 760741 SUM | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-10 | 2025-04-10 | Dump both batches and re batch
5232 | Downstream-Suite 1- Butyl Column Packing- Steps were not performed per EWI | aidan.mula@sanofi.com | 8RM/1303 | 2025-03-29 | 2025-04-10 | Procedures should be followed as written and all material used documented in EBR.
5231 | 8NYA - Support Service - 760738 CZ sample plan requires pre filtration osmo, EUR asks for pre and post filtration osmo. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-10 | 2025-04-10 | Sample plan and EUR to reflect correct sampling needed
5230 | Purification-Thyrogen SP2 EGF2B11A process- pH out of range during elution step at 51NYA | Sylvester.Kegbeh@sanofi.com | 51 NYA ROOMS 132 | 2025-03-25 | 2025-04-09 | It is expected that the SP2 pH is within the range of 8.1-8.9 at end of elution per QU-SOP-0061107 Column Chromatogram Review for Thyrogen at 51 NYA attachment 6 | Yes | Deviation
5229 | 8NYA - Support Services - 760738 Osmo standardization issues, 55 mM sodium citrate, 0.01% Tween 80, pH 6.3 | bailey.gleason@sanofi.com | 8RM1323 | 2025-04-09 | 2025-04-09 | That the EUR would be aligned with the SOP. That the instrument would pass daily standardization.
5228 | X2731 MCC recipe parameter missing | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-04-09 | 2025-04-09 | Recipe would move forward and arrive at prompt for SUT install
5227 | 8NYA - Support service - PH meter put Out of service after 2 false failed attempts | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-09 | 2025-04-09 | Record to provide correct ranges and not provide operator with false fails
5226 | Load Aqusyn valve did not open during Butyl skid set up recipe | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1323 | 2025-04-09 | 2025-04-09 | To get a Phath forward
5225 | Butyl SOP is missing steps to remove bubble trap filter pre skid cip | Pavel.dobrovolskiy@sanofi.com | 8NYA 1323 | 2025-04-09 | 2025-04-09 | Its expected that all steps of the EWI are converted into the new SOP format | Yes | Deviation
5224 | 8NYA - Support Service - CFH status change required | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-09 | 2025-04-09 | Hold Gemba to see path forward.
5223 | RF639F008 - 8NYA - MCC UV Standardization Recipe required automation intervention | Yves.Antoine@sanofi.com | 8NYA Room 1403 | 2025-04-09 | 2025-04-09 | UVs are zeroed during recipe |  | No Deviation
5222 | Pipette ID recorded in OneLIMS Does Not Match Pipette ID on Nexus Worksheet | venessa.carloni@sanofi.com | QC Micro Lab 68NYA Room 4790 | 2025-03-26 | 2025-04-09 | It is expected that the pipette ID recorded on Nexus worksheet matches the pipette ID recorded in bioLIMS. | Yes | Deviation
5221 | Single Wrap TSA Lot 25GN00223 found to be unsealed along the packaging seam | amanda.lavalley@sanofi.com | 68 NYA/ QCM 4700/ cold room 4755 | 2025-04-07 | 2025-04-09 | It is expected that the packaging for these places are fully sealed and intact.
5220 | 8 NYA Endotoxin Sample 25MS00479 (P/N 760731 Lot: RF613F015) -cracked sample tube | roxanne.vacha@sanofi.com | 68 NYA, 4700, QC Micro | 2025-04-09 | 2025-04-09 | Per QU-SOP-0070558 Step 6.2.3, confirm sample containers are not cracked, leaking. Samples are expected to be free of any cracks or leaks.
5219 | 8NYA-GEMBA- For different color clamps on Sub Bag SUT's | Yves.Antoine@sanofi.com | Room 1403 | 2025-04-09 | 2025-04-09 | We expected the SUB SUT's have the same colors clamps
5218 | 8NYA - Support Service - 3k SUM bag failed pre inspection | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-09 | 2025-04-09 | Hold Gemba to see path forward.
5217 | Agitator did not start after media load | Jared.Watson@Sanofi.com | 08RM1403 | 2025-04-08 | 2025-04-08 | the agitator was expected to start
5216 | 2 lots (25MN00195 and -196) of implant cards p/n 909859 from vendor Westrock failed barcode scan | alan.rousseau@sanofi.com | 74 NYA Label Shop Bar during code scan of IC | 2025-04-08 | 2025-04-08 | Per print proof approval and pre-issuance to MFG the implant card must pass barcode scanning
5215 | 8NYA Gemba Support Services-760741 No second verifier for manual weigh out | william.lee@sanofi.com | 8NYA/1350 | 2025-04-07 | 2025-04-08 | There should be a second verifier when performing manual weigh outs of material.
5214 | 8NYA Downstream Gemba Request- Discrepancy between MES instructions and SOP QU-SOP-0084570 V1.0 | katherine.reardon@sanofi.com | 8NYA 1503 | 2025-04-08 | 2025-04-08 | MES points operators to correct section of SOP
5213 | 8NYA Support Services Gemba - Incorrect pump used for 760739 batch | william.lee@sanofi.com | 8NYA/1350 | 2025-04-02 | 2025-03-13 | The correct pump 2008954 P-1520002 should have been used for liquid addition.
5212 | suite 1 ph meter out of service due to 2 failed standardization attempts | Pavel.dobrovolskiy@sanofi.com | 8NYA 1323 | 2025-04-06 | 2025-04-06 | its expected that meters are standardized daily with passing results | Yes | No Deviation
5211 | CZ Upstream, Wave equil Gas flow alarm RF658F006 | jeffrey.boldizar@sanofi.com | 8NYA1403 | 2025-04-06 | 2025-04-07 | Alarm cleared with comment
5210 | 700 Osmo standards expired and no longer carried by Thermo Fisher | Joshua.brooks@sanofi.com | 8NYA Downstream MFG Suite 2 | 2025-04-06 | 2024-04-06 | The calibration standard would be replaced prior to expiry. The standard would be in stock by our approved supplier
5209 | 45NYA Thyrogen APU-Room 190-Dept SP45-Pressure Gauge F-7283 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 45-190 | 2025-04-04 | 2025-04-04 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
5208 | Cz Zinc- Disassembly step not signed at time of completion | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2025-04-06 | 2025-04-06 | Determine impact and path forward | Yes | No Deviation
5207 | 8NYA - Support Services - 760738 buffer 10-hour timer expiry | william.lee@sanofi.com | 8NYA 8RM1350 | 2025-04-06 | 2025-04-06 | That the transfer operation would be completed within the 10 hour timer period. | Yes | Deviation
5206 | 8NYA - Support Services - 760737 OOS ph Value and Polysorbate addition | william.lee@sanofi.com | 8RM1350 | 2025-04-06 | 2025-04-06 | That the polysorbate would have been fully incorporated into the solution during the recirculation process. That the pH would return a result within range.
5205 | Butyl Eluate failed to be assigned to the tote PBT-3020002 | Pavel.dobrovolskiy@sanofi.com | 8NYA 1323 | 2025-04-06 | 2025-04-06 | Its expected that record will assign the correct container type / id to the SU during stock creation
5204 | 8NYA- Support Service - RF618F010 - Pump P1520002 mod bad alarm, Preventing precursor buffer addition. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-06 | 2025-04-06 | Pump to run according to recipe
5203 | Gemba-8NYA-multiple steps not signed on CZ Zinc record | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 room 1503 | 2025-04-05 | 2025-04-06 | N/A
5202 | 8NYA- Support Service - RF1AUF016 - MES record is indicating the PH standard is okay when the meter is standardized to the previous standard | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-05 | 2025-04-05 | hold gemba to see path forward
5201 | POU sample submitted with incorrect solution (1749034) | Pavel.dobrovolskiy@sanofi.com | 8NYA 1323 | 2025-04-05 | 2025-04-05 | Its expected that required samples will be taken from the correct buffer totes / barrels
5200 | 8NYA- Support Service - Transfer record allows operators to perform fill prior to handing out the transfer line | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-05 | 2025-04-05 | MES signal "flag" should not have let the operator proceed, signal flag tote transfer complete.
5199 | MQV Late Training | susan.mcnamara@sanofi.com | N/A | 2025-02-13 | 2025-04-04 | GQT - Curriculum - Managerial Quality Visit (MQV) - English is required training prior to performing MQV.
5198 | Suite 2 - Downstream - Phenyl tCCR does not have enough storage bufffer | Catherine.Kuzdzal@sanofi.com | 8 NYA Suite 2 Downstream | 2023-12-05 | 2025-04-04 | It is expected that the tCCR changes would have the appropriate volumes captured
5197 | 8NYA - Support service - RF610F006 / 760728 Buffer was tested under the incorrect PH standard | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-04-04 | 2025-04-04 | Hold Gemba to see path forward.
5196 | 8NYA-Cz-Failed inspection of SUB bag | Yves.Antoine@sanofi.com | 8NYA Room 1403 | 2025-04-04 | 2025-04-04 | Expect SUB SUT to pass Inspection
5195 | Partial data parsed for zero filter sample (1696492) during FB EGF4B09 process | Sylvester.Kegbeh@sanofi.com | 51 NYA ROOMS 88 | 2025-03-17 | 2025-03-17 | it is expected that all data obtained during zero filter should be parsed in BioLIM QU-SOP-0074750 step 6.8
5194 | 8NYA - Support service - SUM 1510 Ethylene glycol bleed valve left open, glycol spill on mezzanine | bailey.gleason@sanofi.com | 8NYA RM1350 SUM 1510 ethylene glycol drop | 2025-04-04 | 2025-04-04 | hold gemba to see path forward | Yes | No Deviation
5193 | Wrong chart recorder was stopped while performing Reaction for Lot FBYCMC040 GZ2139-32 | Jan.stepanik@sanofi.com | 76-229 | 2025-04-04 | 2025-04-04 | Chart recorder should not have been stopped during the reaction. | Yes | Deviation
5192 | 8NYA-FZ-FGHCE002- No Manual Pinch Clamp Installed On active Eluate SUM | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2025-03-31 | 2025-04-03 | It is expected that SUM Install be performed per SOP, with a manual clamp installed on bottom outlet line
5191 | CZ MEdia 760741 Tote Leaked During Tote Installation ¦ Support Services | christian.galvez-tello@sanofi.com | 8NYA RM1210 | 2025-04-03 | 2025-04-03 | Tote outlet not to leak.
5190 | Particle found in cell count sample during trypan mixing step. | david.loik@sanofi.com | 51NYA | 2025-04-03 | 102504/03/2025 | Particles would not be present in the sample.
5189 | Bioreactor X-2210 Motor block (16.5 degrees) not aligned with X-2200/intended (19.6 degrees) | drew.hopkins@sanofi.com | 1333 | 2022-05-13 | 2025-04-02 | Bioreactor skid X-2210 motor block aligns with X-2200
5188 | 8 NYA P1 Upstream Small Scale Inoculation Prep Room 1343 Weekly Room Cleaning Expired | GARY.DODAKIAN3@SANOFI.COM | 8 NYA room 1343 | 2025-04-03 | 2025-04-03 | QU-SOP-0084636 / MABC-QU-OPE-0524531
5187 | No Tracking Method for Butyl Eluates When Transporting from Cold Room to JIT | danny.martinez@sanofi.com | 8NYA RM 1220, RM 1131 | 2025-04-03 | 2025-04-03 | Have a method to track the Butyl eluates when transporting from the cold room to the JIT.
5186 | 8NYA GEMBA - Support services - gemba #5185, OOS pH on 705673 butyl buffer | JACK.LAPLANTE@SANOFI.COM | 8NYA RM1350 | 2025-04-03 | 2025-04-03 | Path forward
5185 | 8NYA - Support Services - pH result is OOS for 705673 (RF550F007) | christopher.brewster@sanofi.com | 8 NYA, Rm. 1350 | 2025-04-03 | 2025-04-03 | The acceptable pH range for this buffer is 6.85 - 7.34.
5184 | Q Membrane pump chamber inlet broken | charles.salgadinho@sanofi.com | 8NYA, 1503 | 2025-04-02 | 2025-04-02 | Pump heads being intact upon receiving
5183 | High pH values found during chromatogram review | cval269@yahoo.com | 51-132 | 2025-04-28 | 2025-04-02 | it is expected that the pH remains with in specified range
5182 | Thyrogen Overdue Training Items found during Internal Audit | brian.mcgee@sanofi.com | 45/51 NYA | 2024-05-10 | 2025-04-01 | It is expected that employees will be trained on the current versions of documents that they are actively using to perform tasks
5181 | 8NYA - Support Services - RF1QJF007 - pH test OOS for 832327 0.1NaOH/Hexylene Glycol | william.lee@sanofi.com | 8RM1350. | 2025-04-02 | 2025-04-02 | That the pH would be within specification.
5180 | Repligen Chelating Sepharose Column (PN 847442, Batch 23GN01500) Does not meet Flow Rate @ 1.0 Bar Specification | sheila.pandolfo@sanofi.com | 68 New York Ave | 2024-07-23 | 2025-03-28 | It is expected the flow rate @ 1.0 bar is within the specification in MABC-QU-OPE-0323091-Material Specification for Opus Column – Fast Flow GE Chelating Sepharose (v2.0), Section 6.1 Specification Table. It is expected the values reported on the CoA by Repligen are accurate.
5179 | Buffer tote with buffer P/N 832327 leaking at the bag transfer line connection point | Marc.Latta@sanofi.com | 1212 | 2025-04-02 | 2025-04-02 | buffer tote should be leak free
5178 | 8NYA - Support Service - Bimonthly cleaning performed incorrectly | bailey.gleason@sanofi.com | RM1202 | 2025-04-01 | 2025-04-01 | hold gemba to see path forward
5177 | Missing pre-approval signatures for Engineering Change Management (ECM) forms | ana.salichs@sanofi.com | 68NYA | 2025-04-01 | 2025-04-01 | For ECMs to comply with SOP signature requirements.
5176 | MAU-8400001, MAU-8401001 failure | andy.le@sanofi.com | 8NYA | 2025-04-01 | 2025-04-01 | MABC-QU-SOP-0008720
5175 | MAU-8400001, MAU-8401001 failure | andy.le@sanofi.com | 8NYA | 2025-04-01 | 2025-04-01 | MABC-QU-SOP-0008720
5174 | MAU-8400001, MAU-8401001 failure | andy.le@sanofi.com | 8NYA | 2025-04-01 | 2025-04-01 | MABC-QU-SOP-0008720
5173 | Thyrogen-EGF3B04-45NYA-QC Chemistry-OOT per QU-FOR-0024431for Thyrogen Quantitative Purity Determination by SDS-PAGE | Aimee.Park@Sanofi.com | 45NYA, QC Chemistry Room 1136 | 2024-12-20 | 2025-03-31 | Final Assay result: Average % Purity result was 98.30% would have been within Thyrogen trend limits of 98.67-99.72% per QU-FOR-0024431.
5172 | CNC Cleaning Status | michael.yashin@sanofi.com | 8NYA, Room 1210 | 2025-04-01 | 2025-04-01 | Operation should not require a bi-monthly cleaning to be performed.
5171 | Zinc eluate sum agitator failed to start during eluate mixing | Pavel.dobrovolskiy@sanofi.com | 1503 | 2025-03-31 | 2025-03-31 | Its expected that the agitator motor will turn on when prompted by the precipe
5170 | 8NYA - Facilities cleaning performed non contemporaneously | bailey.gleason@sanofi.com | N/A | 2025-03-31 | 2025-03-31 | Hold Gemba to see path forward. |  | Deviation
5169 | Compressed Air went down, causing Clean steam to go down. | DANIEL.BOIES@SANOFI.COM | CUB 55, and 51 Mech RM | 2025-03-31 | 2025-03-31 | unknown
5168 | Cerezyme-8NYA-Downstream-No instructions for selecting the correct standardization range for product in MABC-QU-SOP-0084641 | carl.pierre@sanofi.com | 8 NYA -1503 | 2025-03-31 | 2025-03-31 | Standard to be mentioned for product | Yes | No Deviation
5167 | Air Removal Time Limit Alarm in RQ of HA Load 5 Mops and Buckets | DEVIN.VIEL@SANOFI.COM | 76 NYA room 352 | 2025-03-31 | 2025-03-31 | Per QU-SOP-0048454 the cycle should go into a second air purge. The cycle stayed in the first purge and continued to pull a vacuum
5166 | 68NYA Support APU-Metrology-Room 1830-no potential product impact-50g weight in weight set S-1410 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2025-02-20 | 2025-03-31 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5165 | 8NYA GEMBA - FZ Butyl Column Out of Limit bioburden CV result | alex.tello@sanofi.com | QC | 2025-03-31 | 2025-03-31 | To be within the Limit
5164 | Cz Zinc load pH out of range | edgar.vale@sanofi.com | 8NYA/Room1503 | 2025-03-31 | 2025-03-31 | pH value for adjusted eluate would be within range of 7.85 to 8.05 | Yes | No Deviation
5163 | Particulates found in 3x different ATF 10s during MFG inspection > Pre use | joseph.gonzalez@sanofi.com | 8NYA, room 1403 | 2025-03-30 | 2025-03-30 | New ATF with no particulate's
5162 | Upstream CZ, Media study, ATF 10 tube (756380) drain tee vendor defect | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2025-03-28 | 2025-03-28 | For connections to be integral
5161 | Cold Room 131 at 51 NYA went out of range overnight on 29Mar25 | Alyson.Belleville@sanofi.com | 51 NYA Room 131 | 2025-03-29 | 2025-03-29 | it is expected that cold rooms would maintain temperature. | Yes | Deviation
5160 | Particulate found during ATF inspection. | joseph.gonzalez@sanofi.com | 8NYA, room 1403 | 2025-03-30 | 2025-03-30 | New ATF with no particulate's
5159 | SOP and Record Requiring Different standard selection for 760737 | Bradd.Charles@sanofi.com | 8 NY AVE Room 1323 | 2025-03-29 | 2025-03-29 | Per MABC-QU-SOP-0084642 760737 Has a standardization solution of 290
5158 | Sample bag was found leaking post sampling | jared.Watson@Sanofi.com | 8RM1403 | 2025-03-29 | 2025-03-29 | Sample not leaking
5157 | EM plate, left settling setup found damaged on receipt by QC | Jacob.Roman@sanofi.com | 68RM4720 | 2025-03-29 | 2025-03-29 | Expected that plates are undamaged.
5156 | Missing TOC and Conductivity Data for 68NYA | salam.chamas@sanofi.com | 68 NYA | 2025-03-03 | 2025-03-26 | we should not have missing data for TOC and Conductivity @ 68NYA
5155 | Fz Butyl Column C-3110020 TOC Swab Failure | zachary.rickard@sanofi.com | QC | 2025-03-28 | 2025-03-28 | TOC swab results will be within the range presented in the cleaning verification attachment MABC-QU-FOR-0090360 | Yes | No Deviation
5154 | Monthly Trend Reports not approved in a timely manner. | Meena.Krishnaswamy@sanofi.com | MA BioCampus | 2025-03-27 | 2025-03-27 | Reports are approved in a timely manner.
5153 | Cz Phenyl Blank Run EBR not updated Support Services | robert.morawski@sanofi.com | 8NYA CNC RM 1212 | 2025-03-28 | 2025-03-28 | The Cz Phenyl Blank Run EBR was not revised along with the production EBRs, and the new buffer setup is not captured due to the lack of revision.
5152 | Master batch record change after validation import into Production | tara.girshick@sanofi.com | 8 NYA Manufacturing Execution System | 2024-04-01 | 2025-03-27 | It is expected that all changes
5151 | Monthly Trending Evaluation not performed per QU-SOP-0084232 "MA Biocampus Monthly Deviation Trending Evaluation". | Kaylyn.Manoogian@sanofi.com | MA Biocampus | 2024-12-31 | 2025-03-27 | Per MABC-QU-SOP-0084232, and Assessment Team of APU Management would attend a monthly trend evaluation meeting to identify potential trends and approve a monthly trending report.
5150 | Procedure deviations witnessed during observing 8 NYA cleaners | christine.stevens@sanofi.com | 8 NYA Rooms 1350 and 1342 | 2025-03-25 | 2025-03-25 | QU-SOP-0084646 6.1.1 and 6.2.6 states to disinfect materials prior to entering into controlled spaces and to wait the contact time after cleaning. The vacuum was brought into the Grade C space and no cleaning step was witnessed and no time elapsed if it was cleaned. QU-SOP-0084636 6.4.1.4 states to Vacuum floor. Remove any small items from floor to ensure cleanliness. MAL and PAL were said to have been vacuumed but did not occur 6.11.2.6 states to ensure appropriate contact time was met. No contact times for vacuum and PAL contact time was not going to be followed until insisted upon. 6.2.2.7 states Ensure all doors are kept closed during performance of disinfection and are closed completely upon exiting of any area. Place “Do Not Enter” signage for each area once area cleaning has been completed. Remove signage after required contact time is complete. 6.6.2.3 states to Set up mopping system from furthest point from the exit door and work towards exit. The PAL was cleaned while we were still in solution prep so the order was not followed.
5149 | MABC-QU-SOP-0003344-Portable Media Storage and Usage Log (v6.0) | Robert.petersjr@sanofi.com | 51NYA, rooms 108C and 109 | 2025-03-23 | 2025-03-27 | Operator will record the date, time and initials/date on the Removed line of ATTACHMENT 1. A second individual must sign for witnessing the media removal
5148 | Ineffective Effectiveness Checks EFF-000214 & EFF-000215 Regarding MODBAD Alarms | mark.gwozdzik@sanofi.com | 8NYA | 2025-03-21 | 2025-03-21 | CAPA's to be Effective.
5147 | 68 NYA - Thyrogen - Multiple Lots - QC unable to perform source data review on LAL Plate-Reader | amy.roy@sanofi.com | 68 NYA 4700 QCM | 2025-03-26 | 2025-03-27 | That data is retrievable upon performing source data review
5146 | PH Probe Standardization DLI Instruction unable to follow | Jonathan.Ahiaba@sanofi.com | 1403 | 2025-03-26 | 2025-03-26 | Step 6.6.16 need clarification not able to follow due to firmware update to transmitter.
5145 | Out of Tolerance occurred on 26MAR25 during Wave Rocker BR-2152001 annual PM | dom.arone@sanofi.com | 8 NYA, RM 1403 | 2025-03-26 | 2025-03-26 | GE-QU-SOP-0069490-Calibration Program (version 6.0) instructs to investigate all out of tolerance events during calibration
5144 | 8NYA - CZ RF622E018/EGHAW018 unsampled in biolims | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-10-30 | 2025-03-26 | Hold Gemba to see path forward.
5143 | 8NYA-Gemba-Downstream Fz UFDF mismatch between SOP and EBR | aidan.mula@sanofi.com | 8NYA | 2025-03-26 | 2025-03-26 | EBR and SOP should align
5142 | 4 Plates within Aseptic Qual not collected | elizabeth.endsley@sanofi.com | Sample Receipt Room 4720 at 68 NYA | 2025-03-25 | 2025-03-26 | Plates to arrive collected within the system.
5141 | MES / SOP Room Cleaning Discrepancy | Thomas.Loik@sanofi.com | 8YNA | 2025-03-26 | 2025-03-26 | DCR to align SOP, CCR to Align MES
5140 | daily construction clean walkthrough NOT PERFORMED DURING 45 nya SHUTDOWN | john.mendenhall@sanofi.com | 45 NYA | 2025-03-10 | 2025-03-19 | Per MABC-QU-SOP-0069058, a daily facility walkdown would be completed per 6.2.3 | No | Deviation
5139 | Low Viability x2200 Fz | Todd.niziak@sanofi.com | 8NYA RM 1333 | 2025-03-26 | 2025-03-26 | QU-SOP-0084639 Step 6.2.17
5138 | Low Viability x2200 Fz | Todd.niziak@sanofi.com | 8NYA RM 1333 | 2025-03-26 | 2025-03-26 | QU-SOP-0084639 Step 6.2.17
5137 | Action Notice 2025-I-025 Number of Signs Removed Does Not Equal the Number of Signs Issued | yunbo.lou@sanofi.com | 8 NYA Room 1502 | 2025-03-25 | 2025-03-25 | The number of stop signs removed shall equal the number of stop signs issued.
5136 | 8NYA DS Column Packing Discrepancy with EWI | katherine.reardon@sanofi.com | 8NYA 1303 | 2025-03-24 | 2025-03-24 | SOP or EBR revision.
5135 | Particle embedded on sample bag surface | jacob.roman@sanofi.com | 8RM1403 | 2025-03-24 | 2025-03-24 | Per QU-SOP-0084522, 6.2, we should not have particulates in the SUT.
5134 | ATF 10 Particulates | brian.o'rourke@sanofi.com | 8 NYA Room 1403 | 2025-03-24 | 2025-03-24 | ATF housing should be free of particulates, debris, voids, or liquid upon initial inspection.
5133 | Failed Pe-Use Shipping Bag Inspections | JOEL.HICKEY@SANOFI.COM | 8 NYA/Room: 1313 | 2025-03-20 | 2025-03-20 | According to QU-SOP-0080267 step 6.16.8 there must be no particulates within the product contact surface area and there must be pictures taken of particulates with a Tappi chart. Per QU-SOP-0080267 step 6.16.6 there must be no punctures to bag tubing. |  | Deviation
5132 | Out of Tolerances found on 50L Wave Rocker BR-2162001 during annual PM by performed by Cytiva | dom.arone@sanofi.com | 8 NYA, RM 1403 | 2025-03-24 | 2025-03-24 | GFGE-QU-SOP-0069490-Calibration Program (version 6.0) | Yes | No Deviation
5131 | Rm 1119 inadvertently included in MABC-QU-SOP-0084636-Facilities Disinfection at 8NYA (v2.0) | sadiq.ghojaria@sanofi.com | Rm 1119 (CNC) | TBD | 2025-03-20 | Rm 1119 should not have been included in the SOP listed above | No | Deviation
5130 | Implant card p/n 909859 is rejected for torn frayed edges | alan.rousseau@sanofi.com | NBO QC | 2025-03-24 | 2025-03-24 | expected IC cards to have clear precise cut edges not frayed and peeling
5129 | 8 NYA GEMBA - Cleaning Record missing Verifications | alex.tello@sanofi.com | 8NYA | 2025-03-03 | 2025-03-21 | Manual entries to have a verification by a second operator. |  | Deviation
5128 | ATF10 faulty clamp connection on air filter | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-03-21 | 2025-03-21 | Air filter should come from vendor placed correctly on ATF10 |  | Deviation
5127 | Trend of Dynamic Personnel EM hits at 76NYA, Grade A | michelle.smolinsky@sanofi.com | 76 NYA, Room 355 | Between October 242024 and Feb 2025 | 2025-03-20 | No trends would be identified
5126 | 8NYA - Possible breach in 8 Steam system | Craig.Stevenson@sanofi.com | 8NYA Room 1403 | 2025-03-21 | 2025-03-21 | It is expected that no GMP water be leaking from a gasket connection.
5125 | 8NYA-FZ- FGHSM003-X2510 SIP Started without Skid Hookups P/N 714046 and P/N 714044 | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2025-03-20 | 2025-03-20 | all associated MCC parts required for SIP get installed per SOP | Yes | Deviation
5124 | Skid Setup step does not open in EBR | Juan.Bonilla@sanofi.com | 8NYA/1323 | 2025-03-20 | 2025-03-20 | The Skid setup step should populate after signing for the buffer supply connection.
5123 | 3 QCC Empower Methods unaligned with filings | Sarah.Adams@sanofi.com | 68 NYA, QCC | 2023-05-31 | 2025-02-19 | Empower methods to align with filing.
5122 | 8NYA - GEMBA Downstream - Fz Nano - PV1 SUM Agitator Stayed On Overnight | danny.martinez@sanofi.com | 8NYA RM1313 | 2025-03-19 | 2025-03-20 | Expected results would be for the agitator to turn off after the final sampling for Nano is complete
5121 | asdfasdfdsaf | tanni.arif@sanofi.com | 8NYA Mech Room | 2025-03-01 | 2025-03-04 | CSG feeds autocl
5120 | IGNORE FAKE GEMBA REQUEST TO SHOW AD HOW TO REQUEST GEMBA | tanni.arif@sanofi.com | 8NYA Mech Room | 2025-03-01 | 2025-03-04 | CSG feeds autocl
5119 | 8NYA GEMBA - Support services - FIT 2009073 status was changed but was not allocated to record | JACK.LAPLANTE@SANOFI.COM | 8NYA1350 | 2025-03-20 | 2025-03-20 | Path forward | Yes | Deviation
5118 | visual inspection of two shipping bags fails, went and two more and failed as well. | KELVIN.GLEAN@SANOFI.COM | 8NYA 1313 | 2025-03-20 | 2025-03-20 | QU-SOP-0080267 | No | No Deviation
5117 | 8NYA-No Product-FGHBN003X- Leak identified on Protocol Sample x2 (Potential Vendor Defect) | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2025-03-19 | 2025-03-19 | It is expected that the sample bags remain integral
5116 | Upstream CZ, Media study, ATF 10 (754042) permeate connection and ATF 10 tube (756380)permeate line vendor defect | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2025-03-19 | 2025-03-19 | For connections to be integral
5115 | Change Control Request (CC-001663) In Routine Use Prior To Regulatory Submissions for Fz & Cz | Meena.Krishnaswamy@sanofi.com | QC Laboratory | 2025-03-18 | 2025-03-18 | Switch On Date for this change control documented in the Quality Blocking Plan prior to regulatory approvals
5114 | GQCMS-QU-OPE-2729428-Training System Effectiveness Assessment Guideline execution started prior to document being effective | Erin.Young@sanofi.com | 68 NYA | 2025-02-26 | 2025-03-19 | Per GMP, documents must be effective prior to use.
5113 | 8NYA Downstream GEMBA- Missing verifier for manual entry of biosealer | katherine.reardon@sanofi.com | 8 RM 1323 | 2025-02-20 | 2025-02-18 | Obtain verification from second operator when entering biosealer label into record manually
5112 | PH sampling out of range | maria.palmar@sanofi.com | 8NYA / 1350. | 2025-03-19 | 2025-03-19 | PH to be on range
5111 | CD-CHO Media Sample Leakage | Pushti.Shah@sanofi.com | 4700 68 NYA QC Micro | 2025-03-18 | 2025-03-18 | QU-OPE-2728785 Step 6.3.2
5110 | FZ X-2200 Fluctuating DO and Elevated O2 sparge | Jonathan.Ahiaba@sanofi.com | 1333 | 2025-03-18 | 2025-03-18 | MABC-QU-OPE-0524531, MABC-QU-SOP-0083846 | Yes | No Deviation
5109 | Third Party Change Notification not initiated for revision 3 of Cytiva change notification | Nashmin.Hojatti@sanofi.com | 68nya | 2024-09-12 | 2025-03-18 | It is expected that per QU-GOP-0000619 section 2.2.3 the entity should assess the need to create a change control within 10 working days of receipt of TPCN/TPSC (third party site concerned)
5108 | FZ Butyl unpack record is calling for an incorrect column | Pavel.dobrovolskiy@sanofi.com | 8NYA 1503 | 2025-03-18 | 2025-03-18 | Its expected that the records used will be created for the correct column being unpacked | Yes | No Deviation
5107 | 8 NYA - Support Services - WFI WO's Closed Incorrectly | nicholas.grote@sanofi.com | 8 NYA | 2025-03-15 | 2025-03-18 | WO should be tech completed and then closed by separate operators. CM type work orders should have the required comments. | Yes | Deviation
5106 | 68NYA Support APU-Metrology-Room 1830-no potential product impact-Pressure Calibrator S-1660 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2025-03-11 | 2025-03-18 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5105 | 8NYA APU-Room 1403-Dept CC8- SIP Temperature Indicator FB-2899-044 on Chrom skid X-2560 out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1403 | 2025-03-17 | 2025-03-18 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | No | Deviation
5104 | Bioburden sample 171220 missing test volume | heidy.ruano@sanofi.com | 68 New York Ave, QC Micro | 2025-03-11 | 2025-03-17 | Per QU-SOP-0073109 the testing volume should be 10mL | Yes | Deviation
5103 | Bioburden sample 171220 missing test volume | joseph.napoli@sanofi.com | 68 New York Ave, QC Micro | 2025-03-11 | 2025-03-17 | Per QU-SOP-0073109 the testing volume should be 10mL
5102 | 8NYA Support Services - Spider Found in the CNC Area | Ifeanyichukwu.Akpone@sanofi.com | 8NYA CNC | 2025-03-17 | 2025-03-17 | N/A
5101 | Breach in HA Production Area | DEVIN.VIEL@SANOFI.COM | 76 nya 3105, 363, 352, 360, and 357 | 2025-03-17 | 2025-03-17 | It is expected that after the facilities cleaning operators enter through regular | Yes | Deviation
5100 | Leak in phase 1 hallway. corridor RM 1201 | Gary.Dodakian@sanofi.com | 8NYA RM 1201 | 2025-03-17 | 2025-03-17 | Repairs need to be made to prevent future leakage. | Yes | Deviation
5099 | Upstream CZ, ATF permeate connection vendor defect | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2025-03-16 | 2025-03-16 | For connection to be intergal
5098 | 8NYA-Cz-FGHCG004-Daily Wave Sample Cedex Analysis missing GLC2B testing requirement when executed | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2025-03-16 | 2025-03-16 | It is expected that GLC2B Test be selected when running Cz CEDEX samples
5097 | UFDF Cleaning temperature specs not met | craig.valery@sanofi.com | 51-88 | 2025-03-11 | 2025-03-15 | It is expected that the sanitization buffer temperature is within the specified process range.
5096 | 8NYA-Support services- while performing bag inspection for 3k sum bag for MO RF573F070, recirculation line was found to have two male end connections instead 0ne male one female | Bradd.Charles@sanofi.com | 8 NY AVE Room 1350 | 2025-03-14 | 2025-03-14 | The bag should come from vendor with correct bag fittings.
5095 | 2 Lots in OneLIMS have no Sampling Point designation - Cannot be QC Released | kristin.gothier@sanofi.com | NA | 2025-03-14 | 2025-03-14 | Per QU-OPE-1894750-OneLIMS – Lot Disposition, Certificate of Analysis, Generation and Report Gener (v8.0), 4.1 INTERMEDIATES MANUAL LOT DISPOSITION, do not perform disposition if the summary tab contains errors.
5094 | CZ Samples Lot# FGHDS001 (ID#1607400, 1607401, and 1607402) received by mfg personnel, lot unable to be QC Released | amy.roy@sanofi.com | 68NYA/4700/QC Micro | 2025-01-05 | 2025-03-14 | QU-SOP-0070558-Quality Control Sample Receipt, Section 3, it is the responsibility of QC to receive samples
5093 | A280 Q elution buffer blank not sampled in onelims | craig.valery@sanofi.com | 51 NYA RM132 | 2025-03-13 | 2025-03-13 | it is expected that samples are documented in onelims at time of physical sampling | Yes | Deviation
5092 | Same operator signed as performer and verifier | craig.valery@sanofi.com | 51 NYA RM 138 | 2025-03-01 | 2025-03-13 | It is expected that different operators perform and verify | Yes | Deviation
5091 | 76NYA Biosurgery-Room 3102-Dept HABP-Pressure Gauge F-21886 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 76-3102 | 2025-03-13 | 2025-03-14 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5090 | Cz-ICB Capto Butyl (MCC) Lolo Lolo alarm for the regeneration 2 phase in the Delta V recipe is below the PAR in the control strategy | ilene.robinson@sanofi.com | 8NYA | 2024-03-14 | 2024-03-14 | The lolo alarm for the regeneration step would be within the PAR
5089 | DP alarm from room 234 to 233 | nicole.fragaalvarez@sanofi.com | 76-234 | 2025-03-12 | 2025-03-13 | no DP alarms
5088 | Chart Recorder in room 507 failed to record data during the night of 12Mar25 | nicole.fragaalvarez@sanofi.com | 76 NYa room 507 | 2025-03-13 | 2025-03-13 | Temperature for the room needs to be recorded in chart recorder in order to verify that room had maintaned at 15C to 30C
5087 | 51NYA – Periodic Compliance Activities Not Performed for DeltaV System | vincent.faria@sanofi.com | 51 NYA | 2022-04-01 | 2025-01-03 | Per QU-SOP-0051867, Step 6.2.3 - Admin audit review for a critical system is annually, Per QU-SOP-0051867, Step 6.3.2 - backup and restore verification performed every two years for medium system data risk.
5086 | Seprafilm Drying Chart recorder failed, preventing a batch of Seprafilm to be dried. | jan.stepanik@sanofi.com | 76-238 | 2025-03-12 | 2025-03-13 | Chart recorder should have been working.
5085 | Nanofilter (Autoclaved) Batch Record (MO RF525F001) does not capture NewAge assembly used | sheila.pandolfo@sanofi.com | Support Services | 2025-01-14 | 2025-03-12 | It is expected the batch record accurately reflects the single use technology used in the process. | Yes | Deviation
5084 | 8NYA Support Services Gemba - Autoclave used during WFI outtage | william.lee@sanofi.com | 8NYA/1350 | 2025-03-13 | 2025-03-13 | Autoclave should not have run if WFI was not released |  | No Deviation
5083 | CZ Zinc packing slurry percentage is 60.08% MES range is 40.00-60.00 | Pavel.dobrovolskiy@sanofi.com | 8NYA 1303 | 2025-03-12 | 2025-03-12 | Its expected that the three samples average out to a target of 55%
5082 | Corrective maintenance on L&P assets at 11 Forbes not being recorded in EAM | michael.letendre@sanofi.com | 11 Forbes L&P | 2025-03-12 | 2025-03-12 | QU-SOP-0072346 Step 1.1 | Yes | No Deviation
5081 | Unable to move forward with Column resin prep record | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1303 | 2025-03-12 | 2025-01-30 | To get a Phath forward
5080 | media warm times in QU-OPE-0515888 and QU-SOP-0020192 do not align | nicole.lufkin@sanofi.com | 51 NYA 152 | 2025-03-12 | 2025-03-12 | QU-SOp-0020192 attachment 1
5079 | GEMBA to determine forward processing for HA-CMC Process | jan.stepanik@sanofi.com | 76-229 | 2025-03-12 | 2025-03-12 | Processing should continue
5078 | 8NYA GMEBA: Resizing customized MES templates required for SU, HU, EQM labels | lolav.ahmed@sanofi.com | 8NYA Sol prep | 2025-03-03 | 2025-03-02 | N/A
5077 | 8NYA-No Product-X2220 Placed Back In Service and used for Media Study and ETP while ongoing Calibration/ PM work was needed | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2025-02-20 | 2025-03-11 | It is expected that equipment is confirmed OK to use during area clearance
5076 | GEMBA - 8 NYA DOWNSTREAM - Discrepancy in EWI-16 | adam.keister@sanofi.com | 8 NYA, Room 1303 | 2025-03-11 | 2025-03-11 | For EWI instructions to align with available equipment. Not all columns have these outer fastening nuts and bolts on the bottom bed support.
5075 | UFDF Production went into hold due to Membrane failure High High Alarm | Juan.Bonilla@sanofi.com | 8NYA/1323 | 2025-03-11 | 2025-03-11 | Production recipe to resume as intended.
5074 | 8 NYA Gemba - Support Services - Sodium Chloride Weigh out - 24Kg HU created during weigh out. | nicholas.grote@sanofi.com | 08-1364 | 2025-03-11 | 2025-03-11 | Weigh outs go into different SUs and labels are generated for each label. Operator should hit "new target HU" not "finish" when moving to a new bag within the same batch of material.
5073 | 68 NYA-Support APU-QCM68-Room 4700-Hygrometer F-20756 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-4700 | 2025-02-20 | 2025-03-11 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5072 | GEMBA - 8 NYA DOWNSTREAM - Automatic Equipment Status Changes Fail in MES | adam.keister@sanofi.com | 8 NYA - 1503 | 2025-03-11 | 2025-03-11 | For MES to work as intended.
5071 | 8 NYA Gemba - Support Services - Black Particulate found on inner layer of tote bag during pre-use inspection | nicholas.grote@sanofi.com | 08-1202 | 2025-03-10 | 2025-03-10 | No black particulate is found on the tote bags. | Yes | Deviation
5070 | 76NYA - No incubation time documented | Craig.Stevenson@sanofi.com | 68NYA Room 4700 | 2025-02-27 | 2025-03-10 | It is expected that the Incubation time be recorded onto the paperwork | Yes | Deviation
5069 | No incubation protocol assigned to sample 1717788 | Stefanie.Erickson@sanofi.com | QC Micro bioburden room | 2025-03-08 | 2025-03-08 | for there to be an incubation component built in onelims
5068 | Missing Paperwork for action notice 2025-R-004 | panyin.essilfie@sanofi.com | 76NYA | 2025-03-09 | 2025-03-09 | QU-SOP-0022403, QU-SOP-0113302
5067 | 8 NYA Gemba - Support Services - 760741 Media Batch Size Change - Effecting use of Pre-Weighed DTT | Nicholas.grote@sanofi.com | 08-1350 | 2025-03-10 | 2025-03-10 | Historically the batch size is designed to accept the pre weighed amounts for these two materials.
5066 | 8NYA-Fz-FGHSD002-Inoculation of SUB X2200 scheduled to take place during hours not supported by Upstream Mfg | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2025-03-10 | 2025-03-10 | It is expected that qualified operators execute production activities and are available for process operations
5065 | A GEMBA is needed for the Fz UFDF Conductivity meters switched into manual mode. | Juan.Bonilla@sanofi.com | 8NYA/1323 | 2025-03-10 | 2025-03-10 | The conductivity meters should not be switching to manual mode.
5064 | Upstream FZ, RF503F002, DO spanning must occur within 16Hr after media load | jeffrey.boldizar@sanofi.com | 8NYA 1333 | 2025-03-10 | 2025-03-10 | DO spanning to occur within 16hr at end of Media load | Yes | No Deviation
5063 | Bioburden for Clarification FGF4B01C not sampled in Onelims at time of sampling | Alyson.Belleville@sanofi.com | 51 - 138 | 2025-03-08 | 2025-03-08 | It is expected that the bioburen would be sampled in Onelims at the time it was actually sampled. | Yes | Deviation
5062 | 8NYA-Gemba-Downstream Cz Freeze Ross communication error at end of load dispensing operation | aidan.mula@sanofi.com | 8NYA/RM1703 | 2025-03-09 | 2025-03-09 | DeltaV and Ross should communicate and generate an alarm when communication issues occur
5061 | 8NYA - Support service - RF603F012 filter failed to wet properly | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-03-09 | 2025-03-09 | Hold Gemba to see path forward. |  | No Deviation
5060 | No Product - 8NYA - Media Study Tote--Liquid observed on media tubing | michael.knapp@sanofi.com | 8NYA 1210 | 2025-03-09 | 2025-03-09 | Tubing and media totes remain integral. | Yes | Deviation
5059 | 8NYA- Support service - RF603F012 Filter failed FIT twice with Error: leak test failure | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-03-09 | 2025-03-09 | Hold Gemba to see path forward. |  | No Deviation
5058 | 8NYA-Gemba-Downstream Cz Phenyl Blank Run does not have 705665 on the BOM | aidan.mula@sanofi.com | 8RM1503 | 2025-03-09 | 2025-03-09 | PN: 705665 should be on the BOM
5057 | CZ Zinc failed CV due to hi swab 3 results (8.388ppm) | Pavel.dobrovolskiy@sanofi.com | 8RM1303 | 2025-03-09 | 2025-03-09 | Its expected that all of the swabs will pass TOC (2.4ppm) | No | No Deviation
5056 | 8NYA Support Services - Spider found on floor - 1115 Waste Airlock | bailey.gleason@sanofi.com | 8 NYA - Room 1115 - Waste Airlock | 2025-03-09 | 2025-03-09 | expect no insect activity in this area |  | No Deviation
5055 | 8NYA GEMBA - CZ Nanofiltration post viral mixer not functioning. | alex.tello@sanofi.com | 1703 | 2025-03-08 | 2025-03-08 | For SUM to start mixing for 15 minutes so we can take our samples. | Yes | Deviation
5054 | SUM bag failed visual inspection due to plastic particle on the inside layer | Pavel.dobrovolskiy@sanofi.com | 8RM1323 | 2025-03-08 | 2025-03-08 | Its expected that no particles are found inside the bag
5053 | SuM bag vailed visual inspection | Pavel.dobrovolskiy@sanofi.com | 8RM1323 | 2025-03-03 | 2025-03-03 | its expected that all of the equipment that is attached to the bag will remain connected | Yes | Deviation
5052 | TSH- Buffer Prep - A compromised manifold was used to fill bag 6 of FGFB1007, Part number GZ5644-52. | jeremy.wooden@sanofi.com | 45 NYA Room 190 | 2025-02-21 | 2025-03-07 | It is expected that a compromised manifold would not be used to fill buffer into bag 6.
5051 | 8NYA - Support Services - New EUR for pH testing references incorrect verification pH value | william.lee@sanofi.com | 8RM1350 | 2025-03-07 | 2025-03-07 | That the EUR would have been field tested to catch this issue before releasing it to mfg. That the correct verifying range would be pulled into the result verification step.
5050 | Manufacturing ATF Clamp-on Flow Sensor | benjamin.larson@Sanofi.com | 8NYA | unknown | 2025-01-29 | Determine if any containing action is required.
5049 | FZ, Upstream, RF503F002. MES Errors/ New MES steps not in SOP | jeffrey.boldizar@sanofi.com | 8NYA 1333 | 2025-03-06 | 2025-03-06 | SOP instruction to perform operation properly | Yes | No Deviation
5048 | Unable to connect to FRGD589VDX2 | kayode.yusuf@sanofi.com | 68 NYA / 4700 / CBA | 2025-02-24 | 2025-02-25 | N/A
5047 | 8 NYA Gemba - Support Services - P/N 705667 has incorrect WFI amount defined in Delta V | nicholas.grote@sanofi.com | 08-1350 | 2025-03-06 | 2025-03-06 | Density factor is set correctly.
5046 | Unable to login into FRGD589VDX2 | kayode.yusuf@sanofi.com | 68 NYA / 4700 / CBA | 2025-02-24 | 2025-02-25 | N/A
5045 | Data Loss event for Framingham BMS | thomas.st.pierre@sanofi.com | Framingham Campus | 2025-02-28 | 2025-03-03 | Per MABC-QU-SOP-0073452 alarm reports are expected to be generated daily for reconciliation.
5044 | QU-OPE-2729850- version 1.0 approved in QualiPSO without document control as a final authorizer | Amy.McIntyre@sanofi.com | QualiPSO CMS | 2025-03-05 | 2025-03-06 | Expected results would be that the owner added all required signatories per step 6.2.5 of QU-WIN-0055911 | Yes | Deviation
5043 | Dead rodent mouse discovered in delta v control box | jacob.kallon@sanofi.com | 51 NYA room 139 | 2025-03-06 | 2025-03-06 | Reference MABC-QU-SOP-0005611-Pest Control-12.0-English
5042 | Chart recorder not recording 5hrs 49 mins FBYFRM011 QU-OPE-0309296 | dianne.moustafa@sanofi.com | room 364 at 76NYA | 2025-03-04 | 2025-03-04 | step 7.11.20 QU-OPE-0309296 recording of temperature, pH control, and transfer pressure
5041 | 45NYA-Support APU-Dept QCC-Room 1136-Glass Thermometer F-6375 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 45-1136 | 2025-03-05 | 2025-03-06 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
5040 | Atypical Sensor Array Standardization (Conductivity) | Craig.valery@sanofi.com | 51 NYA RM132 | 2025-01-16 | 2025-03-05 | It is expected that conductivity standardizations are performed accurately | Yes | Deviation
5039 | Stop sign missing for AN 2025-F-006 | david.loik@sanofi.com | 51 NYA | 2025-03-04 | 2025-03-04 | Signage would be present for removal. | Yes | Deviation
5038 | 45/51 NYA EM Plates Received Expired | diane.collins@sanofi.com | 68 NYA ,4720 sample receipt | 2025-03-05 | 2025-03-05 | QU-SOP-0070558 step 6.2.5, EM plates must be incubated with in 24 hours after being sampled.
5037 | 8NYA-GEMBA Request - Nano Load SUM temp excursion | stanley.durand@sanofi.com | 8NYA | 2025-03-04 | 2025-03-04 | SUM temp should maintain set point temp
5036 | RO#2 loop shutdown due low level alarm ( Alarm was triggered when system shutdown not before). RO#1 went into alarm due low pressure and did not called out. | jesus.acosta2@sanofi.com | 51-146 and 51-150 | 2025-03-04 | 2025-03-04 | RO #2 storage tank should have filled RO#1 tank with no issues. | No | Deviation
5035 | 8NYA-Gemba-Downstream Cz CM Eluate exceeded temperature for about an hour | aidan.mula@sanofi.com | 8RM1503 | 2025-03-04 | 2025-03-04 | SUM should remain in temperature
5034 | 8NYA-Gemba-Downstream Fz SP Recipe went into hold due to chilled water loop | aidan.mula@sanofi.com | 8RM1323 | 2025-03-04 | 2025-03-04 | Automated systems should work as intended
5033 | 8NYA Support Services - Failed Athena Transactions, Records needed for Shipment 05MAR2025 | Ifeanyichukwu.Akpone@sanofi.com | 8NYA | 2025-03-04 | 2025-03-04 | N/A | No | Deviation
5032 | 45/51NYA - TSH - Bioburden sample not received and expired | abel.agnes@sanofi.com | 68NYA 4720 | 2025-03-04 | 2025-03-04 | Bioburden samples are submitted to the laboratory for testing prior to the 24 hour expiry. | No
5031 | 2 Late Extended PMs for 8NYA Compressed Air 3153023 & 3153024 | tanni.arif@sanofi.com | 8NYA Mech Room | 2025-03-01 | 2025-03-04 | All PMs completed on time by original due date or, for eligible PMs extend on-time | Yes | Deviation
5030 | Possible Deviation Gemba QU-OPE-0259002 | cynthia.labaire@sanofi.com | F068 | 2025-02-28 | 2025-02-28 | GQCMS-QU-GOP-0000590 section 6.1.2
5029 | WFI range incorrect for 11 trucks option in step 7.4.9, QU-OPE-0511911 | karthik.sriram@sanofi.com | N/A | 2024-09-26 | 2025-02-28 | WFI range for 11 trucks should be +/- 1% from the target (254.9 kg - 260.1 kg) in step 7.4.9 in QU-OPE-0511911
5028 | Approved Paperwork Used by MFG to Execute Routine EM in the 76 HA areas | megan.botti-marino@sanofi.com | 76 NYA | 2025-02-15 | 2025-03-03 | It is expected that effective paperwork be used.
5027 | Conductivity Meter #2008023 Failed 2 Standardization Attempts In A Row | patrick.mitchell@sanofi.com | 8NYA/1323 | 2025-03-03 | 2024-03-03 | expected meter to standardize to 12.88 mS/cm and then reflect that in the reading.
5026 | 8NYA-Gemba-Downstream Fz Butyl NRSR was not taken | aidan.mula@sanofi.com | 8NYA/RM1323 | 2025-03-03 | 2025-03-03 | NRSR should have been taken prior to consumption
5025 | SP2 EGF4B09A pH is out of range | Sylvester.Kegbeh@sanofi.com | 51 NYA ROOMS 132 | 2025-03-02 | 2025-02-07 | it is expected that the SP2 pH is within the range of 4.8-5.6 during elution per QU-SOP-0061107 attachment 3 | Yes | Deviation
5024 | PM 2911570 not completed by extended due date | dom.arone@sanofi.com | 8 NYA, RM 1333 | 2024-03-01 | 2025-03-03 | Per MABC-QU-SOP-0021310-Maintenance Operations (v19.0), if extended PM is not completed by due date, a deviation is required
5023 | the BGA in 08-1343 (X-1342001, 2009183) Not functioning with a Blue Screen | GARY.DODAKIAN3@SANOFI.COM | 8 NYA room 1343 | 2025-02-23 | 2025-02-23 | BGA screen works as intended and operator can use BGA to standardize and test samples
5022 | 76NYA-Biosurgery APU-Dept SEPF-Room 233-F-21880 and F-21881 Microscopes are overdue for calibration | robert.bauch@sanofi.com | 76-233 | 2025-03-01 | 2025-03-03 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
5021 | 45NYA-Thyrogen APU-Dept GCB-Room 112-FB-4531-009 and FB-4531-010 overdue for calibration | robert.bauch@sanofi.com | 45-112 | 2025-03-01 | 2025-03-03 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date | Yes | Deviation
5020 | 8NYA – Alert/Action endotoxin trend for XVF Media w/ DTT | gretchen.brunner@sanofi.com | 8 NYA | 2023-06-30 | 2025-02-28 | The media would have appropriate PPC recovery at the dilution it was validated.
5019 | Autoclave steam failure and clean area impacted | valeria.moreno@sanofi.com | 76-354 | 2025-03-03 | 2025-03-03 | Clean room following QU-SOP-0051603
5018 | 8NYA Downstream - Buffer Port Aquasyn Installation Not Performed | wilfredo.lopez@sanofi.com | 8RM1323 | 2025-03-01 | 2025-03-02 | It is expected that all Aquasyn ports are installed onto the skid during skid setup operations
5017 | 76 NYA - Possible Reveerse Osmosis breach | Craig.Stevenson@sanofi.com | 76NYA, Room 347 | 2025-03-03 | 2025-03-03 | it is expected that water not come out of the look at areas where there are no sampling ports
5016 | 8 NYA Gemba - Support Services - Manual Weigh out missing information - | nicholas.grote@sanofi.com | 08-1364 | 2025-02-07 | 2025-03-02 | Comments in the record have all information needed to perform manual transations in Athena.
5015 | 8NYA GEMBA - Support services - FIT 2009075 status was changed but was not allocated to record | JACK.LAPLANTE@SANOFI.COM | 8NYA1290 | 2025-02-05 | 2025-02-14 | Path forward
5014 | Osmometer Out Of Service | JOEL.HICKEY@SANOFI.COM | 8 NYA/Room: 1703. | 2025-03-01 | 2025-03-01 | Per QU-SOP-0084642 step 6.2.17, there must be two consecutive readings between 696-704mOsm/kg for the 700mOsm standard.
5013 | 8 NYA Gemba - Support Services - Buffer Stock Creation Amounts Incorrect | nihcolas.grote@sanofi.com | 08-1202 | 2025-03-01 | 2025-03-01 | If the pi pull comes through with an incorrect value, we need an opportunity to correct this prior to stock being created. |  | Deviation
5012 | 8NYA-FZ- EGHCE010- Missed verification of Manual Data Entry- Pi Pull Error | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1333 | 2025-01-28 | 2025-03-01 | Per SOP, Manual data entries need a verifier
5011 | 8 NYA Gemba - 3k SUM Bag (692169) Failed Inspection - 760741 Media MO- RF621F023 | Nicholas.grote@sanofi.com | 08-1350 | 2025-01-26 | 2025-01-26 | No black particles on the bag
5010 | 68NYA QC BI EMS Probe F-068-TI-09504-044 went below validated range | patrick.hoyle@sanofi.com | 68 NYA Room 4760 QC Microbiology | 2025-03-01 | 2025-03-01 | QU-SOP-0021304 Attachment 4.
5009 | Probe F-068-TI-09504-044 went out of range in incubator FB-1031 for approximately 2 hours. | patrick.hoyle@sanofi.com | 68 NYA Room 4760 QC Microbiology | 2025-03-01 | 2025-03-01 | QU-SOP-0021304 Attachment 4.
5008 | 8NYA-GEMBA Downstream - UFDF Conductivity Probes Unable To Clear During Load Membrane DV Recipe. | danny.martinez@sanofi.com | 8NYA RM1323 | 2025-03-01 | 2025-03-01 | Expected results would be to figure out why the conductivity probes aren't clearing during the Load Membrane DV Recipe.
5007 | CM pH adjust missing steps in SOP | elias.mcdonnell@sanofi.com | 8NYA, 5103 | 2025-03-01 | 2025-03-01 | Add more detail in set-up portion of SOP; i.e., connecting cord and acknowledging prompt to begin the pH adjust.
5006 | IPA packets are missing the sterility sticker | lauren.ruffing@sanofi.com | 68NYA 4700 QC Micro | 2025-02-27 | 2025-02-27 | IPA packet should have sterility sticker
5005 | 8NYA - Multiple - Daily sampling activity performed without proper OJT qualification | michael.knapp@sanofi.com | 8NYA 1333 | 2025-02-28 | 2025-02-28 | OJT training completed for required activities. | Yes | Deviation
5004 | QE-0125851 & QE-125857 - Aseptic Gowning Qualifications Action Level Excursions | daniel.whitcroft@sanofi.com | 76 New York Avenue | 2025-02-20 | 2025-02-27 | Per QU-SOP-0039519, Attachment 1, Action Level = >0 CFU/plate
5003 | 705665 Batch FGHBE001 New Batch Needed | joshua.brooks@sanofi.com | 8NYA | 2025-01-22 | 2025-01-23 | Stock is created upon goods receipt.
5002 | 8NYA GEMBA Support Services - final Breakdown step not signed for on 27FEB25; RF1AKF029 | JACK.LAPLANTE@SANOFI.COM | 8NYA 1350 | 2025-02-27 | 2025-02-28 | Path forward, non-sig event to refer to for late entry comment
5001 | Cz Zinc and Fz SP Chilled buffers out of temperature range Support Services | robert.morawski@sanofi.com | 8NYA CNC RM1210 and RM1212 | 2025-02-28 | 2025-02-28 | Chilled tote recirculation started immediately following the electrical shutdown.
5000 | 80 NYA Freezer Farm Door Found Propped Open | Craig.valery@sanofi.com | Building 80 NYA Room 150 | 2025-02-27 | 2025-02-27 | it is expected that secured access doors remain closed. | Yes | Deviation
4999 | GEMBA - 8 NYA DOWNSTREAM - Q Membrane Recovery Out Of Spec | adam.keister@sanofi.com | 8 NYA Room 1503 | 2025-02-28 | 2025-02-28 | Path forward. | Yes | Deviation
4998 | Cz Q Load supply line connected to the wrong port | edgar.vale@sanofi.com | 8NYA/Room1503 | 2025-02-28 | 2025-02-28 | Q load supply line would be connected to BV-3332-005
4997 | No Units on Printout from Analytical Balance | venessa.carloni@sanofi.com | QC Micro Lab 68NYA Room 4790 (bioburden room) | 2025-02-19 | 2025-02-27 | As per QU-SOP-0074684 Table 1, Consistency+ Unit of Measure included when handling GxP data. As per QU-SOP-0013325 Step 6.3.3. unit of measure from NIST traceable weights to be printed out on printout worksheet. As per QU-SOP-0052170 Step 6.2 unit of measure to be on printout when weighing out HA Powder in Bioburden Pour Plates.
4996 | Calibrated Clocks not used for Material Transfer | brendan.scanlon@sanofi.com | 8 NYA, 45 NYA, 51 NYA, 76 NYA | 2020-05-13 | 2025-02-26 | Calibrated clocks and or timers would be used to verify contact times
4995 | 8NYA-Gemba-Downstream SUT expired today and no other is available at 8NYA | aidan.mula@sanofi.com | 8NYA/RM1303 | 2025-02-27 | 2025-02-27 | Determination at Gemba
4994 | yellow spots on bowie dick test pack | DEVIN.VIEL@SANOFI.COM | 76 NYA room 352 | 2025-02-25 | 2025-02-27 | Per QU-SOP-0025112 step 6.3.4 the test passes if all the ink has turned blue | Yes | Deviation
4993 | 8NYA-Gemba-Downstream-QCA Alarm events from 16 and 19FEB | aidan.mula@sanofi.com | 8NYA | 2025-02-16 | 2025-02-27 | Alarms should be addressed within 72 hours
4992 | ITCSR from 2021 for conductivity temp correction on 8NYA WFI loops reverted to old values with DeltaV upgrade. This is causing high conductivity alarms that will hold up MFG. | brendan.dickhaut@sanofi.com | 8NYA | 2025-02-26 | 2025-02-26 | ITCSR is a temporary update/change. A permanent solution was needed to drive update.
4991 | 68 NYA - QC Micro - Gas Pack information not recorded | Craig.Stevenson@sanofi.com | 68NYA Room 4700 | 2025-02-19 | 2025-02-26 | It is expected that gaskpak and indicator strip information be documented at the time of use.
4990 | F076-DPI-9524-016A PVLL Urgent alarm 76NYA RM 230 TO RM 228 DP | chris.pedroli@sanofi.com | 76-230 exit airlock to 76-228 common hallway | 2025-02-25 | 2025-02-26 | MABC-QU-SOP-0070733-Alarm Response Procedure for Framingham Campus (v17.0) step 6.3 | Yes | Deviation
4989 | missing test volume MEDIA_760744_BIOBURDEN lot FGHAX002 | Stefanie.Erickson@sanofi.com | QC Micro bioburden room | 2025-02-21 | 2025-02-25 | for the sample volume to be entered prior to testing
4988 | logbook sent with vial shipment to bioreliance not returned | nicole.lufkin@sanofi.com | 45 NYA room 101 | 2025-01-09 | 2025-02-26 | QU-SOP-0070997
4987 | Missed daily ethanol PM for 24FEB25 | JAKE.STANTON@SANOFI.COM | 80NYA ethanol | 2025-02-24 | 2025-02-25 | PM to be completed daily
4986 | 4 instances of MFC failure/replacement during wave Rocker annual PM | dom.arone@sanofi.com | 8 NYA, RM 1403 | 2024-03-11 | 2025-02-25 | GFGE-QU-SOP-0069490, Calibration Program v6.0 requires failed vendor calibrations to be assessed by Quality Assurance | Yes | Deviation
4985 | 8NYA-FZ-Incorrect gowning observed during setup of Inoc Prep room | John.Lytle@sanofi.com | Rm 1333/ 8NYA | 2025-02-25 | 2025-02-25 | It is expected that operators gown according to SOP | Yes | Deviation
4984 | 8 NYA Gemba - Support Services - 08-1361 Interior MAL door will not close | nicholas.grote@sanofi.com | 08-1361 | 2025-02-25 | 2025-02-25 | MAL doors close and do not open at the same time.
4983 | 51 NYA-Thyrogen APU- Room 153- dept CC HEPA filter FB-0416 was found out of tolerance during a scheduled calibration | glen.ciano@sanofi.com | 51-153 warm room | 2025-02-21 | 2025-02-25 | QU-SOP-0069490 section 5.8.3 unit should have generated passing results | Yes | Deviation
4982 | 8NYA-GEMBA- BIRF SOP not assigned to Downstream curriculum | stanley.durand@sanofi.com | 8NYA | TBD | 2025-02-24 | SOP should be assigned to a Downstream curriculum. | Yes | No Deviation
4981 | Warehouse Disinfectant is not prepared to SOP | steven.servais@sanofi.com | 80NYA, 74NYA Warehouse (possibly 8NYA JIT warehouse) | 2025-02-23 | 2024-11-06 | Step 6.4.1 of SOP QU-SOP-0073029-Preparation of Disinfection & Cleaning Solutions
4980 | Interim Control Memo for QE-115096 not executed for UFB (GZ5156-00) EGF3B08 | Alyson.Belleville@sanofi.com | 51 - 88 | 2025-02-21 | 2025-02-24 | It is expected that the interim control memo would be executed and filled out appropriately for all required steps.
4979 | Duplicated CV (2025-76CV-0025) | alexandre.silva2@sanofi.com | 76 NYA APU | 2025-02-24 | 2025-02-24 | Per MABC-QU-SOP-0008338, Step 6.1.3, QA assigns the unique cleaning verification number.
4978 | 45 NYA Support APU-Dept QCC45-Room 1139-EMS temperature indicator FB-1464-300 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 45-1139 | 2025-02-21 | 2025-02-24 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
4977 | 51NYA- TSH- MFG Event- Aseptic Qual performed incorrectly in oneLIMS | judith.keller@sanofi.com | 51NYA | 2025-02-22 | 2025-02-24 | It is expected that samples are labeled correctly and submitted with the correct label to QC
4976 | 51 NYA APU-Room 129A-Dept APT SS-F-17195-020 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 51-129A | 2025-02-22 | 2025-02-24 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
4975 | unplanned HEPA/HVAC outage | brendan.dickhaut@sanofi.com | 8NYA | 2025-02-23 | 2025-02-23 | Engineering controls expected to be in place.
4974 | D1 agitator fell off | nicole.lufkin@sanofi.com | 51 NYA 140 | 2025-02-23 | 2025-02-23 | agitator is expected to be attached
4973 | seed train media filter not crimped | ryan.devanna@sanofi.com | 8 NYA, innoc prep | 2025-02-23 | 2025-02-23 | n/a
4972 | FBYHAM003 terminated due to SWECO motor issue | munish.kumar2@sanofi.com | 76 NYA/Room#3103/Purification | 2025-02-21 | 2025-02-21 | SWECO should be running for at least 66 hours until manually stop for the weight change (max 96hours) QU-OPE-0275728, GZ5615-10, FBYHAM003
4971 | Blue roll- up doors is nonfunctioning and is stuck in the open position. The nonfunctioning door is entrance to the JIT from a common hallway. The JIT is a badge access area. | michael.montiverdi@sanofi.com | RM 1130 JIT 8NYA | 2025-02-19 | 2025-02-19 | Door expected to close and operate properly
4970 | 8NYA - Gemba - FZ nano filter leak | lukus.palzkill@sanofi.com | 8NYA downstream suite 1 postviral room 1313 | 2025-02-20 | 2025-02-20 | x | Yes | Deviation
4969 | 8NYA - Gemba - PCS alarms on MMC record | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 room 1503 | 2025-02-17 | 2025-02-21 | PCS alarms will be addressed
4968 | 8NYA - Gemba - Disassembly step not signed | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 post viral | 2025-02-20 | 2025-02-21 | SUM and all attached SUT should be broken down and discarded | Yes | No Deviation
4967 | 8NYA - Support Service - CZ - RF608F005 760726 Completed batch exceeded 1000kg | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-02-21 | 2025-02-21 | Record should be released for max 856L
4966 | 8NYA Downstream - FGHDT006 - Discrepancies in SOP Instructions | wilfredo.lopez@sanofi.com | 8RM1323 | 2025-02-20 | 2025-02-20 | It is expected that the SOP contains the correct instruction and that not steps be missing. A revision is required to the SOP to provide corrected instructions.
4965 | 8NYA Gemba - Fabrazyme Nano - FGHDT006 - small leak observed pre-nano filter during Nanofiltration | stanley.durand@sanofi.com | 8 NYA room 1323 | 2025-02-21 | 2025-02-21 | No leaks should be observed during NanoFiltration
4964 | 8 NYA - Support Services - Particulates Found in on SUM Bag part # 692170 | nicholas.grote@sanofi.com | 08-1350 | 2025-02-20 | 2025-02-20 | No particulates on the film of the bag
4963 | Fz Seed train thaw SOP misalignment with MES | gary.dodakian3@sanofi.com | 8RM 1343 | 2025-02-20 | 2025-02-20 | Expected to scan and pass flask in prior to thaw to minimize vial culture exposure time.
4962 | Three Shipping Bags Failed Pre-Use Inspection | JOEL.HICKEY@SANOFI.COM | 8 NYA/Room: 1703. | 2025-02-20 | 2025-02-20 | Per QU-SOP-0080323 (Inspect Bag and Shipping Barrel) step 6.15.3-7, shipping bags will not have wrinkles, scratches or abnormal welds.
4961 | 8 NYA Gemba - Support Services - 1,000 liter batches exceeding 1,000 kg | nicholas.grote@sanofi.com | 08-1350 | 2025-02-20 | 2025-02-20 | Weights do not exceed calibrated range of the SUM.
4960 | QE-123041 closed prior to 3rd Party review per Protocol # QU-OPE-2722156 (Version 1.0) requirements. | irene.kuria@sanofi.com | N/A | 2025-02-19 | 2025-02-19 | Per protocol QU-OPE-2722156 (Version 1.0) Section 3, all investigations (including non-significant) created from 17FEB2025 are required to be reviewed by 3rd Party, prior to closure. | Yes | Deviation
4959 | Fz SP Column Cycles | zachary.rickard@sanofi.com | 8NYA 1313 | 2024-10-10 | 2025-02-20 | CST should have been updated to contain the correct value for cycles. | Yes | No Deviation
4958 | Conductivity Single Point Instructions | elias.mcdonnell@sanofi.com | 8NYA, 1324 | 2025-02-20 | 2025-02-20 | Single point conductivity should be listed with single point pH measurement within the chrom SOPs
4957 | 76 NYA - approved paperwork used to execute HA EM for three MFG operations | megan.botti-marino@sanofi.com | 76 NYA | 2025-02-03 | 2025-02-19 | Effective versions of paperwork should be used for test execution.
4956 | 8 NYA approved paperwork used to execute Aseptic Qualification | megan.botti-marino@sanofi.com | 8 NYA | 2025-01-31 | 2025-02-19 | Effective versions of paperwork should be used for test execution.
4955 | Cz Nano post formulation concentration too high | zachary.rickard@sanofi.com | 8NYA 1503 | 2025-02-19 | 2025-02-19 | A280 concentration will fall within the targeted range.
4954 | Fz UFDF failed NWP during CIP | zachary.rickard@sanofi.com | 8NYA/1323 | 2025-02-19 | 2025-02-19 | NWP would be passing and recipe would continue to the next step. | Yes | Deviation
4953 | 8 NYA Gemba - Support Services - Transfer setup complete but not signed for and 5 ft Extension not used for Mod Enzyme Cocktail Transfer (760739) | nicholas.grote@sanofi.com | 08-1350 | 2025-02-19 | 2025-02-19 | All steps that are executed physically are signed for in the EBR at the time of execution and SOP is clear on how to utilize all of the items that are included in the BOM.
4952 | Support Services - 1000L SUM assigned to MO RF615E013 was not the SUM used for PN 760736, Batch EGHAP013 | sheila.pandolfo@sanofi.com | 8 NYA Room 1350 (Solution Prep) | 2024-12-22 | 2025-02-18 | It is expected the batch genealogy is accurate for PN 692170, Batch 23GN02046 and the SUM used is assigned to the order. It is expected that the SUM that was rejected would be removed from the MO per MABC-QU-SOP-0032609-General Manufacturing Athena Transactions, Section 6.12.
4951 | GEMBA - 8 NYA DOWNSTREAM - Overlay Not Working Properly on CZ Nano Load SUM | adam.keister@sanofi.com | 8 NYA Room 1503 | 2025-02-19 | 2025-02-19 | For the overlay to be properly engaged for the nano load SUM.
4950 | Gemba-8NYA-Downstream CZ Zinc- multiple SUT inspection failures | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 room 1503 | 2025-02-18 | 2025-02-18 | Bag will be discarded new bag will be used
4949 | Second Verifier Needed for Manual Entry of Welder ID comment | phoebe.karkos@sanofi.com | 1333 | 2025-02-18 | 2025-02-18 | Per SOP, second verifier is required.
4948 | MCC record not able to progress | RALSTON.NAJARRO@SANOFI.COM | 8NYA1333 | 2025-02-17 | 2025-02-17 | Record should progress
4947 | X2240 PO2 rising | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-02-18 | 2025-02-18 | PO2 should stay below 150
4946 | Fz Zinc- Buffer supply depleted during recipe/ Pump did not stop automatically during recipe hold | nathaniel.barnes@sanofi.com | 8NYA 1323 | 2025-02-18 | 2025-02-18 | Determine cause of buffer depletion- buffer tote was full when setup and should not have been run dry. Also look into automation error for the pump not stopping.
4945 | 8NYA GEMBA - Downstream MMC Step not signed at time of execution. | alex.tello@sanofi.com | 8RM1212 | 2025-02-17 | 2025-02-17 | For step to be signed when the connection was made. | Yes | No Deviation
4944 | 8NYA-Gemba-Downstream Cz Phenyl- Skid device not removed after CIP | aidan.mula@sanofi.com | 8NYA/RM1503 | 2025-02-17 | 2025-02-17 | Procedure should be followed as written
4943 | X2240 DO excursion | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-02-17 | 2025-02-17 | DO should be controlled to setpoint | Yes | No Deviation
4942 | MMC fraction cannot be located | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2025-02-17 | 2025-02-17 | All MMC eluates are to be located and pooled
4941 | 51 NYA Thyrogen Purification: no available sample bag for bioburden during formulated bulk EGF2B07 | Craig.valery@sanofi.com | 51-88 | 2025-02-16 | 2025-02-16 | available, uncompromised sample bags are expected to be present | Yes | Deviation
4940 | 8NYA - Support services. Manual weigh out discrepancy found during review | JACK.LAPLANTE@SANOFI.COM | 8NYA | 2025-02-02 | 2025-02-15 | Path forward | Yes | Deviation
4939 | 8NYA Downstream - Automation Intervention Required to Turn Off Cerezyme Zinc Skid Pump | wilfredo.lopez@sanofi.com | 8RM1503 | 2025-02-14 | 2025-02-14 | It is expected that the recipe progresses through its phases with the equipment functioning as expected | No | Deviation
4938 | 8NYA Downstream - Recipe hold due to incorrect agitator position | KELVIN.GLEAN@SANOFI.COM | 8NYA 1333 | 2025-02-15 | 2025-02-15 | It is expected that the recipe parameters are achieved and the operation progresses through it's phases
4937 | 45NYA - Thyrogen - Unknown Lots - Compressed Air Dewpoint OOS | Craig.Stevenson@sanofi.com | 45 NYA Room 228 | 2025-02-14 | 2025-02-14 | It is expected that compressed air genterate results >-40C for dewpoint testing.
4936 | 8NYA - Support services, manual weigh out discrepancy on RF544F005 | JACK.LAPLANTE@SANOFI.COM | 8NYA RM1350 | 2025-02-12 | 2025-02-14 | Path forward
4935 | Fz UFDF Rinse Sampling not complete | ZACHARY.RICKARD@SANOFI.COM | 8NYA/1303 | 2025-02-13 | 2025-02-14 | All required samples will be taken. | Yes | No Deviation
4934 | CZ - RF639E017 - 8NYA - Buffer PN 760722 takeout qty error--overissued in Athena | michael.knapp@sanofi.com | 8NYA 1403 | 2025-02-13 | 2025-02-14 | Accurate QTY for buffer use is issued against BOM
4933 | GEMBA - 8 NYA DOWNSTREAM - No Sampling Line Available for NRSR Sample | adam.keister@sanofi.com | 8 NYA Room 1503 | 2025-02-14 | 2025-02-14 | Path forward on how to sample NRSR. Can we weld on to sartorius bag to take sample?
4932 | 8NYA - ALL - Gemba Request - Welding/Sealing instructions for IPA use not consistent | nicholas.grote@sanofi.com | 8RM1202, the CNC space outside of Solution Prep | 2025-02-14 | 2025-02-14 | That there would be consistency in task performance for each type of sealer, as IPA use should be consistent across platforms for contamination control and operational performance.
4931 | Wrong replacement hoses installed during Annual PM of SUB X-2250 in 08-1403 | dom.arone@sanofi.com | 8 NYA, RM 1403 | 2025-02-13 | 2025-02-13 | Like-for-Like hoses should have been kitted and installed with the work order | Yes | No Deviation
4930 | 76NYA-Biosurgery APU-Room 352-Dept HAFL-F-4507-021 HEPA filer in Grunberg oven found out of tolerance during a scheduled calibation | glen.ciano@sanofi.com | 76-352 | 2025-02-13 | 2025-02-13 | QU-SOP-0069490 section 5.8.3 unit should have generated passing results
4929 | Bubble trap lo lo sensor but butyl production recipe into hold during load | Pavel.dobrovolskiy@sanofi.com | 8nya 1323 | 2025-02-13 | 2025-02-13 | Its expected that during the load operator will pause the pump at once air is observed in the line
4928 | 8NYA GEMBA - Skid Pump tubing not found in EWI Reference info. | alex.tello@sanofi.com | 1503 | 2025-02-13 | 2025-02-13 | For pump tubing to be shown on the reference picture to know which inserts to use
4927 | 8NYA - Support Service - RF1C3F009_001 PCS alarms need assessment for CZ shipment | bailey.gleason@sanofi.com | 8NYA RM1373 | 2025-02-09 | 2025-02-13 | hold gemba to see path forward | Yes | No Deviation
4926 | 8NYA - Support Service - FIT 2009075 received external valve error twice when testing RF1CBF003 filter | bailey.gleason@sanofi.com | 8NYA RM1290 | 2025-02-13 | 2025-02-13 | hold gemba to see path forward | Yes | No Deviation
4925 | 8NYA-Gemba-Downstream Cz MMC- Pumb tubing membrane found defective during inspection | aidan.mula@sanofi.com | 8NYA/RM1503 | 2025-02-13 | 2025-02-13 | SUT should be integral for use
4924 | Disable capacitance probe due to probe divergence on X2240. | michael.yashin@sanofi.com | 8NYA, Room 1403 | 2025-02-13 | 2025-02-13 | Control strategy should be met FBL-CST-000029, probes should not diverge. |  | No Deviation
4923 | Gemba-8NYA-Downstream FZ-Particles found in SUM bag during pre-inspection | lukus.palzkill@sanofi.com | 8NYA downstream suite 1 room 1323 | 2025-02-12 | 2025-02-12 | Bag will be discarded new bag will be used
4922 | WFI drop valve 138 in 76-364 broke causing a large release of WFI to drain | Charlie.Houle@sanofi.com | 76 NYA; Room 364 (HA Fermentation) | 2025-02-12 | 2025-02-12 | WFI drop should not leak.
4921 | Leaking A280 Sample. 1670348-EGF4B09A-TSH_Q_591167_F8_A280-PI-ELISA-FRAM_A280-2 | elizabeth.endsley@sanofi.com | Sample Receipt Room 1111 at 45 NYA | 2025-02-12 | 2025-02-12 | For the cap to be on tight and not leaking | Yes | Deviation
4920 | NEO DS Solo VPE EGHDA001 ST 9M | david.chan@sanofi.com | 68 NYA Room 2700 | 2025-02-12 | 2025-02-12 | Should have no foreign matter in the sample. | Yes | Deviation
4919 | 76NYA-Support APU-BIF76-Room 352-EMS Humidity indicator F-16427-304 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-352 | 2025-02-12 | 2025-02-12 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4918 | 8NYA - Support Service - RF557F014 application lock, preventing record closure | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-02-12 | 2025-02-12 | hold gemba to see path forward
4917 | REGEMBA for EM climet data | ryan.devanna@sanofi.com | 8 NYA, innoc prep | 2024-12-24 | 2025-02-12 | n/a
4916 | 8NYA - Support Service - RF621F015 incorrect tote was scanned into record | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-01-20 | 2025-02-11 | hold gemba to see path forward | Yes | Deviation
4915 | 8NYA - Support Service - RF545F012 documentation error for manual weigh out | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-02-02 | 2025-02-11 | hold gemba to see path forward | Yes | Deviation
4914 | 51 NYA - Thyrogen: Chromatogram review temperature out of range during equilibration | craig.valery@sanofi.com | 51 NYA RM132 | 2025-01-18 | 2025-02-11 | it is expected that temperature remains with in specified range
4913 | 51 NYA - Thyrogen: Chromatogram review pH out of range during equilibration and load. v | craig.valery@sanofi.com | 51 NYA RM132 | 2025-01-22 | 2025-02-11 | it is expected that the pH remains with in specified range
4912 | 8 NYA - Support Services - 705663 Buffer SUM WT Alarm - 3,168.3 kgs | nicholas.grote@sanofi.com | 08-1350 | 2025-02-11 | 2025-02-11 | No SUM WT alarm
4911 | 8NYA Downstream - Incorrect Manifold Part Listed in BOM for Butyl Standalone CIP | wilfredo.lopez@sanofi.com | 8RM1210 | 2025-02-10 | 2025-02-10 | It is expected that the correct part numbers are listed in the batch record BOMs and the associated work instructions. A standalone CAPA was recommended by QA to reference in subsequent batch records and to capture corrective actions for EBR and SOP updates.
4910 | 8 NYA Gemba - Suport Services - Particulates on 3K Chilled tote bags | nicholas.grote@sanofi.com | 08-1202 | 2025-02-10 | 2025-02-10 | No embedded particulates on the bag. | Yes | No Deviation
4909 | Clean Steam Sampling cooler on 45 NYA Autoclave left running | stephen.peglow@sanofi.com | 45 NYA / Autoclave Mechanical Space | 2025-02-10 | 2025-02-10 | unknown QC Micro clean steam sampling SOP
4908 | GEMBA: SUT issued to multiple Process Orders but not Physically used. | william.bradley@sanofi.com | 8NYA /RM1323 | 2025-01-10 | 2025-02-10 | It is expected that Manufacturing only issues what is being used for processing.
4907 | Batched with incorrect Tote ID in EBR S705662upport Services | robert.morawski@sanofi.com | 8NYA CNC | 2025-02-10 | 2025-02-10 | 705662 EBR RF551F007 was executed with the correct Tote ID information.
4906 | CZ - RF639E017 - 8NYA - SUB X2220 increased pCO2 value resulting in SUB termination | michael.knapp@sanofi.com | 8NYA 1403 | 2025-02-10 | 2025-02-10 | O2 remains within system control
4905 | 51NYA TSH Q fraction 7 lot EGF3B08B re-sampled incorrectly by MFG | judith.keller@sanofi.com | 51NYA | 2025-02-08 | 2025-02-08 | It is expected samples are sampled correctly in oneLIMS | No | No Deviation
4904 | No Product, No Lots – 68 NYA QC Micro - Swab SOP not added to Bioburden curricula | megan.botti-marino@sanofi.com | QC Micro lab/68 NYA | 2025-02-10 | 2025-02-10 | It is expected that all required SOPs be included in their respective curricula. | Yes | No Deviation
4903 | 68NYA Support APU-Validation-Room 1745-Validator FB-5136 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 68-1745 | 2024-12-18 | 2025-02-10 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4902 | 8NYA - Support Service - RF557F020 tote transport now executed, confirmation of stock was executed | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-01-25 | 2025-02-08 | hold gemba to see path forward
4901 | Gemba 8 NYA - Support Services - Sodium Chloride-GZ1335-04 foreign material | william.lee@sanofi.com | 8 NYA Room 1350 (Solution Prep) | 2025-02-09 | 2025-02-09 | expect to observe no particulate
4900 | 8NYA-Gemba-Downstream Fz Zinc- Path forward needed for storage of skid | aidan.mula@sanofi.com | 8NYA/RM1323 | 2025-02-08 | 2025-02-08 | Determination at Gemba with cross-functional group
4899 | 8NYA GEMBA -Cz Downstream, Nano filter broken bleed valve during inspection. | alex.tello@sanofi.com | 1503 | 2025-02-08 | 2025-02-08 | Filter to be free of any damages
4898 | Insect Found in light fixture cover at Thyrogen-51NYA Purification Room 139 | Sylvester.Kegbeh@sanofi.com | 51 NYA ROOMS 136 and 139 | 2025-02-07 | 2025-02-07 | it is expected that insects should not be found inside the lab per QU-SOP-0005611 | No | Deviation
4897 | 8NYA-downstream-FZ Butyl-bubble trap filter blew off after loading | lukus.palzkill@sanofi.com | 8NYA downstream suite 1 room 1323 | 2025-02-08 | 2025-02-08 | filter should release pressure | Yes | Deviation
4896 | Various false alarms sent to remote alarm monitoring system | brendan.dickhaut@sanofi.com | 8NYA | 2025-02-08 | 2025-02-08 | False alarms are not recieved.
4895 | Fz SP- Buffer below 2-10C Temp range | nathaniel.barnes@sanofi.com | 8 NYA 1323 | 2025-02-07 | 2025-02-07 | Buffers should remain in 2-10 temp range and not overchill. Determine potential product impact and root cause of buffer overchilling.
4894 | Cedex not ran for daily sampling on x2240 (FGHBN001) | phoebe.karkos@sanofi.com | 1403 | 2025-02-07 | 2025-02-07 | MES intervention required, resampling need to run cedex | Yes | Deviation
4893 | X-2811 CIP Low buffer Alarm and WFI Expiry | Jonathan.Ahiaba@sanofi.com | 8NYA, 1403 | 2025-02-06 | 2025-02-06 | CIP will run with proper buffer flow
4892 | P/N 813903 (Meissner Dual source) was not assessed for suitability per QU-SOP-0071562 System Suitability SOP for Framingham Mass BioCampus for Cerezyme | stephanie.aceto@sanofi.com | 8 NYA, room 1703 | 2025-02-06 | 2025-02-06 | Parts being implemented into a GMP process for the first time are assessed for suitability for use per QU-SOP-0071562 prior to implementation in manufacturing. | Yes | Deviation
4891 | Unapproved Training Material uSed | gregory.tesmer@sanofi.com | 8 NYA - Room 1220 | 2025-02-05 | 2025-02-06 | All training material must be processed using a TIA and then run through an approval process where an audience is defined and the material run through a workflow
4890 | MCC production recipe failed to launch on X-2831 | dana.hopkins@sanofi.com | 8NYA Upstream Suite 2 room 1403 | 2025-02-06 | 2025-02-06 | DeltaV recipe should start when directed by MES
4889 | Quality observations of the HA Fill operations don't have End time recorded | valeria.moreno@sanofi.com | 76-360 | 2025-02-04 | 2025-02-07 | Start time and End time needs to be recorded by QA operator when observing the Fill operations
4888 | TPCN backlog without completed assessment supplier change notification | maggie.snow@sanofi.com | 68 NYA / 11 Forbes | 2022-03-22 | 2025-02-06 | Per QU-GOP-0000619 section 2.2.3 - Assess whether entity is concerned. Assess the need to create a change control within 10 working days of receipt of TPSC.
4887 | Setup NV data missing from parsing | ryan.devanna@sanofi.com | 8 NYA, innoc prep | 2024-12-24 | 2025-02-07 | n/a
4886 | DeltaV prompt "Confirm ready to transfer" not in the SOP | KELVIN.GLEAN@SANOFI.COM | 8NYA 1503 | 2025-02-07 | 2025-02-07 | Prompt description and instruction on how to acknowledge it included in the SOP
4885 | X-2831 UV Zero Error | brian.o'rourke@sanofi.com | 8NYA Room 1403 | 2025-02-06 | 2025-02-06 | UV transmitters should zero for load inlet and all the 3 columns | Yes | No Deviation
4884 | Training was documented using MABC-QU-FOR-0113823-Group Training Event Completion, to close an event listing training material that was not approved | gregory.tesmer@sanofi.com | N/A | 2025-02-05 | 2025-02-06 | Section 6.3 of QU-SOP-0069666 Training System for Specialty Care Operations Network reviews the requirements for Developing and Designing Training Materials to determine appropriate learning solutions.
4883 | 8NYA Gemba - FGHDT004 - Fabrazyme Nano - welder not identified during Post Formulation Sampling. | stanley.durand@sanofi.com | 8 NYA room 1503 | 2025-02-03 | 2025-02-05 | EBR calls for a Terumo welder to be selected for post formulation sampling.
4882 | CZ - RF639E017 - 8NYA - Unclear SUM install instructions | michael.knapp@sanofi.com | 8NYA 1403 | 2025-02-05 | 2025-02-05 | Clear instructions for SUM installation
4881 | FZ - EGHCE010 - 8NYA - Low viability of SUB X2200 | michael.knapp@sanofi.com | 8NYA 1333 | 2025-02-06 | 2025-02-06 | Viability at 80% or greater
4880 | Temperature went out of specification during reaction for lot FBYCMC001. | jan.stepanik@sanofi.com | 76 NYA room 229 | 2025-02-06 | 2025-02-06 | Temperature should have remained within specification
4879 | 8 NYA Gemba - Support Services - Scale Weigh Out Created additional HU - RF564F005 - Ongoing batch | nihcolas.grote@sanofi.com | 08-1350 | 2025-02-05 | 2025-02-05 | weigh out completed in range | Yes | Deviation
4878 | HA - FBYFRM005 - 76 NYA - Tank scale reading missing minus sign (-) | karthik.sriram@sanofi.com | 76 NYA, Room 364 | 2025-01-22 | 2025-02-05 | Scale reading in step 7.12.12.2, QU-OPE-0309296 is to be reported with minus sign as expected per step 7.12.12 in QU-OPE-0309296
4877 | CZ - RF600F003 - 8NYA - Inadequate instructions for bubble trap calibration | michael.knapp@sanofi.com | 8NYA 1403 | 2025-02-05 | 2025-02-05 | Instructions for bubble trap sensor check include replacing cap
4876 | pH out of range during equil, load, and elution steps of SP1 EGF3B12B. | Alyson.Belleville@sanofi.com | 51 NYA, 138 | 2025-01-26 | 2025-02-05 | It is expected that the pH would remain within range throughout the operation. | Yes | Deviation
4875 | 51 NYA Thyrogen - Missed removal time on BIRF | craig.valery@sanofi.com | 51-109 | 2025-01-02 | 2025-02-05 | it is expected that all applicable fields are filled in | Yes | Deviation
4874 | FZ - EGHCE010 - 8NYA - Low viability of SUB X2200 | michael.knapp@sanofi.com | 8NYA 1333 | 2025-02-05 | 2025-02-05 | Viability at 80%or greater
4873 | CZ - RF600F001 - 8NYA - CNC Monthly MFG Cleaning Not Performed | michael.knapp@sanofi.com | 8NYA 1210 | 2025-01-16 | 2025-02-04 | Cleaning is performed per procedure
4872 | Transfer Disassembly not signed for at time of occurrence. | zachary.rickard@sanofi.com | 8NYA CNC | 2025-01-31 | 2025-02-05 | Steps will be signed for at time of completion.
4871 | 68NYA Support APU-Facilities-Bore Gages F-15367 and F-15368 overdue for calibration | robert.bauch@sanofi.com | 68NYA | 2025-02-01 | 2025-02-05 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
4870 | 8NYA- Support Services - RF618F016 Mod Enzyme Flask Prep SFO - Sialidase enzyme below target | alejandro.martinez2@sanofi.com | 8NYA/8RM1350 | 2025-02-04 | 2025-02-04 | Path-forward to end record. Refer to similar event QE-099607 | No | No Deviation
4869 | Steps on FBL-EWI-00001 need to be rewritten/ Bag Installation and Inflation Steps to be Rewritten | marcelo.dossantos2@sanofi.com | 8NYA | 2025-02-04 | 2025-02-04 | Steps should allign with current operations
4868 | Incorrect Harvest hold time recorded during clarification EGF3B12B process | Luis.dominguez@sanofi.com | 51NYA room 139 | 2025-01-25 | 2025-02-04 | it is expected per QU-OPE-0351657 the correct harvest time be recorded
4867 | Missing Bio & Endo samples on LIMS | genesis.torres2@sanofi.com | 8NYA | 2025-02-04 | 2025-02-04 | Get lims labels
4866 | QCA alarms not acknowledge on time | valeria.moreno@sanofi.com | 80NYA | 2025-01-23 | 2025-02-03 | Gemba to assess impact
4865 | Butyl load filter top bleed is broken | Pavel.dobrovolskiy@sanofi.com | ya | 2025-02-04 | 2025-02-04 | its expected that filters will remain in a safe state while they wait to be tested | Yes | Deviation
4864 | Sample 1601522-FGF4B01-TSH_HARVEST_GZ1476_ELISA_ENDO-H-006 Hyclone bag contained floating red particle | johny.lisitu@sanofi.com | 68 NYA 4700 QCM | 2025-02-04 | 2025-02-04 | Expectation is that there are no particles floating inside the sample hyclone bag
4863 | Retain samples were missed during sampling step for the elution sum on Cz Zinc | patrick.mitchell@sanofi.com | 8NYA/1503 | 2025-02-03 | 2024-02-03 | Per QU-OPE-0044689, a retain sample needs to be taken along with the bioburden, endotoxin, and a280 samples.
4862 | HIHI Alarm on Cz Zinc Eluate Collection Sum During Elution | patrick.mitchell@sanofi.com | 8NYA/1503 | 2025-02-03 | 2024-02-03 | the collection sum should remain between temps of 2 & 10C and not spike causing the recipe to hold,
4861 | 8NYA-downstream-CZ zinc-recipe failed to turn off pump and could not move to next step | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 room 1503 | 2025-02-03 | 2025-02-03 | Recipe should run to completion on its own during production according to QU-SOP-0080419
4860 | MES manifold tube connection step won't allow buffer setup. | genesis.torres2@sanofi.com | 8nya | 2025-02-03 | 2025-02-03 | Allow for buffer setup to continue on MES | Yes | No Deviation
4859 | Retain sample 166273, EGHBN015 submitted in BioLIMs late needs time adjustment. | Gary.Dodakian3@sanofi.com | 8NYA | 2025-02-03 | 2025-02-03 | Per QU-OPE-1913158 adjust sample time.
4858 | Missing a witnessed by on the BIRF for the removal of FB DGF2B20 | Sylvester.Kegbeh@sanofi.com | 51 NYA ROOM 87 | 2025-01-29 | 2025-02-03 | it is expected that the Eluate be removed per QUSOP-0074092
4857 | 8NYA-CZ-EGHBN016 SUB-2250 O2 out of range, 194.4mmHg | Yves.Antoine@sanofi.com | 1403 | 2025-02-03 | 2025-02-03 | Expect to stay within range of the CST
4856 | Sample 1630987-FGFPB002-TSH_PFB1749-716645 sample tube cracked. | johny.lisitu@sanofi.com | 68 NYA 4700 QCM | 2025-02-03 | 2025-02-03 | Per QU-SOP-0070558 Step 6.2.3, confirm sample containers are not cracked, leaking. The expectation is that samples are of any cracks or leaks.
4855 | 76NYA Biosurgery APU-Room 3102 Pressure Gauge F-18956 overdue for calibration. | robert.bauch@sanofi.com | 76-3102 | 2025-02-01 | 2025-02-03 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date | Yes | Deviation
4854 | 68NYA Support APU-Dept QCC-Room 2700-Thermometer overdue for calibation. | robert.bauch@sanofi.com | 68-2700 | 2025-02-01 | 2025-02-03 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
4853 | 2024 - Cerezyme Product Quality Report Not Signed by due date of 31Jan25 | Meena.Krishnaswamy@sanofi.com | N/A | 2025-01-31 | 2025-02-03 | The final element pending i.e Deviation/CAPA Efficiency report was not signed off and available for completing the PQR and routing for signatures from SLT
4852 | 8 NYA GEMBA- SUPORT SERVICES- TOLERANCE DEVIATION , 100L BUFFER 705666 | joshua.crowley@sanofi.com | 8 NYA RM1350 | 2025-02-02 | 2025-02-02 | MABC-QU-SOP-0083660 STEP 6.11 MIXING TIMER, 6.11.1 AKNOWLEDGE PROMPT
4851 | Cz, Harvest Overflow, RF639F001, SOP instructions missing | elaine.jansen@sanofi.com | 8NYA 1403 | 2025-02-02 | 2025-02-02 | Update SOP instructions to complete operations
4850 | Recovery Percentage out of Range for Butyl Eluate | JOEL.HICKEY@SANOFI.COM | 8 NYA/Room: 1323 | 2025-02-02 | 2025-02-02 | Product loss associated with QE-177988. | Yes | Deviation
4849 | 8 NYA Gemba - Support Services - 692302 needed for Nano run was taken out against another record in November, replacemenet parts are in Q at NBO | nicholas.grote@sanofi.com | 08-1373 | 2025-02-02 | 2025-02-02 | SUT is taken out against the correct record.
4848 | CZ, ATF, RF639E016, Multiple SOP instructions to be updated for ATF changeout | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2025-02-02 | 2025-02-02 | SOP instruction to perform operation properly
4847 | 8 NYA Gemba - Support Services - Incorrectly Documented Manual Weigh out - 705688 | Nicholas.grote@sanofi.com | 08-1363 | 2025-01-31 | 2025-02-02 | Manual weigh out information is properly documented so that the correct raw material can be taken out against the record.
4846 | GEMBA - 8 NYA DOWNSTREAM - Leaks observed on CZ Zinc Skid after Blank Run | adam.keister@sanofi.com | 8 NYA 1503 | 2025-02-02 | 2025-02-02 | To investigate leak and determine path forward
4845 | 8NYA-Fz - Butyl skid bubble trap filter fell off after load was complete | KELVIN.GLEAN@SANOFI.COM | 8NYA 1323 | 2025-02-02 | 2025-02-02 | Bubble trap hi hi alarm should stop pump.
4844 | 8NYA - GEMBA Downstream - FGHDS008 - PV Mixing- Automation Intervention Agitator Not Reaching Set Point". | danny.martinez@sanofi.com | 8NYA RM1703 | 2025-02-01 | 2025-02-01 | Agitator is expected to ramp up to set point after launching the mix recipe.
4843 | 8NYA - GEMBA Downstream - FGHDS008 - Automation Intervention Agitator Not Reaching Set Point | danny.martinez@sanofi.com | 8NYA RM1703 | 2025-02-01 | 2025-02-01 | Agitator is expected to ramp up to set point after launching the mix recipe.
4842 | 8NYA-FZ-EGHCE008- Missed Verification of Manual Entry | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1210 CNC | 2025-01-10 | 2025-02-01 | Per MABC-QU-SOP-0019707-Inova Execution at 8 New York Ave (Rev 22.0), Step 6.7.2, "If manual entry is required for any step in the batch record...a second operator is required to create an exception on the step to document verification of the manual entry."
4841 | CZ - RF639F001 - 8NYA - DO Probe Fouling on SUB Train 2240 | michael.knapp@sanofi.com | 8NYA 1403 | 2025-02-01 | 2025-02-01 | Probes are representative of the culture and do not display evidence of fouling. | Yes | No Deviation
4840 | ATF failed during Visual inspection and ATF permeate SUT failure. | joseph.gonzalez@sanofi.com | 1403 | 2025-02-01 | 2025-02-01 | no particulates | Yes | Deviation
4839 | 8NYA Downstream - FGHCD001 - Cerezyme Zinc Blank Run On Hold due to Flow Failures | wilfredo.lopez@sanofi.com | 8RM1503 | 2025-01-31 | 2025-01-31 | It is expected that the recipe runs without holds | No | Deviation
4838 | 8NYA Downstream - Osmometer Out of Service due to Failed Calibration | wilfredo.lopez@sanofi.com | 8RM1503 | 2025-02-01 | 2025-02-01 | Expected that the system passes calibration checks | Yes | No Deviation
4837 | Leak found in CZ MCC buffer 832327 Train 1&2 | christian.galvez-tello@sanofi.com | 8NYA RM1210 CNC | 2025-01-31 | 2025-01-31 | N/A
4836 | 8 NYA Gemba - Membrane tear occurred Aseptiquick X connection on transfer line Buffer 832327 - RF1QJF003 | Nicholas.grote@sanofi.com | 08-1350 | 2025-01-31 | 2025-01-31 | Membrane is intact when it is pulled.
4835 | Particulate found during ATF inspection. | joseph.gonzalez@sanofi.com | 1403 | 2025-01-31 | 2025-01-31 | no particulates | Yes | Deviation
4834 | CZ, MCC, RF639F001, inefficient tube length to complete task per SOP | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2025-01-31 | 2025-01-31 | for tubing to reach the CNC space inorder to complete Buffer prime setup
4833 | 8 NYA Gemba - 760741 WFI Fill Performed without Vent Line Setup Complete | nicholas.grote@sanofi.com | 08-1350 | 2025-01-31 | 2025-01-31 | Vent line setup is performed.
4832 | Filter installed incorrectly during Cz CM Filtration setup | edgar.vale@sanofi.com | 8NYA/Room1503 | 2025-01-31 | 2025-01-31 | Filter would be installed correctly during the setup for filtration
4831 | 8NYA - Support Services - FGHAN005/RF613F005 - Insufficient EWI instruction led to automation intervention | nicholas.grote@sanofi.com | 8RM1350 | 2025-01-30 | 2025-01-30 | That instructions in an EWI/SOP will be accurate and sufficient to allow operator to perform task. | Yes | Deviation
4830 | UP-4011070 sum mixer not mixing | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1323 | 2025-01-30 | 2025-01-30 | To get a Phath forward
4829 | FZ low viability for 2210 and 2200 | phoebe.karkos@sanofi.com | 8NYA 1333 | 2025-01-26 | 2025-01-26 | Viability should be higher
4828 | 8NYA - Support Services - RF609F002 - 3000L SUM bag failed inspection due to two male ends on air line connection | bailey.gleason@sanofi.com | 8RM1350 | 2025-01-29 | 2025-01-29 | That a SUM bag will be correctly assembled to allow for use.
4827 | 8NYA Gemba - Support Services - Extension tubing leak during transfer of 760727 | bailey.gleason@sanofi.com | 8RM1202 | 2025-01-30 | 2025-01-30 | That SUT tubing will be constructed properly and not be susceptible to leaks during pressurized transfer tasks. | Yes | Deviation
4826 | Cerezyme SUM bag Faillure Q Membrane | paul.devaney@sanofi.com | 8NYA/1503 | 2025-01-29 | 2025-01-29 | Bag should have seated properly with drive mixer
4825 | Fz Butyl and SP CNC setup SOPs missing instructions for unique buffers Support Services | robert.morawski@sanofi.com | 8NYA CNC RM 1210 | 2025-01-29 | 2025-01-29 | SOP would have correct verbaige in both the Butyl and SP SOP on the unique buffer setups.
4824 | No Product, No Lots – 68 NYA QC Micro - Analyst Recorded Incorrect Pipettes for Bruker Assay | megan.botti-marino@sanofi.com | QC Micro lab/68 NYA | 2025-01-22 | 2025-01-29 | It is expected that instruments used be recorded accurately and contemporaneously. | Yes | Deviation
4823 | Buffer (705689) leaked by clamp in the UF/DF row line not set up | dennis.bazinet@sanofi.com | 8 NYA CNC space | 2025-01-29 | 2025-01-29 | No liquid after the clamp | Yes | Deviation
4822 | Port reticles printed on SUT manufactured on Thermo Automated Line, SUT in use but SCN needs to be completed | sheila.pandolfo@sanofi.com | Room 1701 | 2025-01-28 | 2025-01-28 | It is expected the SCN would have been completed for the impacted PNs prior to the impacted SUT being used in Manufacturing
4821 | 8NYA - Support Service - Post wetting recipe for CZ nano went into alarm for low low pressure twice | bailey.gleason@sanofi.com | 8NYA RM 1373 | 2025-01-29 | 2025-01-29 | Hold gemba to see path forward
4820 | FZ low viability for 2210 and 2200 | phoebe.karkos@sanofi.com | 8NYA 1333 | 2025-01-26 | 2025-01-26 | Viability should be higher
4819 | 8NYA Downstream - FGHBT002 - Steps cannot be executed per order of operations | wilfredo.lopez@sanofi.com | 8RM1323 | 2025-01-29 | 2025-01-29 | It is expected that the MES and SOP instruction steps align in order of operation execution.
4818 | 8NYA Gemba - FGHDS003 - Cerezyme - Pump Pressure spike experience during shipping bag transfer. | stanley.durand@sanofi.com | 8NYA room 1703 | 2025-01-28 | 2025-01-25 | Recipe to finish without issue
4817 | Version 6.0 of QU-SOP-0071120 approved by final authorizer in error | Amy.McIntyre@sanofi.com | QualiPSO CMS | 2025-01-28 | 2025-01-28 | Expected results would be that the final authorizer would reject the workflow per QU-WIN-0055911 (steps 6.3.4 and 6.5.4)
4816 | 8NYA-FZ- X2200 Low Viability | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2025-01-28 | 2025-01-28 | It is expected that viability remain >/= 80% during ViCell analysis
4815 | 8NYA-FZ-EGHCE008- SUB X2210 Low Viability | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2025-01-28 | 2025-01-28 | It is expected that viability remain >/= 80% during ViCell analysis
4814 | 8NYA-Cz- X2731 Unable to restart following hold due to high pressure | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2025-01-28 | 2025-01-28 | It is expected that pressure remain within allowable ranges per DeltaV alarm limits
4813 | 8NYA-CZ-EGHBN016- X2250 pH adjustment needed via Manual Sparge | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2025-01-28 | 2025-01-28 | it is expected that the pH be kept in range using manual C02 sparge as necessary per SOP
4812 | 8NYA Gemba - FGHBR009 - Cerezyme - Multiple HiHi Bubble Trap alarms during Zinc operations | stanley.durand@sanofi.com | 8NYA room 1503 | 2025-01-25 | 2025-01-25 | Follow alarm response procedure per QU-SOP-0084645.
4811 | 8NYA Gemba - FGHDS004 - Cerezyme - Nanofiltration recipe going into hold due to kink in chase buffer line | stanley.durand@sanofi.com | 8NYA room 1703 | 2025-01-28 | 2025-01-27 | Follow alarm response procedure in response to a kinked line per QU-SOP-0084645.
4810 | 8NYA-FZ-X2200 and X2210 Low Low Capacitance Controllers not enabled after meeting threshold | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2025-01-28 | 2025-01-28 | It is expected that the DeltaV alarm for Low Low CAP re-enable after returning to process allowable range |  | No Deviation
4809 | 8NYA - Support Services - 705682 (MO: RF554F003) Pre-Filtration Timer Expired | christopher.brewster@sanofi.com | 8 NYA, Rm. 1350 | 2025-01-28 | 2025-01-28 | It is expected to not have issues/unplanned outages with MES. If there wasn't an outage, this buffer would have had enough time to be completed.
4808 | No Product - No Lots - 76NYA Biosurgery, Data Integrity | Craig.Stevenson@sanofi.com | 76NYA, Room 364 | 2025-01-20 | 2025-01-27 | It is expected that the all employees complete OJT per ALCOA principals
4807 | CZ - RF639E017 - 8NYA - DO Probe Fouling, Trend Review and Probe Disabling | michael.knapp@sanofi.com | 8NYA 1333 | 2025-01-27 | 2025-01-27 | Probes are representative of the culture and do not display evidence of fouling. |  | Deviation
4806 | FZ - EGHCE008 - 8NYA - SUB 2210 Low Vaibility | michael.knapp@sanofi.com | 8NYA 1333 | 2025-01-27 | 2025-01-27 | Viability is within range | Yes | Deviation
4805 | FBYHAM001 terminated due to SWECO motor issue | munish.kumar2@sanofi.com | 76 NYA/Room#3103/Purification | 2025-01-25 | 2025-01-25 | SWECO should be running for at least 66 hours (max 96hours) QU-OPE-0275728, GZ5615-10, FBYHAM001
4804 | 8NYA Gemba - Support Services - WFI pulled before sani after outage from process air issue | arnold.christopher@gmail.com | 8RM1350 | 2025-01-26 | 2025-01-26 | That communications after a process upset would reach all team members. That team members will be careful about using Teams chat vs email as not everyone is included in Gemba chats.
4803 | 51 RO loop 1 went down due low level at 11:20am and went back 11:40. RO loop 2 went down from 10:41am to 11:40 am | jesus.acosta2@sanofi.com | 51-146 and 51-150 | 2025-01-26 | 2025-01-26 | RO #2 storage tank should have filled RO#1 tank with no issues.
4802 | 8NYA - Support Services - The filter bleed failure. | maria.palmar@sanofi.com | 8 New York Ave | 2025-01-26 | 2025-01-26 | To successful transfer of the buffer.
4801 | 8NYA - Support Services - The filter bleed failure. | maria.palmar@sanofi.com | 8 New York Ave | 2025-01-26 | 2025-01-26 | To successful transfer of the buffer.
4800 | 8NYA - Support Services - Expired tote bag | mariapalmar1810@gmail.com | 8 New York Ave | 2025-01-26 | 2025-01-26 | To successful transfer of the buffer.
4799 | 8NYA-Gemba-Downstream Cz Zinc- Automation intervention during Regen/Store was required | aidan.mula@sanofi.com | 8NYA/RM1503 | 2025-01-24 | 2025-01-24 | Record and recipe should align
4798 | pH probe signal dropping out causing holds on bioreactor. | brendan.dickhaut@sanofi.com | 8NYA 1403 | 2025-01-25 | 2025-01-25 | pH probe signal is not interrupted.
4797 | data loggers not stopped upon receipt of vials | nicole.lufkin@sanofi.com | 45 NYA | 2025-01-10 | 2025-01-24 | dataloggers stopped upon arrival
4796 | incorrect drawing rev. for 691805 in SPC (no product impact) | sarah.francis@sanofi.com | 8 NYA | 2023-07-12 | 2025-01-24 | SPCs contain the updated drawings
4795 | 8NYA - Support Service - Clarification needed in regard to pump specific operations listed under QE-097032 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-01-24 | 2025-01-24 | hold gemba to see path forward
4794 | Monitoring of HA Cell Banks Missing Paperwork | alexandre.silva2@sanofi.com | 76 NYA APU | 2024-12-20 | 2025-01-23 | Form supposed to be archived per QU-SOP-0033141-Record Management and Retention
4793 | 8NYA GEMBA - Downstream, UFDF CIP record steps misaligned | marco.vera@sanofi.com | 1323 | 2025-01-24 | 2025-01-24 | for Skid setup step to be opened at the time of launching recipe | Yes | No Deviation
4792 | POU Endotoxin sample for FGHAV019 missing. | genesis.torres2@sanofi.com | 8nya CNC | 2025-01-24 | 2025-01-24 | Locate Endo POU
4791 | Endo sample missing on lims. | genesis.torres2@sanofi.com | 8nya CNC | 2025-01-24 | 2025-01-24 | Locate Endo POU | Yes | No Deviation
4790 | Monitoring of HA Cell Banks Missing Paperwork | alexandre.silva2@sanofi.com | 76 NYA APU | 2022-05-27 | 2024-12-20 | Form supposed to be archived per QU-SOP-0033141-Record Management and Retention
4789 | 8NYA - Support Service - RF621F020 Recirculation line popped open during recirculation | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-01-24 | 2025-01-24 | hold gemba to see path forward
4788 | No Athena Stock Creation from 21JAN2025 | joshua.brooks@sanofi.com | 8NYA | 2025-01-21 | 2024-01-25 | Goods receipt in SAP from stock creation in MES or Zumerrors from failed goods receipt.
4787 | 8 NYA Gemba - Support Services - Incorrect Peristaltic Pump Used for Transfer | nicholas.grote@sanofi.com | 08-1350 | 2025-01-06 | 2025-01-23 | Pumps are used for their intended purpose.
4786 | Load volume calculated for QU-OPE- 0202442 in OneLims does not aligned with paper worksheet | nancy.watkins@sanofi.com | 45NYA, QC Chemistry | 2024-11-05 | 2025-01-23 | Calculated load volume should align with paper and OneLims
4785 | 8NYA Gemba - Support Services - Initial WFI fill deviation has automation in a loop | william.lee@sanofi.com | 8RM1350 | 2025-01-24 | 2025-01-24 | That the recipe, as designed, allows operations to proceed past the first WFI dispense knowing there is a second that will adjust accordingly to the final set point. | Yes | Deviation
4784 | CZ, RF639F001, SUB Hi-Hi temp alarm during temp stabilization | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2025-01-23 | 2025-01-23 | Temp Alarm not to be active during temp control stabilization timer | Yes | No Deviation
4783 | 51NYA GEMBA: missing witness on BIRF | adam.rowe@sanofi.com | 51NYA, room 147 | 2025-01-22 | 2025-01-23 | Barrel should have been witnessed | Yes | Deviation
4782 | 8NYA-FZ-EGHCE008- X2210 temporary weight excursion while troubleshooting sparge lines | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2024-12-28 | 2025-01-23 | It is expected that CST ranges are maintained
4781 | Technicians working in 8NYA suites without badge access Support Services | robert.morawski@sanofi.com | 8NYA CNC RM 1210 | 2025-01-23 | 2025-01-23 | Techs remained off the floor until the criteria to regain gowning access was completed.
4780 | Analyst Recorded Incorrect Pipettes for Bruker Assay | kyle.lebherz@sanofi.com | QC Micro | 2025-01-15 | 2025-01-22 | For the pipettes that were used to be recorded consistently on assay paperwork and logbook. | Yes | Deviation
4779 | Weekly alarm totalization BMS report | william.moniz@sanofi.com | All associated ENE BMS affected buildings | 2025-01-03 | 2025-01-23 | per procedure 6.8.1 weekly report is sent out along with WO generated from CMMS, EAM
4778 | 8NYA Downstream - FGHBT005 - BOM Discrepancy Between Record and SOP | wilfredo.lopez@sanofi.com | 8RM1503 | 2025-01-23 | 2025-01-23 | It's expected that the MES steps and the SOP instructions align
4777 | CZ - EGHBN015 - 8NYA - Unclear SOP instructions during ATF changeout activities | michael.knapp@sanofi.com | 8NYA 1403 | 2025-01-22 | 2025-01-22 | Instructions are located for air line connection within the ATF10 changeout section
4776 | 8 NYA Gemba - Support Serivices - Manual Weigh Transactions Not Performed | nicholas.grote@sanofi.com | 08-1364 | 2025-01-22 | 2025-01-22 | MIGO transactions are performed before records are reviewed and submitted. All required information to perform MIGO is included in the comment on the EBR.
4775 | 8NYA Hexylene Glyocol tank transmitter in high pressure state and cannot receive fill. | andrew.apel@sanofi.com | 8NYA Tank Farm | 2025-01-22 | 2025-01-22 | Pressure transmitter should be reading properly (~1 inch of Water column (inWc))
4774 | 8 NYA Gemba - SUT part # 714043 scaling up based on Batch Size | nicholas.grote@sanofi.com | 08-1350 | 2025-01-20 | 2025-01-22 | It is expected that raw materials scale up but not the SUT items.
4773 | 8NYA, viable sample data missing for EM collection batch 202412080-1 | megan.botti-marino@sanofi.com | 8 NYA | 2024-11-14 | 2024-01-02 | It is expected that all EM samples have reportable results.
4772 | Particles observed during casting of Seprafilm Lot FBYSEP003 | jan.stepanik@sanofi.com | 76-238 | 2025-01-22 | 2025-01-22 | No particles should have been visible. | Yes | Deviation
4771 | Fz Butyl CNC EBR BF Branching Change Support Services | robert.morawski@sanofi.com | 8NYA CNC RM 1210 | 2025-01-22 | 2025-01-22 | After executing the Manifold Setup step both Buffer Setup and Manifold Tube Connection should become executable.
4770 | 8NYA Gemba - FGHBR008 - Cerezyme Zinc Eluate SUM Low temperature excursion | stanley.durand@sanofi.com | 8NYA room 1503 | 2025-01-21 | 2025-01-21 | It's expected that SUM maintains set temperature 2-8 C.
4769 | 8 NYA - Support Services 692169 SUM SUT - Agitator appears bent | nicholas.grote@sanofi.com | 1350 - Solution prep | 2025-01-21 | 2025-01-21 | SUM agitator should not be bent--possibility of leak, or other such failure | Yes | Deviation
4768 | 8NYA-Gemba-Downstream- Cz Zinc Load SUM Low Temp Excursion | aidan.mula@sanofi.com | 8NYA/RM1503 | 2025-01-21 | 2025-01-21 | Chilled water should not exceed defined temperatures of 2-8C
4767 | Filter tested using the incorrect program during UF/DF EGF2B07 process | Sylvester.Kegbeh@sanofi.com | 51 NYA ROOM 88 | 2025-01-21 | 2025-01-21 | It is expected per QU-OPE-0526918 Step 4.2.1.5 that the filter be tested using program 1083-01 per QU-SOP-0072849
4766 | 8NYA - Support Service - RF573F001 incorrect extension used for transfer | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-01-21 | 2025-01-21 | hold gemba to see path forward
4765 | work order 3203988 required daily readings of humidity and temperature on 74NYA humidor; readings missed 01DEC24 | Janet.langon@sanofi.com | 74CCE-UT1-02 | 2024-12-01 | 2025-01-21 | Reading recorded
4764 | TSK EGF4B09A Percent recovery out of the range (66%) of 69%-112% | Sylvester.Kegbeh@sanofi.com | 51 NYA ROOM 132 | 2025-01-13 | 2025-01-21 | It is expected per QU-OPE-0295249 step 4.9.19 that the percent recovery is within the of range 69%-112%
4763 | 8NYA - Support Service - P/N 713972 HU 1001351752 documented in 2 different records | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-01-21 | 2025-01-21 | hold gemba to see path forward
4762 | X2240 Unable to get 3 tabs attached during SUB setup | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-01-20 | 2025-01-20 | Operators should be able to attach a minimum of three tabs
4761 | 8NYA GEMBA - Downstream, CZ Shipping barrels pump HI pressure | alex.tello@sanofi.com | 1703 | 2025-01-20 | 2025-01-20 | for recipe to hit SUM target weight and complete recipe. | Yes
4760 | 8NYA - Support Service - RF607F003 / 760725 PH is OOS | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-01-20 | 2025-01-20 | hold gemba to see path forward
4759 | Unknown product - Multiple Buildings - Critical Utility NitratesChP | Craig.stevenson@sanofi.com | TBD | 2024-12-30 | 2025-01-20 | It is expected that all annual sampling required per the ChP, generate passing results | Yes | Deviation
4758 | Gemba needed for 8 NYA to discuss Action notice 2025-I-001. | kenny.leland@sanofi.com | 8 NYA 1200, 1201, 1210, and 1211 | 2025-01-16 | 2025-01-19 | Aseptic connections should not be made until AHU work order 3298769 has been completed.
4757 | 8NYA-Fz-FGHSD001- SOP and MES do not align for Wave 2 Operations, placing Production activities on hold | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2025-01-19 | 2025-01-19 | it is expected that the SOP and Batch record order of operations align for correct execution of production activities
4756 | pH for SP1 EGF2B11A out of range (low) during equil through elution | Alyson.Belleville@sanofi.com | 51-138 | 2024-12-28 | 2025-01-18 | It is expected that the pH would remain within range throughout the operation. | No | Deviation
4755 | 8NYA - GEMBA Downstream Cz- 2 CM Retain Eluate Samples Leaking At Sample Pickup | danny.martinez@sanofi.com | 8NYA | 2025-01-18 | 2025-01-18 | Path forward to re-sample the 2 CM eluate retain samples
4754 | 8NYA Support services - "discharge tubing from the SUM to the pump pancaking under transfer vacuum" | JACK.LAPLANTE@SANOFI.COM | 8NYA RM1350 | 2025-01-10 | 2025-01-11 | Path forward | Yes | No Deviation
4753 | 8NYA Downstream - FGHBP007 - Incorrect Filter Test CX Value in Parent MO Parameter Code | wilfredo.lopez@sanofi.com | 8RM1350 | 2025-01-18 | 2025-01-18 | It is expected that the correct parameters be written into the batch record code so the correct CX values populate into the subsequent operations.
4752 | 8NYA - Incorrect amount on P/N: 692302 issued to the record RF525F002. | JACK.LAPLANTE@SANOFI.COM | 8NYA | 2025-01-15 | 2025-01-17 | Path forward
4751 | Personnel improperly gowned | michelle.smolinsky@sanofi.com | 11 Forbes Road, Room 1204A | 2024-12-23 | 2025-01-17 | Personnel would be properly gowned in all areas
4750 | DO probes divergence for DO probe | drew.Hopkins@sanofi.com | 8 NYA / 1403 | 2025-01-17 | 2025-01-17 | Probe 1 and 2 should not reading same value and control the SUB based on two probes reading average | Yes | No Deviation
4749 | Effectiveness check EFF-004837 has no criteria, CAPA action is closed and completion of the effectiveness check cannot be completed | james.hubbard@sanofi.com | 8 NYA | 2025-01-16 | 2025-01-16 | EC would have critera and would be able to close on time
4748 | FZ- RF503E010 - 8NYA - 8NYA MCC 2510 Load Pressure Indicator Intermittent ModBad | michael.knapp@sanofi.com | 8NYA 1333 | 2025-01-17 | 2025-01-17 | Functional Pressure Indicator | Yes | No Deviation
4747 | 8NYA - GEMBA Downstream, FZ Butyl CIP Record unable to move forward. | alex.tello@sanofi.com | 1323 | 2025-01-17 | 2025-01-17 | For the record to open up provisional release section.
4746 | automation intervene needed for sub 2210x to disable harvest sum low weight limit to allow drain and replace with new bag. | joseph.gonzalez@sanofi.com | 1333 | 2025-01-17 | 2025-01-17 | low weight limit alarm disabled |  | No Deviation
4745 | 8 NYA Support Services - 3K tote bag particle inside plastic | nicholas.grote@sanofi.com | 8 NYA -CNC Tote filling area | 2025-01-17 | 2025-01-17 | FBL=EWI-000015 Step 4.4.3.1.3
4744 | 8 NYA Support Services - 1k Tote bag debris - inside outer packaging | nicholas.grote@sanofi.com | 8 NYA -CNC Tote filling area | 2025-01-17 | 2025-01-17 | FBL=EWI-000015 Step 4.4.3.1.3
4743 | 8NYA - MFG ALL - Supply fan supporting that AHU is down at 8NYA | bailey.gleason@sanofi.com | 8NYA | 2025-01-16 | 2025-01-16 | N/A
4742 | 8NYA GEMBA - Downstream, FZ- No Waste stream reference for in SOP for FZ Mannitol | alex.tello@sanofi.com | 1323 | 2025-01-16 | 2025-01-16 | to have a reference to where to drain the buffer
4741 | Two SUM bags found with defective propellers inside | pavel.dobrovolskiy@sanofi.com | 8NYA 1503 | 2025-01-16 | 2025-01-16 | its expected that SUM bag driver unit come from the vendor assembled correctly
4740 | No Product - No Lots - 76NYA - Room breach during active Action Notice | Craig.stevenson@sanofi.com | 76 NYA Room 421 and 381 | 2025-01-15 | 2025-01-15 | It is expected that correct gowning be used while in classified areas during Action Notices.
4739 | FZ nano SOP is missing steps for load SUM disassembly | pavel.dobrovolskiy@sanofi.com | 8NYA 1333 | 2025-01-16 | 2025-01-16 | its expected that the SOP will provide the operator with the guidlines for removal and discard of all SUT
4738 | 8 NYA - Support Service - L - Glutamine P/N 726192 was left in cold room 2-10C overnight, Storage condition should be 15-30C | bailey.gleason@sanofi.com | 8NYA RM1220 | 2025-01-16 | 2025-01-16 | N/A
4737 | Cz MMC- Eluate Storage SUM: Low Temp Excursion | nathaniel.barnes@sanofi.com | 8NYA / 1503/ UP-4103050 | 2025-01-16 | 2025-01-16 | Determine product impact from temp excursion
4736 | 8NYA - Support Services - 3k tote bag incorrect (tail points out of tote) | william.lee@sanofi.com | 8RM1202 | 2025-01-16 | 2025-01-16 | That the tote bag will be manufactured correctly.
4735 | MMV retain vial cracked on receipt to QC: EGHBN015-CB-2 Retain SUB 500L_Capto Butyl_MMV 10ML Retain (Lot: 61126, sample 166248) | phoebe.karkos@sanofi.com | QC | 2025-01-15 | 2025-01-15 | 10ML MMV retain
4734 | Cz MCC SOP missing Daily checklist instructions | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-01-15 | 2025-01-15 | SOP should have instructions on how operators should complete daily checklists
4733 | 8NYA- Gemba - downstream - Q-Membrane SOP issues | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 room 1503 | NA | NA | Changes will be made to the SOP
4732 | Climet data not parsed onto tablet | ryan.devanna@sanofi.com | 8 NYA, innoc prep | 2025-01-10 | 2025-01-15 | n/a
4731 | 8NYA-Gemba-Downstream Pompe ANX Repack record calls for incorrect resin quantity | aidan.mula@sanofi.com | 8NYA/RM1303 | 2025-01-15 | 2025-01-15 | BOM should be accurate
4730 | 8NYA Downstream - FGHBR001 - Additional Post pH Adjustment Sampling/Testing Performed without Gemba | wilfredo.lopez@sanofi.com | 8RM1503 | 2025-01-05 | 2025-01-05 | It is expected that a preliminary Gemba is held prior to performing a resampling and retesting operation. It is expected that all results are recorded into the batch record accordingly.
4729 | 2024 Quarterly Walkthrough not Completed by Facilities (F23-016 Major 2 see attached) | mabc_EandM_compliance@sanofi.com | 68NYA | 2024-03-31 | 2025-01-15 | QU-SOP-0071975 section 6.3 (entire section for not completing steps)
4728 | Late PM (WO 3281726) Original Due Date 31Dec2024 for not completing quarterly walkthrough | mabc_EandM_compliance@sanofi.com | 68NYA | 2025-01-01 | 2025-01-14 | All PMs completed on time by original due date or, for eligible PMs extend on-time
4727 | Mutiple overdue actions for Internal and Supplier audits | Nashmin.Hojatti@sanofi.com | 68nya | 2024-02-15 | 2025-01-14 | Internal and Supplier audit reports would be completed per the timeframes listed in QU-GOP-0000618 & QU-GOP-0000733. 2025 internal audit risk assessment would be completed prior to 2025
4726 | 8NYA - GEMBA Downstream, Record for Zinc Skid has no Provisional release section. | alex.tello@sanofi.com | 1323 | 2025-01-15 | 2025-01-15 | For the record to have Provisional release section.
4725 | 8NYA Downstream - FGHBE002 - Automation Intervention Required During CV Rinse Sample Collection | wilfredo.lopez@sanofi.com | 8RM1323 | 2025-01-15 | 2025-01-15 | Expected that enough flow and back pressure exists in the system to be able to collect the required samples. Assessment needed to determine if this is a one-time occurrence or if recipe parameters need to be updated.
4724 | Quarterly PM Walkthrough not completed by due date | david.zhong@sanofi.com | 68 NYA 1745 | 2025-01-01 | 2025-01-14 | Quarterly PM Walkthrough completed by due date 31DEC24 | Yes | Deviation
4723 | 76 NYA WFI pumps faulted | john.mendenhall@sanofi.com | 76 NYA, room 347 | 2025-01-14 | 2025-01-14 | WFI pumps would stay running
4722 | Alignment issue with QU-SOP-0083844 and Butyl Load Adjust EBR | katherine.reardon@sanofi.com | 8 NYA 1323 | 2025-01-14 | 2025-01-14 | work towards alignment in procedural documents
4721 | Re-Sampling incorrectly performed in OneLIMS | timothy.ryley@sanofi.com | 68 | 2025-01-01 | 2025-01-14 | Per section 24.1 "The user ensures that the original sample is properly dispositioned as rejected or cancelled." | Yes | Deviation
4720 | An empty bag of chips was found in the men's outer gowning area | valeria.moreno@sanofi.com | 76-204 | 2025-01-13 | 2025-01-13 | Disinfect area, and notify supervisor
4719 | Multiple missing signature steps in Fz SP and Zinc Records | zachary.rickard@sanofi.com | 8NYA/1323 | 2025-01-11 | 2025-01-13 | Steps will be signed for at time of completion and any necessary comments will be added.
4718 | 8NYA Downstream - Fabrazyme SOPs - Incorrect Instructions Provided for CNC Breakdown | wilfredo.lopez@sanofi.com | 8RM1210 | 2025-01-14 | 2025-01-14 | It is expected that SOP instructions are accurate for the process and operation being performed. Document revision required to properly carry out the CNC breakdown operation. Other Fabrazyme SOPs must be assessed for alignment as well.
4717 | CZ- EGHBN017/EGHBGN016 - 8NYA - First Harvest Day Sampling Error (wrong record used) | michael.knapp@sanofi.com | 8NYA 1403 | 2025-01-11 | 2025-01-14 | Sampling is performed and documented on correct systems and records.
4716 | 8NYA Downstream - FGHBE002 - SUT Takeout Step Missing in MES Record for WFI Setup | wilfredo.lopez@sanofi.com | 8RM1210 | 2025-01-14 | 2025-01-14 | Manufacturing personnel expect the records to have takeout steps to perform transactions of the required specific materials to complete the operation. Gemba to be held to update the record to include the missing step.
4715 | 8NYA - Support Services - FIT - Ext valve Error - 762736 / RF1CBF001 | alejandro.martinez2@sanofi.com | 8NYA/8RM1290 | 2025-01-13 | 2025-01-13 | Filter integrity tester launches FIT recipes with no issues.
4714 | 8NYA - Support Services - Leaking Tote Bag - 705673 / RF550F001 | alejandro.martinez2@sanofi.com | 8NYA/CNC Tote Filling station | 2025-01-13 | 2025-01-13 | Tote bag should remain integral and not leak.
4713 | Re-Gemba for QC sample duplicates | Milva.Walsh@sanofi.com | 68 NYA | 2025-01-13 | 2025-01-13 | It is expected that there are not duplicate samples in the OneLims system.
4712 | 8NYA-Cz EGHBN014- Harvest volume too low to perform Eluate Day sampling per SOP | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2025-01-13 | 2025-01-13 | It is expected to have adequate volume of Harvest material to sample per Sampling SOP
4711 | 8 NYA Gemba - Support Services - Incorrect Tote Scanned into record during stock creation for 760720 | nicholas.grote@sanofi.com | 08-1202 | 2025-01-06 | 2025-01-13 | SU is sent into the correct tote when performing stock creation.
4710 | 8NYA Support Services - Incorrect Aseptic connecting made with transfer line | JACK.LAPLANTE@SANOFI.COM | 8NYA CNC | 2025-01-13 | 2025-01-13 | Looking for path foward.
4709 | 8NYA Upstream Empty Endotoxin Sample Submitted to QCM (EGHCE008, p/n 706930) | abel.agnes@sanofi.com | 68NYA 4700 | 2025-01-12 | 2025-01-13 | Sample vials contain sufficient sample volume to proceed with analysis.
4708 | 8NYA=Fz EGHCE008 MCC X2543 Leak identified at Elution Buffer Inlet port to skid | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2025-01-13 | 2025-01-13 | It is expected that the MCC remain integral and free of leaks while production activities are ongoing
4707 | GEMBA - 8 NYA DOWNSTREAM - UFDF Recipe Alarmed Due to Retentate Flow Controller Low Low Alarm | adam.keister@sanofi.com | 8 NYA room 1323 | 2025-01-13 | 2025-01-13 | For recipe to finish and pump to reach SP as expected.
4706 | X-2220 ATF Guard Filter Bleeds Left Open | brian.o'rourke@sanofi.com | 8NYA 1403 Train 1 | 2025-01-06 | 2025-01-13 | Flowpath walkdown should occur prior to perfusion start. Daily checklist walkdown could have caught this earlier as well.
4705 | Failed Conductivity Result for HA Fermentation Filter Housing 2025-76CV-0006 at 76 NYA | dianne.moustafa@sanofi.com | 76 NYA Room 364 Fermentation | 2025-01-11 | 2025-01-11 | Expected to have passing conductivity results
4704 | 8NYA-Gemba-Downstream Fz Zinc CV Rinse sample was found leaking | aidan.mula@sanofi.com | 8NYA/RM1220 | 2025-01-12 | 2025-01-12 | Samples should remain integral at all times
4703 | X2811 Leak on HiBOD waste filter outlet | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-01-12 | 2025-01-12 | Filter should push to drain without leaking
4702 | 8NYA-DOWNSTREAM-FZ-UF/DF-FGHUP001- Leak observed on manifold in CNC | Marc.Latta@sanofi.com | 1323 | 2025-01-12 | 2025-01-12 | no leak should occur per process
4701 | 8NYA-DOWNSTREAM-FZ-UF/DF-FGHUP001-SETUP RECIPE-LEAK OBSERVED ON THE NOVASEPTUM CONNECTION POINT ON THE PERMEATE LINE | Marc.Latta@sanofi.com | 1323 | 2025-01-11 | 2025-01-11 | no leak should occur per process
4700 | 8NYA Gemba - Support Services - late entry for setup | william.lee@sanofi.com | 8NYA/1350 | 2025-01-11 | 2025-01-10 | Step for setup should be signed for at time of bag setup.
4699 | Gemba - 8NYA - Downstream - Freeze record misaligned with SOP | lukus.palzkill@sanofi.com | 8NYA freezer farm | 2025-01-11 | 2025-01-11 | Record should be modified so that there is a step to sign for the storage of drug product in the -80 freezers
4698 | 8NYA - GEMBA - Downstream, UFDF Discrepancy on SOP | alex.tello@sanofi.com | 1323 | 2025-01-11 | 2025-01-11 | for step in MES to be after skid setup. For SOP to not mention a valve? | Yes | Deviation
4697 | 8NYA GEMBA - Downstream, Bag underfilled during freeze ops. | alex.tello@sanofi.com | 1703 | 2025-01-10 | 2025-01-10 | For bag to fill to minimum of 7.3kg | Yes | Deviation
4696 | TSH- 45 NYA - Buffer prep - multiple DP alarms occurred on 10JAN25 for several areas at 45 NYA. | jeremy.wooden@sanofi.com | 45 NYA (186 to 172, 186 to 170, 188 to 186, 190 to 188, and 189 to 186) | 2025-01-10 | 2025-01-11 | It is expected that there would be no DP alarms impacting 45 NYA. | No | Deviation
4695 | 8NYA Gemba - Downstream - Osmometer taken out of service after second failed standardization attempt | aidan.mula@sanofi.com | 8RM1503, | 2025-01-11 | 2025-01-11 | That any standardization will result in a successful attempt.
4694 | 8NYA-DOWNSTREAM-CZ-PHENYL-FGHBT002- SKID PUMP NOT REACHING THE FLOW RATE | Marc.Latta@sanofi.com | 1503 | 2025-01-11 | 2025-01-11 | pump must reach the flow rate defined in the recipe | Yes | Deviation
4693 | 8NYA Gemba - Support Services - 3k SUM bag rejected during inspection due to particulate | william.lee@sanofi.com | 8RM1350. | 2025-01-11 | 2025-01-11 | That all SUM bags will not have observable particulates. | Yes | Deviation
4692 | Tubing installed incorrectly for MCC skid buffer tote 1602001/Valve BV-1729007 | brendan.dickhaut@sanofi.com | 8NYA 1403 | 2025-01-10 | 2025-01-10 | CNC lines from the supply totes to the MCC skid are installed correctly for proper flow path.
4691 | HOF recipe going into hold due to high pressure | edgar.vale@sanofi.com | 8NYA/1365 | 2025-01-10 | 2025-01-10 | No holds during freeze process
4690 | 8NYA-Gemba-Downstream-SP record does not have pH adjust disassembly step | aidan.mula@sanofi.com | 8NYA/RM1323 | 2025-01-10 | 2025-01-10 | Record should have all steps required to perform operation | Yes | Deviation
4689 | EFF-004545 found Ineffective; SOP requirements not met. | Catherine.Collins@Sanofi.com | EC Scope was 8,45,51 & 76NYA MFG areas, on a rotating schedule. | 2024-11-22 | 2025-01-09 | Step 6.3.2 States to document remediations in the checklist/comments. Sec.6.5.2 Weekly WT are done by MFG Management and Quality Operations. Step 6.5.5 WT is documented following a team debrief following the WT. Step 6.5.6 NOTE Holidays & Shut Down are exceptions and WT will occur when area(s) are back in service.
4688 | Cleaning verification option no longer on 51NYA parts washer recipes | nicholas.shoenfelt@sanofi.com | 51-142 | Unknown | 2025-01-10 | It is expected that the parts washer recipe will prompt operator with the option for a CV, which will hold the recipe to at the final rinse to facilitate rinse water sampling. | Yes | Deviation
4687 | Gemba 8 NYA Support Services 760737 MO RF616F001 incorrect procedure | william.lee@sanofi.com | 8NYA/1350 | 2025-01-10 | 2025-01-10 | Bottle should not be used for 760737 batch for liquid weigh out.
4686 | Failure to follow GDP for material activation in Athena | Nashmin.Hojatti@sanofi.com | n/a | 2025-01-09 | 2025-01-09 | QU-SOP-0074684, step 6.11.1: Personnel will not review, verify or witness a step they have performed
4685 | 8NYA - Support Service - Step 6.9.1.9 in QU-SOP-0083660 wording does not reflect pre weighed items. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-01-10 | 2025-01-10 | N/A
4684 | Missed sample weight slip for pH assay | vincent.palumbo@sanofi.com | 11 Forbes | 2024-11-20 | 2025-01-10 | It is expected that any sample weighed out would have a balance weight slip for traceability
4683 | Clean Steam generator #2 MECO down due Low pressure plant steam alarm. | jesus.acosta2@sanofi.com | 76-347 | 2025-01-07 | 2025-01-07 | Clean steam generator#2 should be running with no alarms, it should get normal plant steam pressure with no issues.
4682 | Climet FB4138 not in MES | ryan.devanna@sanofi.com | 8NYA innoc prep | 2025-01-09 | 2025-01-09 | n/a | Yes | No Deviation
4681 | 8NYA Gemba - Support Services - 3K SUM Bag Leak During Initial WFI - 705665/RF544F004 | alejandro.martinez2@sanofi.com | 8NYA/8RM1350 | 2025-01-09 | 2025-01-09 | No leaks occur, SUM bag is integral
4680 | 8NYA Gemba request - Cerezyme Thaw/Pool EGHDT012 - The second filter is noted as being used in the record but does not appear to have been transacted in SAP. | stanley.durand@sanofi.com | 8NYA room 1703 | 2024-05-24 | 2025-01-09 | It's expected that all SUTused in a record to be transacted against process order in SAP.
4679 | Double amber clamps observed- BHT-2042002, Part #760741, MO#: RF621F001, Chemical Name: XVF Media with DTT Expiration Date: 21JAN25, Production CZ (Note: 2 sets of double amber clamps on hoses) Double amber clamps Observed - Part #760744, Batch #: EGHBE024, Chemical Name: XVF Media (Note: 1 set double amber clamps) | randall.porcella@sanofi.com | 8NYA Logistics Corridor 1210 (CNC space) | 2025-01-09 | 2025-01-09 | Effectiveness Reviews would be created
4678 | 8 NYA Gemba - Support Services - Batching Samples Missed - RF603F001 - 760720 | nicholas.grote@sanofi.com | 08-1202 | 2025-01-06 | 2025-01-09 | Product is selected in BioLims (Cz or ASM) which will generate the samples. Then the operators will proceed to take the correct bacthing samples and submit to QC. | Yes | Deviation
4677 | 8 NYA Gemba - Support Services - Incorrect Media Record Versions used in MES | nicholas.grote@sanofi.com | 08-1350 | 2025-01-09 | 2025-01-08 | The effective record version is used for execution. | Yes | Deviation
4676 | Issues with QU-SOP-0080419 during RF643F003 processing | katherine.reardon@sanofi.com | 1503 | 2025-01-09 | 2025-01-09 | work towards alignment in procedural documents
4675 | Cerezyme and Fabrazyme Control Strategy (CST) Process Flow Diagram incorrect | gregory.klotz@sanofi.com | 8 NYA | NA | 2025-01-09 | It is expected that the process flow diagram accurately reflects the process steps for downstream manufacturing.
4674 | Cz Zinc pH bags leaking | zachary.rickard@sanofi.com | 8NYA 1503 | 2025-01-08 | 2025-01-08 | SUT should not be defective. | Yes | Deviation
4673 | 8NYA Downstream - Chromatography Operations EWI Instructions Not Accurate for Standalone CIP | wilfredo.lopez@sanofi.com | 8RM1323 | 2025-01-08 | 2025-01-08 | It's expected that the procedure contains only instructions for steps that can be executed during operations.
4672 | 68NYA Support APU-Dept QCM-room 4700-Thermometer F-16480 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-4700 | 2025-01-08 | 2025-01-09 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
4671 | 8NYA Downstream - FGHBR003 - Discrepancies in SOP Instructions | wilfredo.lopez@sanofi.com | 8RM1503 | 2025-01-08 | 2025-01-08 | It's expected that the procedure documents include all instructions for steps to be performed including references. | Yes | No Deviation
4670 | 8NYA Downstream - Flexsafe Mixer bag with leaking pH probe | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2025-01-09 | 2025-01-09 | pH housing is not supposed to be leaky
4669 | 8NYA Downstream - CNC Manifold Tube Connection Step Signed Prematurely | elias.mcdonnell@sanofi.com | 8NYA, 1503/phase 2 CNC | 2025-01-08 | 2025-01-08 | Step must be signed once the task is performed.
4668 | 8NYA Downstream - SUM Bag Leak in pH Sensor Housing Plastic | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2025-01-08 | 2025-01-08 | No leak in any components of the pH sensor housing (or leaks in general)
4667 | No Product - No Lots - 8NYA - Upstream and Downstream - No Training Completed for MABC-QU-SOP-0079722 | julianna.macconnell@sanofi.com | 8NYA | 2025-01-08 | 2025-01-08 | It is expected for procedure to be assigned and training completed.
4666 | 8 NYA GEMBA Request- Discrepancy between the SUM FS and the code around pH probe High-High and Low-Low alarms | richard.manasian@sanofi.com | N/A-Automation | 2025-01-08 | 2025-01-08 | alarms would have been set as critical | Yes | No Deviation
4665 | Incorrect Version of BIRF QU-SOP-0074092 Attachment 1 Used for Pooled Formulated Bulk EGFPB009 | Alyson.Belleville@sanofi.com | 51 NYA, RM 87 | 2025-01-02 | 2025-01-08 | It is expected that the current effective version of a document would be used at time of execution.
4664 | 8 NYA - Support Services - XVF Media and DTT stored in 6-10 Storage | nicholas.grote@sanofi.com | 08-1420 | 2025-01-08 | 2025-01-08 | XVF Media (758892) and DTT (767677) are stored in the 2-8 cold room. | Yes | Deviation
4663 | Lost of clean steam generators and loops for 76NYA and 8NYAdue to 55 NYA CUB loss of Plant steam pressure due to loss of water in the surge tank. (this event started at about midnight) Facilities received alarm of a low pressure to the clean steam . Alarm was traced back to the high pressure boilers in 55NYA. | MACJONES.OSEIKUFFOUR@SANOFI.COM | 8NYA, ROOM 1170 MECHANICAL ROOM LOST CLEAN STEAM PRESSURE. 45nya This building was close enough to the CUB to not lose any equipment. 51nya This building was close enough to the CUB to not lose any equipment | 2025-01-08 | 2025-01-08 | MABC-QU-SOP-0018877 and MABC-QU-SOP-0070529
4662 | 8NYA - Support service - 705688 EGHBY010 time of batching samples missed. Record needed for shipment. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-10-26 | 2025-01-07 | N/A | Yes | Deviation
4661 | CU not in scope of QU-SOP-0070272-FM - Management of Action Notice Notification within Biologics R&D | Craig.stevenson@sanofi.com | 76 NYA | 2024-12-17 | 2025-01-06 | It is expected that all CU A/Ns be routed though QU-SOP-0022403
4660 | 8NYA-Gemba-Downstream SUM Bag had fraying on aseptiquik tab noticed during installation | aidan.mula@sanofi.com | 8NYA/RM1503 | 2025-01-06 | 2025-01-06 | Bag should remain integral
4659 | 8NYA-Gemba-Downstream Pompe ANX Column Packing Rinse Nozzle hose could not be located for skid setup | aidan.mula@sanofi.com | 8RM1303 | 2025-01-06 | 2025-01-06 | Determination at Gemba
4658 | ServiceNOW MES change proceeded without QA pre-approval | peter.jansen@sanofi.com | 68NYA | 2024-11-22 | 2025-01-06 | QA pre-approval would be performed before work proceeded.
4657 | X2240 pCO2 levels rising | RALSTON.NAJARRO@SANOFI.COM | 8NYA1403 | 2025-01-06 | 2025-01-06 | pCO2 levels should be stable |  | No Deviation
4656 | Gemba - 8NYA - Downstream - MMC A280 RSD value out of range | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 | 2025-01-04 | 2025-01-06 | RSD should be <=3%
4655 | 8NYA - Support service - Discussion regarding H-glycol, and E-glycol after the absence of N2 overlay | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-01-06 | 2025-01-06 | hold gemba to see path forward
4654 | Unqualified GZ-34 Control used within QC Chemistry QU-OPE-0045741 assay | aimee.park@sanofi.com | 45NYA, QC Chemistry Room 1136 | 2025-01-02 | 2025-01-03 | The GZ-34 requalification would have been fully approved extending the expiration date before use within QU-OPE-0045741.
4653 | Non-contemporaneous entries within QC Chemistry Cerezyme Activity, QU-OPE-0011377 Onelims assay | aimee.park@sanofi.com | 68NYA, QC Chemistry laboratory | 2024-12-14 | 2025-01-03 | All Onelims components within batch FBL087-TEC-00199 documented on 13Dec24 would have been reentered on 14Dec24 when the assay initiation occurred.
4652 | 8 NYA Gemba - Support Services - Alarm during Nano Flush Ramp Down | nicholas.grote@sanofi.com | 08-1323 | 2025-01-05 | 2025-01-05 | Recipe completes with no alarms
4651 | CZ- RF639E017 - 8NYA - Failed weld identified on SUB bleed line | michael.knapp@sanofi.com | 8NYA 1403 | 2025-01-05 | 2025-01-05 | Bleed line is not compromised. | Yes | Deviation
4650 | Column Packing CNC Setup Steps Missing | zachary.rickard@sanofi.com | 8NYA/CNC | N/A | 2025-01-05 | Steps needed to perform a task will be documented in an effective EWI or SOP.
4649 | 8 NYA Gemba - Support Services - Data Entry Error During Enzyme Value Entry - Record not used Aborted | Nicholas.grote@sanofi.com | 08-1202 | 2025-01-05 | 2025-01-05 | Bottles that are removed for thaw are used and do not need to return to the freezer.
4648 | CZ post viral SUM mixing recipe not functioning | zachary.rickard@sanofi.com | 8NYA/1703 | 2025-01-05 | 2025-01-05 | Recipe will function as intended.
4647 | Thyrogen bioburden sample 1603745 dropped and cracked before testing | Stefanie.Erickson@sanofi.com | QC Micro bioburden room | 2025-01-05 | 2025-01-05 | for the sample to be sterile prior to testing
4646 | No luerlock sample syringes available to aliquot | Jacob.Roman@sanofi.com | 8RM1403 | 2025-01-04 | 2025-01-04 | Expectation is that samples are aliquoted using syringes per QU-WIN-0046060, single use bag sampling, 4. Aliquot Sample | Yes | No Deviation
4645 | Issues with QU-SOP-0080418 | katherine.reardon@sanofi.com | 8 NYA, 1503 | 2025-01-04 | 2025-01-04 | fix document to be able to operate as written | Yes | Deviation
4644 | Skid Disassembly Step Not Signed For | zachary.rickard@sanofi.com | 8NYA 1503 | 2025-01-03 | 2025-01-03 | Steps will be signed for in real time and at time of occurrence. | Yes | Deviation
4643 | FZ-RF503E008-8NYA-High foam levels identified in SUB X2210 | michael.knapp@sanofi.com | 8NYA 1333 | 2025-01-04 | 2025-01-04 | Foam level is managed by automated antifoam dosing
4642 | 8 NYA Gemba - Support Services - 705691 Media Outlet Line Sealed - RF573E036 | Nicholas.grote@sanofi.com | 08-1202 | 2025-01-03 | 2025-01-04 | Only one the line that was used ot fill the tote is sealed off. The other two outlet lines are left available to connect to the SUB and for the first tote the second outlet is also used for ATF wetting.
4641 | 8NYA Downstream - EGHAJ011 - Leak on Outlet of Recirculation Pump Tri-clamp | wilfredo.lopez@sanofi.com | 8NYA Downstream Phase 2 CNC, Zinc row | 2025-01-03 | 2025-01-03 | It is expected that all connections remain integral and all material is contained within the SUT
4640 | 51NYA - Pest discovered in 143 | Craig.stevenson@sanofi.com | 51NYA Room 140 | 2025-01-03 | 2025-01-03 | It is expected that no pests are in vlassified clean rooms.
4639 | 8 NYA Upstream GEMBA FZ SUB 2200 TCU X-2203001 Alarm Status PN: 706930 Lot: EGHCE010 | GARY.DODAKIAN3@SANOFI.COM | 8 NYA upstream, 1333 sub 2200 | 2025-01-02 | 2025-01-02 | TCU is not in a state alarm in deltav and will send out notifications if there is an alarm locally or failure
4638 | 8 NYA GEMBA- Need path forward on testing blank run samples | richard.manasian@sanofi.com | QC lab 68NYA 3rd floor | 2025-01-03 | 2025-01-03 | We would have testing instructions for routine samples performed during a blank run
4637 | Upstream, FZ, RF503E010 ATF Build instructions missing to clamp off Drain Tee | jeffrey.boldizar@sanofi.com | 8NYA 1333 | 2025-01-03 | 2025-01-03 | Instructions to clamp off drain Tee to insure the aseptic connection does not get wetted out during wetting
4636 | CNC transfer disassembly step was not signed for | katherine.reardon@sanofi.com | 8 NYA, CNC | 2025-01-01 | 2025-01-03 | Step in EBR should have been signed for per QU-SOP-0083844 step 6.55.18 "Sign EBR for transfer disassembly". | Yes | Deviation
4635 | SWECO CIP Requalification Discrepancy | relja.nedeljkovic@sanofi.com | 76 NYA room 3103 | 2024-11-13 | 2024-12-20 | It is expected that minimum soak/rinse time and or volume should not be exceeded to ensure a worst case scenario during the RQ.
4634 | SWECO CIP Requalification Discrepancy | relja.nedeljkovic@sanofi.com | 76 NYA room 3103 | 2024-11-13 | 2024-12-20 | It is expected that minimum soak/rinse time and or volume should not be exceeded to ensure a worst case scenario during the RQ.
4633 | Cz SUB Growth Media Disposal not executed in EGHBN011. Support Services. | robert.morawski@sanofi.com | 8NYA CNC RM 1210 | 2024-10-30 | 2025-01-02 | Steps in EGHBN011 for Growth Media Disposal should have been executed in the EBR. A direct take-out of SU 212356 was taken out against Cost Center US04F0029 on 30OCT24 @07:39. This take out should have been executed and issued to the BOM of EGHBN011. | Yes | Deviation
4632 | Two extended Membrane storage cart PM's went overdue (Work orders 3168925, and 3168926) | Barry.Haines@Sanofi.com | 8NYA 1503 | 2025-01-02 | 2025-01-02 | That PM's are completed on time and before the due date
4631 | 8 NYA Support Services - Particle observed in 3K SUM Bag pre use inspection | bailey.gleason@sanofi.com | 1350 | 2025-01-01 | 2025-01-01 | no particles should be visible in bag
4630 | holds on X2210 for perfusion failure, weight hi hi alarms | brendan.dickhaut@sanofi.com | 8NYA 1333 | 2025-01-02 | 2025-01-02 | SUB maintains weight limits
4629 | N2 bulk supply tank is empty | brendan.dickhaut@sanofi.com | 8NYA | 2025-01-01 | 2025-01-01 | Bulk tank should be refilled and supplied so it does not run dry.
4628 | 8NYA Upstream X-2220 CZ pH Probe divergence during inoculation | kurt.marsocci@sanofi.com | 8 NYA Phase 2 Upstream | 2025-01-01 | 2025-01-01 | pH set point is 7.10 and probe 1 was reading around 6.98 |  | No Deviation
4627 | 8NYA - Support Service - insufficient comments for manual weigh out. RF545E122, RF549E047. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2025-01-01 | 2025-01-01 | hold gemba to see path forward | Yes | Deviation
4626 | Cedex Issue from 31Dec24 | ryan.devanna@sanofi.com | 1333 and 1403 8NYA | 2025-01-01 | 2025-01-01 | n/a | Yes | Deviation
4625 | Cedex ran of out reagent, unable to calibrate new reagent, Cedex 2009178 used | gary.dodakian3@sanofi.com | 8RM1333 | 2024-12-31 | 2024-12-31 | FBL-EWI-000026, sample analysis, review results, results should have been available. | No | Deviation
4624 | 8NYA-Cz- Cz Probe Prep executed in Fz records | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-12-30 | 2024-12-30 | It is expected that probe records for the correct area/ product line are used
4623 | 8NYA Downstream - Cerezyme MBRs Populating the Incorrect CX Values for Filter Part Numbers | wilfredo.lopez@sanofi.com | 8RM1503 | 2024-12-30 | 2024-12-30 | It is expected that the correct CX values are recorded to test the filters with correct recipes. If not corrected prior to launching FIT recipes, the incorrect test may be used. | Yes | Deviation
4622 | 1 operator signed both performed and witnessed boxes in QU-OPE-0295249 Lot: EGF3B08A | craig.valery@sanofi.com | 51NYA RM132 | 2024-12-26 | 2024-12-30 | it is expected that a second operator witnesses step
4621 | 8NYA-Cz-EGHBN014 -X2240 Bleed Clamp found closed after Bleed initiated in DeltaV | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2024-12-22 | 2024-12-23 | Per Step 6.11.28 of QU-SOP-0080429, the operator is expected to "Remove pinch clamps from Bleed flow path for active bleed tote."
4620 | Upstream, CZ, Rf600E054, MCC X2560 WFI ran out during post swab samples rinse for CIP | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-12-30 | 2024-12-30 | Enough WFI volume to finish rinse to complete CIP
4619 | 8NYA- Cz- EGHBN016SUB X2250 Sample Line and Syringe welded to Incorrect Line during Media Charge | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2024-12-30 | 2024-12-30 | It is expected that the sample syringe and 5ft extension be welded to the Sample Line of the SUB
4618 | 8NYA - Support services - Continuous alarm during NTO nano wetting, can't complete. | JACK.LAPLANTE@SANOFI.COM | 8NYA | 2024-12-30 | 2024-12-30 | Looking for path foward.
4617 | Cz Phenyl Eluate- Step Recovery below target | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2024-12-29 | 2024-12-29 | Determine path forward for processing
4616 | Protocol samples were taken prior to mixing | pavel.dobrovolskiy@sanofi.com | 8NYA 1503 | 2024-12-27 | 2024-12-27 | its expected that if mixing is required or suggested as part of the study the protocol would be written to include the mixing speed and duration
4615 | X-2210 Fz train 1 media rate of change issue at perfusion doubling stage | dana.hopkins@sanofi.com | 8NYA 1333 | 2024-12-27 | 2024-12-27 | DeltaV recipe should not hold
4614 | CZ Zinc POU samples were not delivered during 830 sample drop off | pavel.dobrovolskiy@sanofi.com | 8NYA 1503 | 2024-12-27 | 2024-12-27 | its expected that samples will be dropped off and delivered to QC from testing | Yes | Deviation
4613 | unable to sample E3 reactor | nicole.lufkin@sanofi.com | 51 NYA 140 | 2024-12-27 | 2024-12-27 | metaflex would be in cal | Yes | Deviation
4612 | Fz sub sampled under Cz cell type on Vicell | ryan.devanna@sanofi.com | 1333 8NYA | 2024-12-27 | 2024-12-27 | n/a | Yes | Deviation
4611 | 8 NYA APU-room 1333-Dept CC8- MCC skid X-2543 load pressure indicator FB-4773-016 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 8-1333 | 2024-12-26 | 2024-12-27 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4610 | A Gemba is needed for: 8NYA - Downstream - EGHBV030 Automation intervention needed to step back Q Membrane recipe. | kenny.leland@sanofi.com | 8 NYA 1503 | 2024-12-26 | 2024-12-26 | Recipe should have started pH adjust. | Yes | No Deviation
4609 | 8NYA- Cz-X2240 DO Probe Divergence | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-12-26 | 2024-12-26 | It is expected that The DO probes performance do not diverge, and are assessed if they do | Yes | No Deviation
4608 | 8NYA- Downstream - EGHBT041 Cerezyme Phenyl Eluate Post Chem13 % Recovery Out of Range | wilfredo.lopez@sanofi.com | 8RM1324 | 2024-12-25 | 2024-12-25 | Expected that all values during A280 analysis remain within specified ranges. Gemba to discuss possible root causes and capture the event in a QE. | Yes | No Deviation
4607 | 2831 High-High Pressure during Blank Run 1 | John.Lytle@sanofi.com | 8NYARM1403 Upstream Phase 2 | 2024-12-25 | 2024-12-25 | Pressure maintains under High-High limits. |  | No Deviation
4606 | 8NYA CZ 3000L media tote failed to automatically switchover. | christian.galvez-tello@sanofi.com | 8NYA RM1210 CNC | 2024-12-25 | 2024-12-25 | N/A | No | No Deviation
4605 | 8NYA - Support Services - Buffer shared between Fz/Cz not able to be sampled for the individual product | jack.laplante@sanofi.com | 8RM1350. | 2024-12-25 | 2024-12-25 | That the procedure listed in LIMS Workflow would apply for all buffers shared amongst multiple product streams.
4604 | 8NYA - Support Service - Biolims received by MFG under incorrect batch. RF618E023 - 760739 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-12-24 | 2024-12-24 | Hold gemba to see path forward
4603 | 8NYA-Cz- Clamp on X-2240 SUB Bleed Line | Yves.Antoine@sanofi.com | 8NYA Room 1403 | 2024-12-24 | 2024-12-24 | It is expected the flow path to clear of any clamp
4602 | -Cz-SUB X-2240 Capacitance probe 1 divergence | Yves.Antoine@sanofi.com | 8NYA Room 1403 | 2024-12-23 | 2024-12-23 | It is expected that capacitance probe 1 to work properly |  | No Deviation
4601 | Minimum pressure test hold time not met during blue axichrom column pressure test | jacob.kallon@sanofi.com | 51 NYA room 122B | 2024-12-21 | 2024-12-23 | Per QU-SOP-0035605 step 6.9.15 the total pressure hold time should be 25-30 minutes
4600 | Gemba 8 NYA - Support Services - 1K SUM Bag failed pre use inspection-particles | bailey.gleason@sanofi.com | 8 NYA Room 1350 (Solution Prep) | 2024-12-23 | 2024-12-23 | expect to observe no particulate
4599 | 8NYA-Gemba-Downstream Cz- Records executed were not the current version at time of operation | aidan.mula@sanofi.com | 8NYA | 2024-12-23 | 2024-12-23 | Records should be aborted/cancelled in Athena, MES, and LIMS as soon as a new record version is released per SOP.
4598 | Seprafilm Lot DBYSEP086 12M: An adverse trend was detected for Linear Swelling with a predicted value of 31.02% at 36 months (specification ≤30%). Seprafilm Lot CBYSEP023 30M: Linear Swelling result of 24.7% is OOT (95% prediction interval at 30 months 17.62619- 23.41214%). | yekaterina.astakhov@sanofi.com | 76 NYA | 2024-12-19 | 2024-12-19 | It is expected that there is no adverse trend. It is expected that there is no OOT.
4597 | 8NYA X2210 Perfusion start failure | ralston.najarro@sanofi.com | 8NYA1333 | 2024-12-23 | 2024-12-23 | ATF should begin perfusion
4596 | 8NYA - Support service - RF623E021 200L SUM bag leaked during initial WFI | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-12-23 | 2024-12-23 | N/A
4595 | 8NYA - Support Service- RF545E119 manual weigh out performed with insufficient comment. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-12-23 | 2024-12-23 | N/A | Yes | Deviation
4594 | Missed step on MES/Missed MES Record Step on MCC CIP for X2811 | marcelo.dossantos2@sanofi.com | 1403 | 2024-12-22 | 2024-12-23 | Per MABC-QU-SOP-80250, step 6.18 | Yes | No Deviation
4593 | 8 NYA Gemba - Support Services - Date of MFG on Enzyme COA did not match Bottle Date | nicholas.grote@sanofi.com | 08-1202 | 2024-12-22 | 2024-12-22 | Dates match
4592 | 8 NYA Gemba - Support Services - 1K SUM Bag Failed Inspection 692170 | nicholas.grote@sanofi.com | 08-1350 | 2024-12-22 | 2024-12-22 | No particulates and bag is in good condition
4591 | FZ, Upstream, RF503E008, SUB selective dev alarm | jeffrey.boldizar@sanofi.com | 8NYA 1333 | 2024-12-22 | 2024-12-22 | SUB probes to be close in value to each other
4590 | Cold room 147 went out of range | adam.rowe@sanofi.com | 51NYA, rm147 | 2024-12-21 | 2024-12-21 | The cold room should maintain control to its setpoint of 4c.
4589 | 8NYA-Cz-X2230 pH probe divergence following daily sampling activities | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2024-12-22 | 2024-12-22 | It is expected that while performing one point calibrations of the probes per EWI, the correct values are pulled into the record to be applied to further processing decisions
4588 | 8 NYA - Gemba - 760741 Media Setting Incorrect Expiry in MES/Athena | Nicholas.grote@sanofi.com | 08-1202 | 2024-12-21 | 2024-12-21 | 760741 being stored at ambient temperatures should be assigned 14 day expiry.
4587 | 8NYA-Cz- EGHBNE015- ATF10 assembly error | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-12-21 | 2024-12-21 | It is expected that the A2B connector be setup on the Right side of the retentate elbow per step 6.1.22 (rev 5.0)
4586 | GEMBA - 8 NYA DOWNSTREAM - Seal made in wrong spot during "Shipping Bag Replacement" | adam.keister@sanofi.com | 8 NYA 1703 | 2024-12-21 | 2024-12-21 | For operators to correctly seal the lines per SOP instructions. | Yes | Deviation
4585 | 8NYA DS - Shipping bags failed pre-inspection | elias.mcdonnell@sanofi.com | 8NYA, DS phase 2, room 1703 | 2024-12-20 | 2024-12-20 | No hair/microfibers, embedded particulate nor surface scratches are permitted for shipping bags
4584 | CZ, RF639E012, Both equil buffer line sealed off | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-12-20 | 2024-12-20 | MCC buffers to be connected and able to feed the skid with the requested buffer
4583 | Gemba - 8 NYA - Support Services - 3K Tote Bag (692174) damaged tubing connecter - Pre Use for RF621E257 - 760741 Cz media | nihcolas.grote@sanofi.com | 08-1202 | 2024-12-20 | 2024-12-17 | Operator stated it was like this before install. Bag connecters are not damaged.
4582 | DOT Internal Audit FRAD23E0814 event was advanced to closure in Phenix system however CAPA listed as requiring effectiveness review (ER) but the ER were never created. | randall.porcella@sanofi.com | 68NYA | 2024-12-20 | 2024-12-20 | Effectiveness Reviews would be created | Yes | No Deviation
4581 | 68NYA - QC Micro - Late entry with no verifiable source | Craig.stevenson@sanofi.com | 68NYA Room 4700 | 2024-11-19 | 2024-12-20 | It is expected that a verifiable source be used to generate a late entry and correction.
4580 | 8NYA-Cz- MCC X2831 Leak observed at Equilibration/ Load Port#1 Sanitary connection | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2024-12-20 | 2024-12-20 | It is expected that the MCC is free of leaks and integral for production activities
4579 | Pre Autoclave Aber probs gasket not change out per FBL-EWI-000011 | bhargav.patel@sanofi.com | Room 1333 | 2024-12-19 | 2024-12-20 | Per FBL-EWI-000010, during step execution Atuclave prep, gaskets should be change out
4578 | Upstream GEMBA SUB X-2230 Inoculation BV-2230107 not in correct state in DeltaV | GARY.DODAKIAN3@SANOFI.COM | 8 NYA / 1403 | 2024-12-19 | 2024-12-19 | valve is in the correct state and controls O2 levels per the recipe
4577 | Expired mixed gas bottles found at 51 NYA | Jason.Silva@Sanofi.com | 51 NYA gas shed. | 2024-11-10 | 2024-12-19 | bottles replaced prior to expiration within 1week | Yes | Deviation
4576 | 8NYA - Support Service - RF621E254 / EGHAV254 POU samples received in biolims inadvertently | bailey.gleason@sanofi.com | 68 QC lab | 2024-12-18 | 2024-12-19 | hold a gemba to see path forward
4575 | Ineffective CAPA FRAD24A0006, FRAD24A0008 | kyle.tesorero@sanofi.com | 8NYA | 2024-12-18 | 2024-12-18 | MABC-QU-SOP-0079722 - 6.11.26 In the case of an out of specification result or step failure, inspect all vessels and fibrettes, hold a gemba to determine a path forward and document gemba outcome as exception in batch record. MABC-QU-SOP-0019707 - 6.7 Execution Mitigation - Steps should include additional detail on what information is required within Exception for manually recorded GXP data (i.e. Weight indicator)
4574 | 8NYA FZ media pou705691 tested on 29Sep24 incorrect sample/batch number | heidy.ruano@sanofi.com | 8NYA | 2024-09-29 | 2024-12-18 | Correct sample and lot number
4573 | Gemba 8 NYA - Support Services - 705691 Transfer timer expired | william.lee@sanofi.com | 8NYA/1350 | 2024-12-18 | 2024-12-18 | Transfer should have occurred before timer expired at 1640.
4572 | CD40L Placebo Fill 20R Vial-EFGC0203-76 NYA-Unaccounted data by QC Micro | Alyssa.Barrett@sanofi.com | 68 NYA/Room 4700/QC Micro | 2024-11-25 | 2024-12-18 | Deviated from QU-SOP-0019490 Counting of Microbial Plates in Quality Control (V4.0) step 6.2.16 stating to record the counted by section following completion of all counting activities for a given counting package. Also QU-SOP-0074684 Good Documentation Practices (V8.0) step 6.1.2 states all signatures and entries are to be made sequentially at the time of execution. | Yes | Deviation
4571 | No Products – No Lots- 76NYA – No Reportable result for non-viable site 210 for routine HA Fill EM performed 21NOV24 | megan.botti-marino@sanofi.com | 76NYA | 2024-11-21 | 2024-12-18 | It is expected site 210 to have reportable EM results.
4570 | 8 NYA GEMBA -Downstream CZ MMC, Particulates seen inside the bubble trap | alex.tello@sanofi.com | 1503 | 2024-12-18 | 2024-12-18 | No particulates to be inside the bubble trap. | Yes | Deviation
4569 | 8 NYA GEMBA - CZ Phenyl, Devices not removed before storing skid. | alex.tello@sanofi.com | 1503 | 2024-12-17 | 2024-12-18 | All SUT to be removed when removing devices | Yes | Deviation
4568 | 45NYA QCC - TSH Missed Stability Window for Particulates | abel.agnes@sanofi.com | 45NYA QCC Lab Rm 1130 | 2024-12-18 | 2024-12-18 | Particulates are tested within test window of timepoint. Particulates cannot be tested with frozen contingency.
4567 | Gemba - 8 NYA - Support Services - 705691- RF573E033 | nicholas.grote@sanofi.com | 08-1202 | 2024-12-17 | 2024-12-17 | Otiker clamps do not leak.
4566 | 8 NYA Gemba - SOP Lacking Direction for 200 Liter Tote Fill - 705667 - RF538E006 | nicholas.grote@sanofi.com | 08-1202 | 2024-12-17 | 2024-12-17 | EWI-SOP provide instruction on proper bag setup and transfer. | Yes | Deviation
4565 | A280 SOP Alignment Issue | katherine.reardon@sanofi.com | 8 NYA, downstream suite 2, room 1503 | 2024-12-16 | 2024-10-16 | QU-SOP-0079722 V 2.0 multiple steps | Yes | Deviation
4564 | EWI-000015 Wrong Part # for 3k tote bag in tote to tote transfer. Support Services. | robert.morawski@sanofi.com | 8NYA CNC RM 1210 | 2024-12-17 | 2024-12-17 | Part# 682174 should have been the 3K tote bag named in the tote to tote transfer steps. | No | No Deviation
4563 | Repligen Response to NBOD23E0073 indicates some hollow fibers used in ATF10 were polysulfone prototype, impacted ATF10 batches unknown | sheila.pandolfo@sanofi.com | ATF was rejected by Upstream Mfg operators in 8 NYA Suite 2 and investigated at Repligen | 2023-04-26 | 2024-12-16 | It is expected all hollow fiber tubes have PES (polyethersulfone) as the material of contruction.
4562 | rupture disk alarm came out from Clean steam generator #2 at 45 NYA. Alarm did not clear. Clean steam is down until further notice. | andrew.apel@sanofi.com | 45-180 | 2024-12-17 | 2024-12-17 | Clean steam generator#2 should be running with no alarm.
4561 | GEMBA 8NYA Upstream CRYO Labels EGHSD011 | GARY.DODAKIAN3@SANOFI.COM | 8 NYA / 1280 | 2024-12-16 | 2024-12-16 | all cryo labels are present on Cryo | Yes | Deviation
4560 | GEMBA 8NYA Upstream CRYO Labels EGHSD011 | GARY.DODAKIAN3@SANOFI.COM | 8 NYA / 1280 | 2024-12-16 | 2024-12-16 | all cryo labels are present on Cryo | Yes | Deviation
4559 | 68NYA Support APU-Validation-FB-5146 Sensor Input Module found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1745 | 2024-12-10 | 2024-12-16 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4558 | SUB probes are pulling incorrectly. Probe 1 locations pull probe 2 results. Probe 2 locations pull probe 1 results. | brendan.dickhaut@sanofi.com | 8NYA, Room 1403 | 2024-12-16 | 2024-12-16 | Probe results should pull to correct input. Probe 1 should pull to input 1. | Yes | Deviation
4557 | 8NYA -CZ - EGHBN015 Failed inspection of SUB bag | Yves.Antoine@sanofi.com | 8NYA Room 1403 | 2024-12-15 | 2024-12-15 | Passing bags
4556 | 8NYA Support Services Incomplete Media install Record RF639E013 | genesis.torres2@sanofi.com | 8nya CNC | 2024-12-15 | 2024-12-15 | Media install on train 1
4555 | 8NYA - Support Service - Media spill from waste bag in CNC | bailey.gleason@sanofi.com | 8NYA CNC | 2024-12-14 | 2024-12-14 | hold gemba to see path forward
4554 | 8 NYA - GEMBA Downstream CZ. Step not signed during execution. | alex.tello@sanofi.com | CNC | 2024-12-12 | 2024-12-14 | Sign step right after execution
4553 | 8NYA - Gemba - Shipping bag failed pre-inspection | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 post viral | 2024-12-14 | 2024-12-14 | no particles are allowed on the inside of the bag | Yes | Deviation
4552 | 8NYA - ALL - Vante 4160 Sealer Instructions missing key references with last EWI revision | christopher.arnold@sanofi.com | 8RM1202 | 2024-12-13 | 2024-12-13 | That SOP/EWI revisions will not substantially alter the document and that all necessary parts will be in the document.
4551 | 8NYA Downstream - EWI-000031 - Sealer Head Set-Up References Not Available | elias.mcdonnell@sanofi.com | 8NYA, 1503 | 2024-12-12 | 2024-12-12 | Reference links are available for use
4550 | 8 NYA Gemba - Support Services - 3K SUM Bag Leak | nicholas.grote@sanofi.com | 08-1350 | 2024-12-12 | 2024-12-12 | SUM bag does not leak. | Yes | Deviation
4549 | GZ1337-01 Weighing Information was not recorded in EBR or Athena. | yunbo.lou@sanofi.com | 8 NYA | 2024-12-10 | 2024-12-12 | Information of manually weighed material is expected to be recorded in EBR. | Yes | Deviation
4548 | 8NYA-Cz-Rf639E007- Missed verification of Weld/ Seal | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-09-05 | 2024-12-12 | Per EWI, all welds/ seals need a contemporaneous verification of activities | Yes | Deviation
4547 | 8NYA-Cz-RF639E007- Missed material takeout and verification during pC02 probe setup operation | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-08-28 | 2024-12-12 | It is expected that all material used in cGMP production be consumed against the Batch Bill of Materials or in SAP/ Athena | Yes | Deviation
4546 | 8NYA Gemba - EGHDS013 - Cerezyme BIRF printed on approved version instead of effective version. | stanley.durand@sanofi.com | 8NYA | 2024-12-12 | 2024-12-12 | It's expected only the effective version of the BIRFs should be used per QU-SOP-0074092.
4545 | No Products – No Lots- 68NYA QC Micro – Gram stain form not printed for 2 bioburden samples for growth determination | megan.botti-marino@sanofi.com | QC Micro lab/68 NYA | 2024-11-12 | 2024-12-11 | It is expected per step 6.3.10 of QU-SOP-0019490, attachment 1B from MABC-QU-SOP-0070845, “Processing Microbial Isolates Sent to the QC Microbiology Identification Lab (v8.0),” should have been used to perform the gram stains and attached to the associated bioburden paperwork afterwards. This lead to no verifiable source for the calibrated timer required per MABC-QU-SOP-0073940-Gram Stain Method for Staining Microorganisms (v4.0). | Yes | Deviation
4544 | Thyrogen formulated bulk A280 significant figure discrepancy raw data vs reported | craig.valery@sanofi.com | 51 NYA | 2024-12-08 | 2024-12-08 | significant figure results should align | Yes | Deviation
4543 | Water bath not wiped with IPA prior to filling with water for cell thaw | johny.lisitu@sanofi.com | 68 NYA CBA | 2024-12-09 | 2024-12-11 | Per QU-SOP-0024808 Step 6.5.1, the water bath is to be wiped with 70% IPA. | Yes | Deviation
4542 | 8NYA Material Air Door Interlock Breached Due to E-Stop Trigger | tanni.arif@sanofi.com | Material Air Lock | 2024-12-11 | 2024-12-11 | Only one door expected to stay open per requirements | Yes | No Deviation
4541 | 8NYA-Fz-1L Shake Flask Vi-Cell sample run assigned incorrect product line at testing | JOHN.LYTLE@SANOFI.COM | Rm1343 8NYA | 2024-12-11 | 2024-12-11 | ViCell input should be product code "Fabrazyme" for Fabrazyme sample activities
4540 | 705680 - Zinc acetate weigh out -- HU weights do not equal material requirement | nicholas.grote@sanofi.com | 9 NYA / 1350 Solution Prep | 2024-12-11 | 2024-12-11 | 16 kg of zinc acetate material appeared to be weighed out and ready to add to batch, but 3 HU's totaled only 13.182 kg (gap of 3.018)
4539 | 51 NYA Thyrogen Purification-Clarification- documented temperature of diluted harvest out of range. | jacob.kallon@sanofi.com | 51 NYA | 2024-12-10 | 2024-12-10 | Per QU-OPE-0351657, the expected temperature range is 6-10 degrees Celsius.
4538 | Leaking bottle manifold post filter during media batching. | jesse.demers@sanofi.com | 51RM124 | 2024-12-10 | 2024-12-10 | Bottle manifold will not leak. | Yes | Deviation
4537 | 8NYA-Gemba-Downstream Cz Phenyl Blank Run CNC breakdown step was not signed at time of performance | aidan.mula@sanofi.com | 8NYA/CNC | 2024-12-09 | 2024-12-09 | Steps should be signed at time of performance
4536 | WFI Soak and rinse for Bioreactor E3 not recorded in MES during maintenance activities | adam.hildebrand@sanofi.com | 51-140 | 2024-12-08 | 2024-12-09 | QU-OPE-0525093 step 4.4.3
4535 | 8 NYA - Support Services - 3K tote bag failed inspection | william.lee@sanofi.com | 8NYA/CNC | 2024-12-09 | 2024-12-09 | 3K tote bag should not be damaged prior to installation | Yes | Deviation
4534 | Q flowthrough sitting unadjusted for 6+ hours | pavel.dobrovolskiy@sanofi.com | 8nya 1503 | 2024-12-09 | 2024-12-09 | its expected that material will not go expired in MES
4533 | 76NYA Support APU-Room 347 Conductivity Indicator F-20550-006 Found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-347 | 2024-12-06 | 2024-12-09 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4532 | 76NYA Support APU-Room 347 - Temperature indicator F-8740-006 had no as founds taken during a scheduled Calibration | robert.bauch@sanofi.com | 76-347 | 2024-12-06 | 2024-12-09 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4531 | 76NYA Support APU-Room 347 - Conductivity Indicator F-8740-007 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-347 | 2024-12-06 | 2024-12-09 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4530 | 8NYA - Gemba - Sample transport coolers not being conditioned for 24hrs | lukus.palzkill@sanofi.com | cold room 1220 | 2024-12-09 | 2024-12-05 | sample transport coolers should be conditioned for 24hrs in the cold rooms
4529 | WFI loop system ( WFI Pumps) Down affecting also Clean Steam Generator #1 and #1 Down. Both Wfi tank are full and Pressure in Clean steam loop is at 41 Psi at 0640 am this morning 09dec2024 | Justin.Carpenter@sanofi.com | 51-150 | 2024-12-09 | 2024-12-09 | WFI loops should be running feeding Clean steam generators | Yes | Deviation
4528 | Visual inspection Fail on SUB Bag x2240 ( train 3) | John.Lytle@sanofi.com | 8NYARM1403 Upstream Phase 2 | 2024-12-08 | 2024-12-08 | Further inspection will be needed to determine if the SUB bag has failed inspection . | Yes | Deviation
4527 | Cerezyme - RF639E013 - 8NYA - High Harvest Pump Pressure due to SUB media line restriction | michael.knapp@sanofi.com | 8NYA 1403 | 2024-12-07 | 2024-12-07 | No restrictions or high pressure events on the SUB and harvest lines | Yes | Deviation
4526 | 8 NYA Gemba - Support Services - Atria left in Autoclave from 05DEC24 11PM - 07DEC24 11AM | nicholas.grote@sanofi.com | 08-1373 | 2024-12-05 | 2024-12-07 | Atriums are removed from autoclave and the record is completed. Incomplete tasks from shift are communicated and passed off to incoming shift. | Yes | No Deviation
4525 | 8NYA-Cerezyme Upstream-EGHAV213 Bioburden sample 1491238 MEDIA_POU760741_BIOBURDEN sample cap dislodged and leaking sample | elizabeth.endsley@sanofi.com | Micro Lab 4700 | 2024-12-06 | 2024-12-06 | For the cap to be on tight and not leaking | Yes | Deviation
4524 | 1L wave temperature excursion | david.loik@sanofi.com | 51 NYA | 2024-12-06 | 2024-12-06 | Controls would be in place to allow culture to stay within range while sampling.
4523 | 8NYA-Gemba-Downstream Shipping Bag Fill, clamp was not fully seated during barrel transfer causing liquid to enter the unused end of wye | aidan.mula@sanofi.com | 8NYA/RM1703 | 2024-12-05 | 2024-12-05 | Clamp is fully seated for transfer operations to ensure no liquid passes through | No | Deviation
4522 | Missing a verification on a manual takeout | dennis.bazinet@sanofi.com | 8 NYA Phase 2 | 2024-12-04 | 2024-12-05 | Manual step needs to be verified | Yes | No Deviation
4521 | GEMBA for Missed Personnel Monitoring Counts/ Cz Spreadplate and 1 Missing Count Entry on Batch INC-18606M | jessica.kychler@sanofi.com | 68 NYA, 4700, QC Micro | 2024-11-09 | 2024-11-20 | Per step 6.4.1 of QU-SOP-0019490, data entry into OneLIMs should occur concurrently with counting activities.
4520 | 8NYA Cz post Vital shipping bag inspection, embedded particulate found | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1702 | 2024-12-05 | 2024-12-05 | To get a Phath forward | Yes | Deviation
4519 | 8NYA-Gemba-Downstream Shipping Bag Failed Pre-Inspection due to particulate | aidan.mula@sanofi.com | 8NYA/RM1703 | 2024-12-05 | 2024-12-05 | Bags should not contain particles within the product contact area. | Yes | Deviation
4518 | Cz Nano Record- Post formulation concentration check error | nathaniel.barnes@sanofi.com | 8NYA 1323 | 2024-12-05 | 2024-12-05 | Correct the record formula to pull in the correct value to base its calculations off. Step also seems redundant if SOP is having us perform this same calculation manually
4517 | Incorrect OneLIMS reagent dilution volumes recorded for QU-OPE-0297218 assay | aimee.park@sanofi.com | 45NYA, QC Chemistry Room 1136 | 2024-10-16 | 2024-12-04 | The 0.1%TFA volume would have been recorded as 100uL and the HPLC water 900uL for Step 6.1.7 of QU-OPE-0297218
4516 | OOT: No Product Impact- 45 NYA- TSH UFB: EGF4B05 P/N GZ5156 7.0 | nancy.watkins@sanofi.com | 45NYA, QC Chemistry | 2024-12-03 | 2024-12-03 | per QU-FOR-0024431 V12 Trend Limit = 0.8569-0.9885 mg/mL
4515 | GZ1412 (Filter WFI) not created in OneLIMS | cindy.guerra@sanofi.com | 45-190 | 2024-12-04 | 2024-12-04 | LAL sample would be available in OneLIMS for GZ1412 | Yes | No Deviation
4514 | 51 NYA Thyrogen Purification-UF/DF collected diafiltration volume less than calculated volume. | jacob.kallon@sanofi.com | 51 NYA | 2024-11-30 | 2024-12-03 | It is expected the collected dialfiltrration volume (71 kg) is greater than or equal to the calculated diafiltration 71.050kg. volume | Yes | Deviation
4513 | 76NYA Support APU-room 347- conductivity meters F-13587-001 and F-13587-002 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-347 | 2024-12-03 | 2024-12-04 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
4512 | 8 NYA Upstream Cz Eluate Barrel Leak | kurt.marsocci@sanofi.com | 8 NYA Room 1220 | 2024-12-03 | 2024-12-03 | No leak
4511 | 8NYA-Upstream-RF639E005-Cz Eluate Barrel Leak Rm1220 | ricardo.rodriguez@sanofi.com | 8RM1220 | 2024-12-03 | 2024-12-03 | The 200L bag would be properly installed to not create a pinch point which may result in a leak.
4510 | 68 NYA - No Product - No Lot. Open Time not recorded for settling plates for Bioburden Testing | alyssa.barrett@sanofi.com | 68 NYA/ 4700 QC Micro | 2024-11-30 | 2024-12-02 | Per QU-SOP-0057388 step 6.2.4 complete exposure time should be recorded. | Yes | Deviation
4509 | 8 NYA Gemba - Support Services - Manual weigh out missing second operator verification | nicholas.grote@sanofi.com | 08-1350 | 2024-12-01 | 2024-12-02 | Manual weight will have a second operator signing for verification of the information.
4508 | NPI (Anagni TT): Formulation Buffer Bag Leak at Anagni Site | elissa.karol@sanofi.com | 8NYA Room 1210 and Anagni, Italy Site | 2024-09-12 | 2024-09-13 | Shipping bags do not leak
4507 | 76NYA Biosurgery APU-Room 3102 Temperature indicator F-6934-007 overdue for calibration | robert.bauch@sanofi.com | 76-3102 | 2024-12-01 | 2024-12-02 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date | Yes | Deviation
4506 | 68NYA Support APU-Room 4700-QCM68-Conductivity meter FB-3488 overdue for calibration. | robert.bauch@sanofi.com | 68-4700 | 2024-12-01 | 2024-12-02 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
4505 | 8NYA - Support Services - RF1B3E016 - Conductivity result OOS | christopher.arnold@sanofi.com | 8RM1350 | 2024-12-01 | 2024-12-01 | That the buffer sample will test within range when the buffer formulation is performed correctly. | No | Deviation
4504 | Conductivity test failed, result out of range. | maria.palmar@sanofi.com | 8 New York Ave, room 1350. | 2024-12-01 | 2024-12-01 | Conductivity testing to not fail.
4503 | 8NYA X-2250 Hold during ATF changeout, Batch:EGHBN012 | Yves.Antoine@sanofi.com | 1403 | 2024-12-01 | 2024-12-01 | We Expect the Recipe to work to properly after an ATF changeout
4502 | 8NYA - Support Service - RF618E019 range of weighed out enzyme does not match the target value for consumption. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-12-01 | 2024-12-01 | hold gemba to see path forward
4501 | Gemba for Open time not recorded for EM settle plate for bioburden testing | nathaniel.milde@sanofi.com | 68 NYA/ 4700 QC Micro | 2024-11-30 | 2024-11-30 | Per MABC-QU-SOP-0057388 step 6.2.4 complete exposure time should be recorded.
4500 | 8NYA GEMBA - Automation intervention for phenyl ETP | lukus.palzkill@sanofi.com | 8NYA Suite 2/CNC | 2024-11-29 | 2024-11-29 | WFI is usually connected to port 4 | Yes | No Deviation
4499 | 8 NYA Upstream Eluate SUM P-2570002 | kurt.marsocci@sanofi.com | 8 NYA Phase 2 | 2024-11-28 | 2024-11-28 | Transfer eluate SUM to a barrel and then into a new eluate bag before deployment to purification. | Yes | Deviation
4498 | 8NYA-X2560- EGHBN013- Eluate SUM 2 failed to fully deploy and needs intervention | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2024-11-28 | 2024-11-28 | It is expected that SUM bag deployment is in able to perform eluate collection
4497 | 8NYA-Cerezyme Upstream-EGHAV200 LAL sample 1490962 MEDIA_POU760741_ENDOTOXIN sample tube cracked and leaking | larissa.fairweather@sanofi.com | 68 NYA / 4700 / QCM | 2024-11-28 | 2024-11-28 | Per QU-SOP-0070558 step 6.2.3, confirm sample containers are not cracked, leaking. The expectation is that samples are to be free of any cracks or leaks.
4496 | pH out of range during load and wash. Q process Lot EGF4B05B | craig.valery@sanofi.com | 51-130 | 2024-11-09 | 2024-11-28 | it is expected that the pH remains with in specified range | Yes | Deviation
4495 | 8NYA Upstream GEMBA Oscillating Temperature on Train 4 Harvest SUM PBM-2916001 / TCU X-2835001 | GARY.DODAKIAN3@SANOFI.COM | 8 NYA Upstream room 1403 | 2024-11-26 | 2024-11-26 | Temperature is maintained at a stready rate |  | No Deviation
4494 | 692169 bag in JIT doesn't match the label in the JIT | michael.montiverdi@sanofi.com | Solution prep 8NYA | 2024-11-18 | 2024-11-19 | HU Label applied with corresponding material information is accurate | Yes | Deviation
4493 | 8 NYA Upstream - Action Level Endotoxin Cerezyme Point of Use Media PN 760741 Lot EGHAV188 | craig.stevenson@sanofi.com | 68 NYA 4700 QCM | 2024-11-26 | 2024-11-26 | CZ Media POU760741 endotoxin result to be below alert limit |  | No Deviation
4492 | Particulate found during bag inspection | joseph.gonzalez@sanofi.com | 1403, 8NYA | 2024-11-24 | 2024-11-23 | n/a | Yes | Deviation
4491 | 8NYA-Cz-EGHBN013- X2220 DO Probe 1 and Aber Probe 2 divergence | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2024-11-27 | 2024-11-27 | It is expected that probe 1 and probe 2 remain within allowable limits | Yes | No Deviation
4490 | 8 NYA - Support Services - 3K SUM bag failed inspection | william.lee@sanofi.com | 8NYA/CNC | 2024-11-26 | 2024-11-26 | 3K SUM bag | Yes | Deviation
4489 | Temperature out of rage during regen and storage for Q EGF4B05B | Luis.dominguez@sanofi.com | 51NYA room 132 | 2024-11-05 | 2024-11-26 | per QU-SOP-0061107 attachment 7 page 1 of 3, temperature must be between 6-10C from rinse to end of storage. | Yes | Deviation
4488 | 8NYA - Support service - RF614E014 PI failed to pull value was entered incorrectly and verified by a second operator | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-11-26 | 2024-11-26 | Hold Gemba to see path forward. | Yes | Deviation
4487 | Perfusion failure on bioreactor X2230 phase 2 Upstream | brendan.dickhaut@sanofi.com | 8NYA/Room 1403 | 2024-11-25 | 2024-11-25 | Perfusion failure does not occur as media supply is maintained. clamp should have been removed prior to media changeover.
4486 | Missing Validation Study 99-0151 | Amy.McIntyre@sanofi.com | Framingham | 2024-11-25 | 2024-11-25 | That all Validation Studies are accounted for and completed by the audit due date per QU-SOP-0009729.
4485 | Thyrogen – 45 NYA – Clear Liquid Found on Floor Originating from HEPA Grate in 189B | cindy.guerra@sanofi.com | 45-189B | 2024-11-25 | 2024-11-25 | It is expected that there would be no clear fluid dripping from the HEPA grates per QU-SOP-0074732. | Yes | Deviation
4484 | 8 NYA Upstream Clamp on X-2240 Antifoam Line | kurt.marsocci@sanofi.com | 8 NYA Phase 2 | 2024-11-25 | 2024-11-25 | N/A | Yes | No Deviation
4483 | Missing witness step on BIRF for removal of media from Warm room 148. | adam.hildebrand@sanofi.com | 51 NYA RM 140 | 2024-11-05 | 2024-11-25 | It is expected that removal times are witnessed and documented at time of removal.
4482 | 8 NYA UPSTREAM X-2220 TCU Issue | todd.niziak@sanofi.com | 8NYA Phase 2 Upstream | 2024-11-25 | 2024-11-25 | Fix the issue causing the problem and restart the recipe once the SUB and TCU are at temp.
4481 | Missing Witness signature on BIRF for Inoculation Media | adam.hildebrand@sanofi.com | 51 NYA RM 140 | 2024-10-30 | 2024-11-25 | It is expected that movement of media is signed by the witness at time of movement.
4480 | Unavailable 20L Carboy for SP1 EGF4B09C process at 51NYA Purification | Sylvester.Kegbeh@sanofi.com | 51 NYA ROOM 138 | 2024-11-24 | 2024-11-24 | It is expected that 20L carboy is used per QU-OPE-0079501 Step 4.4.6, 4.8.1.3, 4.8.1.8 | Yes | Deviation
4479 | Autoclave SOP does not align with qualification for backup use of bioshield | nicholas.shoenfelt@sanofi.com | 51-129 | 2016-10-12 | 2024-11-25 | It is expected that the SOP for assembly aligns with the qualification and that the same number of sheets of bioshield would be in both. | Yes | Deviation
4478 | Fz flask not growing properly | ryan.devanna@sanofi.com | 8NYA innoc prep | 2024-11-25 | 2024-11-25 | n/a
4477 | 8NYA-Fz-EGHSD008- Unexpected Daily Sampling results at 1L Flask stage | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1343 | 2024-11-24 | 2024-11-24 | it is expected that dialy sampling produce results in line with positive cell growth
4476 | 8 NYA - Support Services - Leak during transfer of 760712 (RF601E013) | christopher.brewster@sanofi.com | Rm. 1350 at 8 NYA. | 2024-11-23 | 2024-11-23 | That the SUM bag and transfer tubing would not leak during buffer transfer.
4475 | 8NYA RM1220 ¦ Detected a leak on the bottom of the tote after swapout was completed | christian.galvez-tello@sanofi.com | 8NYA RM1220 CNC | 2024-11-22 | 2024-11-23 | N/A
4474 | Fraction 8 bag GZ6049-03 not available for ongoing Blue EGF3B08 operation. | Alyson.Belleville@sanofi.com | 51-138 | 2024-11-23 | 2024-11-23 | It is expected that all material would be available for an operation.
4473 | Gemba - 8 NYA - Support Services - 3K SUM Bags (692169) Incorrect Connection on Recirculation Line Port | Nicholas.Grote@sanofi.com | 08-1350 | 2024-11-22 | 2024-11-23 | Bag is received form the vendor with appropriate connections to perform SUM setup.
4472 | Gemba - 8 NYA - Support Services - Overfilled 760744 Media Bottles | Nicholas.grote@sanofi.com | 08-1350 | 2024-11-22 | 2024-11-22 | Bottles are not overfilled, and leak does not occur.
4471 | Cerezyme - RF639E012 - 8NYA - Col 3 High Pressure during Flush phase | michael.knapp@sanofi.com | 8NYA 1403 | 2024-11-23 | 2024-11-23 | No high pressure events on columns.
4470 | Cerezyme - RF639E012 - 8NYA - Column 2 High Pressure following Bubble Trap alarm | michael.knapp@sanofi.com | 8NYA 1403 | 2024-11-22 | 2024-11-22 | No air in system and no high pressure events on columns.
4469 | Gemba Upstream 8NYA SUB X-2250 PH Probes Trend Path Forward Discussion | GARY.DODAKIAN3@SANOFI.COM | 8 nya room 1403 sub 2250 train 4 | 2024-11-22 | 2024-11-22 | PH Probes perform as intended
4468 | Cerezyme Seed Train - Expired endotoxin sample not captured in deviation (DGHCG017 50 L) | drew.hopkins@sanofi.com | 8NYA | 2023-11-01 | 2024-11-22 | Expired samples would have an associated deviation | Yes | Deviation
4467 | 8NYA Phase 2 upstream ¦ Media tote began to leak after media installation | christian.galvez-tello@sanofi.com | 8NYA RM1220 CNC | 2024-11-22 | 2024-11-22 | N/A
4466 | 8NYA-Fz-RF503E006 MES Steps executed with manual data entry did not get verified by a second operator at the time of execution | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2024-10-17 | 2024-11-21 | Per MABC-QU-SOP-0019707-Inova Execution at 8 New York Ave, all manual data requires real time verification
4465 | Verifier missing for manual entry of sealer ID | zachary.rickard@sanofi.com | 8NYA/1323 | 2024-10-30 | 2024-11-20 | When making a comment about equipment being used there should be a verifier if the information is not coming from a verifiable source.
4464 | Fz Nano- Incorrect Seal location on Shipping Bag breakdown | nathaniel.barnes@sanofi.com | 8 NYA 1313 | 2024-11-21 | 2024-11-21 | Determine path forward to ensure shipping bag has adequate outlets after breakdown | Yes | Deviation
4463 | Gemba Request: Growth on TSA Lot 24GN01917 | andres.huerta@sanofi.com | 68 NYA QC Micro | 2024-11-20 | 2024-11-20 | Plates are not expected to have growth upon opening
4462 | 68NYA - QC Micro - Seprafilm plate accidentally discarded | Craig.stevenson@sanofi.com | 68NYA Room 4700 | 2024-11-06 | 2024-11-21 | It is expexcted that if any results greater than 0CFU are generated, then the Seprafilm would be sent for microbial ID.
4461 | External DeltaV Resource no longer at site missing core training | Seyda.Dumrul@sanofi.com | 68NYA | 2024-06-11 | 2024-11-21 | Malec would have the training for DeltaV change management completed | Yes | No Deviation
4460 | 8NYA X2230 Hold during ATF changeout | ralston.najarro@sanofi.com | 8NYA1403 | 2024-11-21 | 2024-11-21 | ATF should resume perfusion following acknowledgement of prompt | Yes | No Deviation
4459 | GEMBA Upstream X-2250 Train 4 Probe Trend Discussion | GARY.DODAKIAN3@SANOFI.COM | 8 nya room 1403 sub 2250 train 4 | 2024-11-21 | 2024-11-21 | Do Probes do not foul
4458 | TSH-Buffer prep- 45 NYA/room 190 Rupture disc on V-1502 burst during tank fill. | jeremy.wooden@sanofi.com | 45 - RM 190 | 2024-11-21 | 2024-11-21 | It is expected that the vent filter main valve would be open during the tank fill.
4457 | 8 NYA GEMBA Downstream Fabrazyme - shipping bag pre-inspection failure | lukus.palzkill@sanofi.com | 8NYA downstream suite 1 postviral room 1313 | 2024-11-21 | 2024-11-21 | Inspect the smooth surface of entire bag for damage. Visible holes are never permitted. If determination of integrity cannot be made with confidence, the bag should not be used.
4456 | Training form not completed on day of training | maggie.snow@sanofi.com | 68NYA | 2024-11-20 | 2024-11-21 | Signatures should be applied to training record on the same day of training contemporaneous with training. QU-SOP-0069666 step 5.4.2 allows training to be recorded per LMS session roster. . |  | No Deviation
4455 | No Counts entered for Environmental monitoring plates for batch FBL1327-000000492 | roxanne.vacha@sanofi.com | 68 NYA, 4700, QC Micro | 2024-11-09 | 2024-11-20 | Per step 6.4.1 of QU-SOP-0019490, data entry into OneLIMs should occur concurrently with counting activities.
4454 | Media transfer pump would not run during media prep | nicholas.shoenfelt@sanofi.com | 51-124 | 2024-11-20 | 2024-11-20 | It is expected that the media transfer pump will not stop during processing. | Yes | Deviation
4453 | Material and Equipment Information Not Completed for Two Occasions-POMPE Stability for E11 Testing. | Sarah.Adams@sanofi.com | 68 NYA, QCC | 2024-10-01 | 2024-11-20 | All equipment and materials entered contemporaneously.
4452 | Carbon Dioxide Bulk with Incorrect Source Location | david.chan@sanofi.com | 68NYA, 2700 | 2024-11-20 | 2024-11-20 | The supplier source location should say 2 Van Buren Blvd, Guilderland, NY 12085
4451 | OneLims issue with the TSH ELISA plates incubation | pinky.patel@sanofi.com | 4700, QC Micro | 2024-11-20 | 2024-11-20 | Per QU-OPE-0208394, step 6.5.5 "incubate for 1 hour at 37oC ± 1°C with shaking (50-60 RPM)
4450 | Critical Utility Modified Without CCR (but Approved FE) | mabiocampus_f&ecompliance@sanofi.com | sitewide, 76NYA and 51NYA | 2024-11-20 | 2024-11-20 | Critical Utility modifications to warrant a CCR | Yes | Deviation
4449 | Gemba needed-51NYA Thyrogen Purification -80 freezer temperature not stabilizing after product transfer | jacob.kallon@sanofi.com | 80 NYA | 2024-11-19 | 2024-11-20 | Expectation is not clear | Yes | No Deviation
4448 | Phase 1 UF\DF CIP recipe not moving forward | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1323 | 2024-11-20 | 2024-11-20 | To get a Phath forward
4447 | Missed MES Step for ATF Wetting Drain/Upstream Process | marcelo.dossantos2@sanofi.com | 1403 | 2024-11-20 | 2024-11-20 | Drained step should have been signed on MES
4446 | No Agitator Run Confirm Feedback for 250L SUMs | Dang.tran@sanofi.com | 8NYA Suite 1 Upstream | 2024-11-14 | 2024-11-14 | Agitator to have run confirm to notify that the agitator is properly working.
4445 | No products, no lots. 68NYA QCM. Analyst did not record multiple items for ID plate F24-0454-1. | kathryn.lentoni@sanofi.com | 68NYA 4700, QCM | 2024-11-12 | 2024-11-20 | Expected results are the analyst should have recorded media in logbook as well as record the incubation temperature and time. Expected results as well are the ID plate sent out to Accugenix should have been recorded. | Yes | Deviation
4444 | 76NYA - Deviation not initiated <24 hours of discovery | Craig.stevenson@sanofi.com | 76 NYA | 2024-11-15 | 2024-11-15 | It is expected that a deviation be initiated in 24 hours of discovery
4443 | 76NYA - Data Integrity issues with Action Notice 2024- | Craig.stevenson@sanofi.com | 76NYA | 2024-11-07 | 2024-11-14 | It is expected that all GDP rules and SOPs are followed.
4442 | New Zebra Scanners do not work with MES. | joel.henchar@sanofi.com | 8NYA | 2024-11-15 | 2023-08-01 | MABC-QU-SOP-0024269-SD-61/SR-61 Scanner and Base Installation and Qualification Work Instruction
4441 | 76NYA Biosurgery APU-Dept Pakg-Room 233-Balance F-18057 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 76-233 | 2024-11-19 | 2024-11-19 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4440 | An extra pouch was found in lot EBYSEP005 during packaging of unlabeled p/n 768275/FG p/n 789446 | alan.rousseau@sanofi.com | room 1017 at 74 NYA | 2024-11-18 | 2024-11-18 | Expected result is that the unlabeled SAP inventory quantity would match actual quantity to be labeled and packaged
4439 | 76NYA Biosurgery APU-Dept SEPF-Room 209-5KG weight F-0892 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-209 | 2024-10-25 | 2024-11-19 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4438 | Nanofilter for Formulated Bulk GZ5157-00 lot EGF3B04 failed FIT 3 times. | Alyson.Belleville@sanofi.com | 51 NYA 88 | 2024-11-18 | 2024-11-18 | It is expected that the filter would pass the integrity test. | Yes | Deviation
4437 | pH for SP1 EGF3B08B low during load, wash, and elution on chromatogram. Highest was 4.7 with a range of 4.8 to 5.6 | Alyson.Belleville@sanofi.com | 51 NYA 138 | 2024-10-30 | 2024-11-18 | It is expected that the pH would remain within range throughout the operation. | Yes | Deviation
4436 | Two instances of defective SUTs found over the weekend at 45/51 | Alyson.Belleville@sanofi.com | 51 NYA 134, 45 NYA 190 | 2024-11-17 | 2024-11-17 | It is expected that all SUTs would be without defect.
4435 | 8NYA - RF525E037 705653 Record repeatly going into hold for high or low pressure | bailey.gleason@sanofi.com | 8NYA 1373 | 2024-11-17 | 2024-11-17 | Hold Gemba to see path forward.
4434 | 8NYA- Support Services Gemba-leak on 760741 3K tote | william.lee@sanofi.com | 8NYA/CNC | 2024-11-17 | 2024-11-17 | Media should not have leaked through clamped line
4433 | GEMBA for Leaking Bioburden Sample 45 NYA Buffer | jessica.kychler@sanofi.com | QC Micro 68 New York Avenue, floor 4, room 4700 | 2024-11-16 | 2024-11-16 | It is expected that bioburden samples are not compromised (cracked or leaking in any way) | Yes | Deviation
4432 | 8NYA-Downstream Cz MMC multiple manual entries without verifier | aidan.mula@sanofi.com | 8NYA/RM1212 | 2024-11-16 | 2024-11-16 | Good documentation practices should be followed when performing manual entries | Yes | Deviation
4431 | 8 NYA GEMBA - Support Services - Leaking 3K tote bag | william.lee@sanofi.com | 8NYA/CNC | 2024-11-15 | 2024-11-15 | Tote bag should not have leaked | Yes | Deviation
4430 | 8NYA-Pompe-DGHDA001- PPQ Pre-requisite not completed prior to execution | Catherine.Kuzdzal@sanofi.com | 8 NYA Suite 2 Downstream | 2023-08-25 | 2024-11-13 | It is expected that pre-requisites are completed prior to performing the PPQ batches.
4429 | 8 NYA plates that were delivered today were expired | dennis.bazinet@sanofi.com | 8 NYA sample room pick up area | 2024-11-15 | 2024-11-15 | Samples transported on day they are received
4428 | 8NYA - Support Service - RF607E015 760725 Recipe needs automation assistance to progress record. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-11-15 | 2024-11-15 | Hold Gemba to see path forward.
4427 | 8NYA - Support Services - Incorrect scale used in weighbooth | william.lee@sanofi.com | 8NYA/1350 | 2024-11-15 | 2024-11-15 | Label for weigh scale and actual scale should be the same
4426 | 8NYA - Support Services - Incorrect scale used in weighbooth | william.lee@sanofi.com | 8NYA/1350 | 2024-11-15 | 2024-11-15 | Label for weigh scale and actual scale should be the same
4425 | 8NYA-Gemba-Cz Nanofiltration SOP has incorrect instructions for flowmeter | aidan.mula@sanofi.com | 8RM1503 | 2024-11-15 | 2024-11-15 | SOP should be accurate to operational steps
4424 | 68NYA Support APU-Validation-room 1745-Pressure Transmitter F-15475 found non-operational during a scheduled calibration | robert.bauch@sanofi.com | 68-1745 | 2024-11-11 | 2024-11-15 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | No Deviation
4423 | 8NYA - Support Services - Cleaning contractors observed improperly performing verification steps | christopher.arnold@sanofi.com | 8RM1300, 8RM1301, 8RM1302 | 2024-11-14 | 2024-11-14 | That a verifying operator will be at the place where the work is occurring and able to perform the task of verification.
4422 | 8 NYA Gemba - Support Services - Leaking 20 Liter 760744 Media Discovered in Cold Room | nicholas.grote@sanofi.com | 08-1220 | 2024-11-14 | 2024-11-14 | No leak in bag.
4421 | QU-SOP-0012102-User Administration approved without final authorizer | peter.jansen@sanofi.com | 68NYA | 2024-11-08 | 2024-11-14 | document would be routed with document controller or other as final authorizer
4420 | Pre Use Shipping Bag Inspection not Commented in the Record | zachary.rickard@sanofi.com | 8NYA/1313 | 2024-11-11 | 2024-11-13 | SOP direction will be followed and shipping bag inspection will be completed and documented accordingly
4419 | 8NYA APU-Dept CC8-Speed Indicating controller FB-2936-002 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 8-1333 | 2024-11-13 | 2024-11-14 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4418 | GEMBA - 8 NYA DOWNSTREAM - Incorrect Instructions in CZ CM SOP | adam.keister@sanofi.com | 8 NYA 1503 | 2024-11-13 | 2024-11-13 | For SOP to have correct instructions
4417 | Leak on Process Waste Line for X2560 pre-filter | steven.rampello@sanofi.com | 1403 | 2024-11-14 | 2024-11-14 | It is expected that all welds and inspected and tested for integrity prior to major operations
4416 | pH probe on Bag leaked during standardization | Juan.bonilla@sanofi.com | 8NYA/1503 | 2024-11-13 | 2024-11-13 | No leaks during standardization.
4415 | 8NYA CZ Zinc set up issue, insufficient volume for production run. | christian.galvez-tello@sanofi.com | 8NYA RM1422 CNC | 2024-11-13 | 2024-11-09 | N/A
4414 | 8NYA-Cerezyme-EGHBY021 LAL sample 1486920 CZ2_CM_761222_EQUIL_ENDO sample tube cracked and leaking | larissa.fairweather@sanofi.com | 68 NYA / 4700 / QCM | 2024-11-12 | 2024-11-13 | Per QU-SOP-0070558 step 6.2.3, confirm sample containers are not cracked, leaking. The expectation is that samples are to be free of any cracks or leaks.
4413 | Pinch Clamp leaked during CZ Zinc Load transfer | Juan.Bonilla@Sanofi.com | 8NYA/1503 | 2024-11-13 | 2024-11-13 | Pinch clamp not to leak
4412 | Load Inlet UV transmitter zeroing error during UV Standardization | jared.Watson@Sanofi.com | 8RM1403 | 2024-11-13 | 2024-11-13 | UV transmitter zero'd
4411 | 8NYA-Cz-EGHBN011- MCC Waste stream leak identified in CNC area </= 5gal | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1210 | 2024-11-12 | 2024-11-12 | It is expected that the waste streams associated with MCC activities be connected to the deciated waste tees/ ports located in CNC, in order to avoid unintended release of waste stream | Yes | No Deviation
4410 | Base tote MES issues/ Fz Base Tote Issues | marcelo.dossantos2@sanofi.com | CNC Space | 2024-11-12 | 2024-11-12 | MES allowing for tube attachment |  | No Deviation
4409 | 8NYA - Gemba - Fabrazyme Eluate removed from QI without microbial results. | Timothyryley@sanofi.com | 68NYA | 2024-11-08 | 2024-11-11 | Forward processing statement for QE-089429 should have been updated with passing microbial data before eluate release.
4408 | TSH 45 NYA Buffer Prep - Settle plates for setup portion of hood op were not stopped in Onelims until half way through the operations portion. Additionally, Climet F-13599 was not updated for daylight savings time | jeremy.wooden@sanofi.com | 45 NYA room 190 | 2024-11-12 | 2024-11-12 | It is expected that the settle plates would be scanned to close them out at the end of the set up operation. | Yes | Deviation
4407 | 51 NYA Thyrogen personnel not properly gowned while working in grade C airloc | jacob.kallon@sanofi.com | 51 NYA room 89 | 2024-11-12 | 2024-11-12 | All personnel entering the area must be properly gowned per QU-SOP-0006952 step 6.5.1.
4406 | Cold room 108C at 51 NYA incurred HH temp alarm at 1151 on 11NOV24 | adam.hildebrand@sanofi.com | 51-108C | 2024-11-11 | 2024-11-12 | It is expected that the cold room maintains temperature for storage of material
4405 | Capacitance controller probe AI2 disabled on bioreactor X-2230 | brendan.dickhaut@sanofi.com | 8NYA/Room 1403 | 2024-11-12 | 2024-11-12 | Probes operate within their set-points and do not require disabling due to malfunction.
4404 | 8NYA - Support Services - 705676 Insufficient 5N HYDROCHLORIC ACID to complete weight out. | daniela.medina2@sanofi.com | 8NYA/ 1350 | 2024-11-11 | 2024-11-11 | Performing area clearance will verify sufficient material on hand. | Yes | No Deviation
4403 | 8NYA - Support Services - 705665 missing bag of dry material for buffer batch. | princess.miller@sanofi.com | 8 RM 1350 | 2024-11-11 | 2024-11-11 | That all material be accounted for.
4402 | 8NYA - Support Service - RF552E008 QA kickback regarding manually deallocation of weigh booth. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-11-11 | 2024-11-11 | Hold gemba to see path forward
4401 | Questionable Weld X-2230 daily sub sample/ Questionable Weld on X-2230 | todd.niziak@sanofi.com | 8 NYA 1403 | 2024-11-11 | 145 | Good weld |  | No Deviation
4400 | Growth on Negative Control plate for 76NYA Manufacturing Operations | amanda.lavalley@sanofi.com | 76 NYA Biosurgery | 2024-11-04 | 2024-11-11 | Negative Control Plates are not expected to have growth.
4399 | Incorrect final PI ELISA result reported for TSH UFB | pinky.patel@sanofi.com | 4700 | 2024-09-30 | 2024-11-10 | QU-OPE-0198462, and step 6.12.3 and QU-FOR-0131484, attachment # 4
4398 | Incorrect final PI ELISA result reported for TSH UFB | pinky.patel@sanofi.com | 4700 | 2024-09-30 | 2024-11-10 | QU-OPE-0198462, and step 6.12.3 and QU-FOR-0131484, attachment # 4
4397 | No Product - No Lot. Bioburden and Endotoxin not collected daily as per Action Notic 2024-N-060 | amy.roy@sanofi.com | 76 NYA Room 107B | 2024-11-06 | 2024-11-11 | it is expected that all samples were to be collected for site 222 per 2024-N-060. | Yes | Deviation
4396 | Version 15.0 of QU-OPE-0182570 (Production Record) Approved without required signatories | Amy.McIntyre@sanofi.com | QualiPSO CMS | 2024-11-06 | 2024-11-11 | Expected results are that the workflow initiator adds all required signatories to the workflow.
4395 | SLED?BBD Date needs to be changed in Athena by authorized user to reflect correct expiry of material | adam.hildebrand@sanofi.com | 51 NYA, RM 140 | 2024-10-20 | 2024-11-08 | It is expected that goods receipt be performed in the correct process order at time of collection. | Yes | Deviation
4394 | 8 NYA GEMBA Downstream Fabrazyme - Discrepancy on NANO SOP | alex.tello@sanofi.com | 1323 | 2024-11-11 | 2024-11-11 | To only have 1 instruction for flow meter set up.
4393 | 8NYA - Support Service - RF557E043 705663 FIT was not set at correct PSI, Filter failed BP test twice. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-11-11 | 2024-11-11 | N/A | Yes | Deviation
4392 | Release/T0 Stability Testing No Performed for EGHDT011 FZ ICB Appearance within window | Sarah.Adams@sanofi.com | 68 NYA, QCC | 2024-11-08 | 2024-11-15 | Sample to be tested within stability window.
4391 | Gemba - 8 NYA - Support Serices - 3K SUM Bag (XXX) failed pre use inspection | nicholas.grote@sanofi.com | 08-1350 | 2024-11-09 | 2024-11-09 | Bag passes inspection
4390 | 8 NYA Loose Cz Media Bioburden Sample Cap - Leaking Sample Batch: EGHAV173 | anastacia.hachem@sanofi.com | 68 NYA QC Microbiology | 2024-11-09 | 2024-11-09 | Per QU-OPE-0329679 Step 3. Responsibilities, QC Personnel are supposed to note if there are any issues with samples and notify management. Samples should not be leaking prior to testing as it will compromise the integrity of the sample and/or test result.
4389 | 8 NYA Loose Cz Media Bioburden Sample Cap - Leaking Sample Batch: EGHAV090 | anastacia.hachem@sanofi.com | 68 NYA QC Microbiology | 2024-11-09 | 2024-11-09 | Per QU-OPE-0329679 step 3. Responsibilities, QC Personnel are supposed to note if there are any issues with samples and notify management. Samples should not be leaking prior to testing as it will compromise the integrity of the sample and/or test result.
4388 | 8NYA Downstream ¦ RF643E015 Can not perform buffer take out for CZ Zinc breakdown. | christian.galvez-tello@sanofi.com | 8NYA RM1422 CNC | 2024-11-08 | 2024-11-09 | N/A | Yes | No Deviation
4387 | 8 NYA Gemba - Support Services - Formulation Dry Addition amounts not showing 0.0 as required by SOP | nicholas.grote@sanofi.com | 08-1350 | 2024-11-09 | 2024-11-09 | Working with MES team it appears that the value displayed in the remaining quantity field will not display a negative number. When a material is weighed out above target, but within range, the remaining quantity in the formulation dry addition screen will show 0. When the target weighed out is below target, but within range, it will show the remaining amount to reach the target resulting in this field reading a value > 0. | Yes | No Deviation
4386 | 8NYA Upstream ¦ BHT-2041002 has a broken provibus cable is broken | christian.galvez-tello@sanofi.com | 8NYA RM1210 CNC | 2024-11-09 | 2024-11-09 | N/A | Yes | No Deviation
4385 | 8NYA - Support Services - All Products - Multiple issues with MABC-QU-SOP-0084636 Facilities Disinfection at 8NYA | christopher.arnold@sanofi.com | 8RM1202 (also impacts other CNC spaces impacted by new SOP) | 2024-11-08 | 2024-11-08 | That the EUR for facilities/cleaning will be aligned with the SOP
4384 | Cz Nano Concentration too high | zachary.rickard@sanofi.com | 8NYA 1503 | 2024-11-08 | 2024-11-08 | Values will be consistent with past production runs.
4383 | GEMBA - 8 NYA DOWNSTREAM - Several Steps in CZ Nano SOP Cannot be Executed as Written | adam.keister@sanofi.com | 8 NYA 1503 | 2024-11-08 | 2024-11-08 | To be able to execute SOP as written
4382 | 760741 Storage Unit take out in MES prior to being used. Support Services. | robert.morawski@sanofi.com | 8NYA CNC RM 1210 | 2024-11-08 | 2024-11-08 | Storage Unit 218753 would still exist in tote.
4381 | Annual PM for HEPA Vacuum 5001594 extended overdue (WO# 3005892) | dom.arone@sanofi.com | 8 NYA, RM 1373 | 2023-11-01 | 2024-11-07 | MABC-QU-SOP-0072346-Maintenance Work Order Management Using Infor EAM (v15.0)
4380 | Missing Certificate of Destruction (COD) | dawn.lam@sanofi.com | 76NYA QA | 2024-10-03 | 2024-11-06 | Certificate of destruction (COD) should be provided for evidence of destruction.
4379 | 8NYA GEMBA Upstream FZ SUB Bag Inspection Observation | gary.dodakian3@sanofi.com | 8 NYA upstream room 1333 | 2024-11-08 | 2024-11-08 | Tubing is not present on the inlet of sparge filter | Yes | No Deviation
4378 | 51NYA - Dewpoint failure for site 51-CA-11 | Craig.stevenson@sanofi.com | 51 NYA Room 134 | 2024-11-07 | 2024-11-07 | It is expected that dewpoint testing generate a result of >-40C.
4377 | 8NYA - Support Services - MES/BRR Kickbacks: RF552E008 / 31OCT240029 "Failure to confirm allocation of equipment" | alejandro.martinez2@sanofi.com | 8RM1364 (DFB), 8RM1350 (Refractometer), 8RM1373 (Autoclave) | 2024-11-05 | 2024-11-06 | Equipment should allocate to manufacturing orders, so there is traceability. | Yes | No Deviation
4376 | Credo cooler had no TIC panels in cooler for sample transport | dennis.bazinet@sanofi.com | 68 NYA / sample management drop off room | 2024-11-07 | 2024-11-07 | Samples were transported without TIC panels installed in credo cooler | Yes | Deviation
4375 | 8NYA Gemba - EGHBT014/RF645E014 - Cerezyme Phenyl blank run Justification missing for late entry | stanley.durand@sanofi.com | 8NYA rm1503 | 2024-11-01 | 2024-09-06 | All entries are expected to be contemporaneous | Yes | No Deviation
4374 | 8NYA Gemba - EGHBT014/RF645E014 - Cerezyme Phenyl blank run Justification missing for late entry | stanleydurand@gmail.com | 8NYA rm1503 | 2024-11-01 | 2024-09-06 | All entries are expected to be contemporaneous
4373 | 8 NYA Upstrem GEMBA SUB Train 4 Capacitance probe Disabling path forward | gary.dodakian3@sanofi.com | 8 NYA Upstream room 1403 | 2024-11-07 | 2024-11-07 | capacitance probes must have a strong signal to operate reliably
4372 | Discarding ASM Frozen product SAP Handling Unit Not aligned with MES Storage Unit | joshua.brooks@sanofi.com | Freezers at 8NYA | 2024-11-07 | 2024-11-07 | HUs align for reconciliation
4371 | 8 NYA GEMBA- Downstream - Piece of hair on weld on SUM bag sample port | alex.tello@sanofi.com | 1503 | 2024-11-07 | 2024-11-07 | Bag to pass inspection
4370 | 760721 1M Citric Acid Expiry Issue | nathaniel.barnes@sanofi.com | 8NYA !503 | 2024-11-07 | 2024-11-07 | Correct the way expiry is assigned in the system to align with actual buffer expiry time
4369 | Incorrect Cx Value populated in SUB Batch Record during probe install. | timothy.ryley@sanofi.com | 08NYA 1403 | 2024-10-30 | 2024-11-06 | CX value for Autoclave run should be updated prior to entering performing Probe Inspection | Yes | Deviation
4368 | TSH - Buffer Prep - 45 NYA - When performing magnahelic gauge checks the DP for cold room 174 was at 0 | Jeremy.Wooden@Sanofi.com | 45 NYA cold room 174 | 2024-11-06 | 2024-11-06 | It is expected that the DP for room 174 would be in range
4367 | Analyst Executing Method without being Trained on Effective Version | Sarah.Adams@sanofi.com | 45 NYA, QCC | 2024-06-26 | 2024-11-06 | Analyst to be trained on effective document version prior to execution.
4366 | 51NYA, 148, warm room out of temp | nicole.lufkin@sanofi.com | 51NYA 148 | 2024-11-05 | 2024-11-06 | warm room would be in temp | Yes | Deviation
4365 | DP alarm 051-DPI-9500-029A from RM88 TO RM87 at 51NYA | Sylvester.Kegbeh@sanofi.com | 51 NYA ROOMS 87 amd 88 | 2024-11-06 | 2024-11-06 | There should be no DP alarm in the rooms at 51NYA | Yes | Deviation
4364 | Media tote reading 2234.5L and on deltaV its reading 3007.1L 8NYA CNC | sarah.flattich@sanofi.com | 8 NYA CNC upstream | 2024-11-06 | 2024-11-06 | Expected results on deltaV should have the exact same reading as the loadcell on the media tote
4363 | CZ media tote reading 2234.5L, DeltaV reading 3007.1L 8NYA CNC | sarah.flattich@sanofi.com | 8 NYA CNC upstream | 2024-11-06 | 2024-11-06 | Expected results on deltaV should have the exact same reading as the loadcell on the media tote | Yes | No Deviation
4362 | MEDIA BIRF's not witnessed | nicole.lufkin@sanofi.com | 51NYA 148 | 2024-11-05 | 2024-11-06 | witness step would be signed | Yes | Deviation
4361 | Cracked LAL Cerezyme Buffer POU 8NYA Sample Submitted | Pushti.Shah@sanofi.com | 4700 68 NYA QC Micro | 2024-11-05 | 2024-11-05 | QU-SOP-0014245
4360 | 8NYA - Support Services - Discussion - Lack of instruction for proper use of the Amber clamp alternative(s) | alejandro.martinez2@sanofi.com | 8NYA / CNC | N/A | N/A | Determine if quality event is needed to address lack of instruction for proper use of the amber clamp alternative(s). Confirm that these clamps were properly on-boarded per current procedure. Finally, deciding if these clamps should be removed from the floor, or if they can continue to be used with an interim control to capture where/when/how they should be used. | Yes | No Deviation
4359 | Phenyl SOP Missing Mod Enzyme Disassembly Step | danny.martinez@sanofi.com | 8NYA, RM 1503, Downstream Ballroom 2 | 2024-11-05 | 2024-11-05 | QU-SOP-0080431 needs a Mod Enzyme disassembly step
4358 | ETP F008-ED-ETP-0297 associated with CAPA-010309 was not QA Approved | kelly.desousa@sanofi.com | 8 NYA | 2024-10-25 | 2024-11-04 | MABC-QU-SOP-0001431 section 4 Support, as necessary, development, review, and approval of ETP, as well as associated reports when critical aspects are impacted, and when test results are intended to be leveraged within qualification protocols.
4357 | Buffer GZ2217-51 lot EGFB2046 found to be leaking into secondary container when operators went to use it for processing. | Alyson.Belleville@sanofi.com | 51NYA room 132 | 2024-11-05 | 2024-11-05 | It is expected that a buffer would not leak. | Yes | Deviation
4356 | 8NYA GEMBA Upstream 2250 pH One-Point Executed without Adjusted Value | gary.dodakian3@sanofi.com | 8 NYA Upstream room 1403 | 2024-11-05 | 2024-11-05 | PH Probes are one-pointed with the adjusted value for cerezyme
4355 | 8 NYA Gemba - Support Services - Hole discovered on 3K SUM bag during pre use inspection | nicholas.grote@sanofi.com | 08-1350 | 2024-11-04 | 2024-11-04 | Pass bag inspectoin
4354 | Cerezyme - RF639E013 - 8NYA - ATF Wye connection aseptiquik tear | michael.knapp@sanofi.com | 8NYA 1403 | 2024-11-04 | 2024-11-04 | No tear during connection.
4353 | GEMBA : PI user SOP approved without doc control team | bibhuti.swain@sanofi.com | NA | 2024-10-25 | 2024-10-31 | Approver workflow need to include doc control team per sop (MABC-QU-WIN-0055911)
4352 | 76NYA, Room 410, Site 410, Reverse Osmosis Breach | Craig.stevenson@sanofi.com | 76NYA, Room 140 | 2024-11-04 | 2024-11-04 | The RO system to not be breached
4351 | 8NYA - Support Services - Cerezyme - Conductivity test failed | maria.palmar@sanofi.com | 8NYA / 1350. | 2024-11-03 | 2024-11-03 | Conductivity testing to not fail. | No | Deviation
4350 | 8NYA - Support Service - Cerezyme EGHBV014 - RF647E014_002 Filter was manually launched using the incorrect recipe. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-10-06 | 2024-11-03 | N/A
4349 | 8NYA-Gemba-Fz Downstream Butyl record has multiple exceptions for sealer ID but no verifier | aidan.mula@sanofi.com | 8NYA/RM1323 | 2024-11-03 | 2024-11-03 | Manual entry requires verifier or reference to verifiable source
4348 | X3017 UFDF skid pump will not achieve a set flowrate | pavel.dobrovolskiy@sanofi.com | 8 RM 1323 | 2024-11-01 | 2024-11-01 | Its expected that during the cleaning, skid pump will achieve the necessary pump speed.
4347 | 8NYA - Support Service - RF603E030 760720 Powder bag clip fell into SUM, unable to retrieve. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-11-01 | 2024-11-01 | Hold gemba to see path forward
4346 | Filter test gas not correct 76 NYA EBYFRM049 QU-OPE-0300206 | dianne.moustafa@sanofi.com | 364 76 NYA | 2024-10-29 | 2024-11-01 | QU-SOP-0046034 Attachment 2 states only use of AIR.
4345 | 8NYA APU-Dept PF8-room 1503-Weight Indicator FB-4310-007 overdue for calibration | robert.bauch@sanofi.com | 8-1503 | 2024-11-01 | 2024-11-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
4344 | Cerezyme - EGHBN012 - 8NYA - Leak of pump tubing during ATF wetting | michael.knapp@sanofi.com | 8NYA 1403 | 2024-10-31 | 2024-10-31 | No leaks occur during ATF wetting | Yes | Deviation
4343 | Tampered Cable Seals on Shipper | jonathan.lucas@sanofi.com | 74NYA, rm.166 | 2024-10-31 | 2024-10-31 | Per QU-SOP-0070997, step 6.5.5, the cable seals must match. | Yes | Deviation
4342 | Total amount of precipitate remaining on Tank V-3532 out of specification. | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2024-10-31 | 2024-10-31 | Per QU-OPE-02020445 step 7.12.30 the amount of precipitate remaining on Tank V-3532 should be between 280-320KG | Yes | Deviation
4341 | 8NYA-Cz Zinc load filter clogging | zachary.rickard@sanofi.com | 8NYA 1503 | 2024-10-31 | 2024-10-31 | Pump will run for the duration of the filtration process. Filter will not clog.
4340 | 8NYA-No Product- SUM Bag Inspection Failure and subsequent miss during elevation | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2024-07-25 | 2024-10-30 | It is expected that SUT that fails Mfg inspection be elevated accordingly per procedure, and in a timely manner | Yes | Deviation
4339 | Repligen complaint materials picked up by Repligen from 8 NYA without Athena RTV transactions completed performed | sheila.pandolfo@sanofi.com | 8 NYA Dock | 2024-03-15 | 2024-10-31 | It is expected the Athena transactions required for returning material to the vendor would be performed and Athena inventory would be accurate. This includes the creation of an inspection lot and the return to vendor transaction (122).
4338 | 68NYA Support APU-Metrology standard S-1337 with RTD S-1337-005 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2024-10-22 | 2024-10-31 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4337 | 8NYA-Cz-X2220 Aseptiqwik tear while installing probe set to SUB | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2024-10-30 | 2024-10-30 | It is expected that a succesful completion of this Aseptiqwik connection results in no damage considered impacting the sterility of the tubing
4336 | 8NYA-Cz-EGHCG019- Gowning Miss while entering Class C area | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1343 | 2024-10-30 | 2024-10-30 | It is expected that the correct gowning is donned based on classification of room and activity
4335 | 51NYA - Expired Sporklenz used during BSC cleaning | Craig.stevenson@sanofi.com | 51NYA | 2024-09-01 | 2024-10-30 | It is expected that BSC post cleaning be performed with non expired cleaning reagents. | Yes | Deviation
4334 | Version 1.0 of QU-SOP-0083900 approved in QualiPSO without Document Control as final authorizer | Amy.McIntyre@sanofi.com | QualiPSO CMS | 2024-10-30 | 2024-10-30 | Expected results are that the workflow initiator adds all required signatories to the workflow.
4333 | Gemba 8NYA Upstream MSAT Request to Disable Capacitance Probe 1 on X-2250 due to Offset CZ pn 761217 Lot EGHBN012 | gary.dodakian3@sanofi.com | 8 NYA Upstream room 1403 | 2024-10-30 | 2024-10-30 | Probes remain aligned and do not diverge
4332 | Facilities and NPI: Incorrect Gowning During Tier 1 Action Notice in Solution Prep | elissa.karol@sanofi.com | 8NYA 1350 and 1126 | 2023-08-31 | 2024-10-30 | Gowning SOP QU-WIN-0024475 is followed correctly
4331 | 12L Wave bag gassing was not changed from 10% to 5% CO2 | david.loik@sanofi.com | 51 NYA | 2024-10-21 | 2024-10-30 | gassing would be changed to 5% CO2 based on the pH | Yes | Deviation
4330 | Gemba 8NYA Upstream X-2240 Missed Antifoam Dosing (Friday PM 25OCT4 -Saturday AM 26OCT24 ) | gary.dodakian3@sanofi.com | 8 NYA Upstream room 1403 | 2024-10-26 | 2024-10-30 | Antifoam doses as scheduled per the recipe.
4329 | Potential Product Impact/Product Release: FZ-ICB Aggregation Empower Processing Method incorrectly calculating Aggregation result. | Sarah.Adams@sanofi.com | 68 NYA, QCC | 2024-10-01 | 2024-10-01 | All aggregation peaks to be included in calculation of aggregation result
4328 | GEMBA Upstream CZ Lot EGHBN010 Train 3 Harvest Vessel X-2816 Temperature Oscillation | gary.dodakian3@sanofi.com | 8 NYA Upstream room 1403 | 2024-10-29 | 2024-10-29 | The TCU is suppose to control temperature for the harvest vessel within the range of 15-20 degrees celcius
4327 | GEMBA - 8 NYA DOWNSTREAM - Small leak observed in VI buffer bag after VI buffer addition | adam.keister@Sanofi.com | 8 NYA - 1323 | 2024-10-30 | 2024-10-30 | For buffer not to leak. Path forward on processing. | Yes | Deviation
4326 | Low Osmo result for pre-filtration process test for 760744 Media | christopher.arnold@sanofi.com | 8RM1350 | 2024-10-30 | 2024-10-30 | that the Osmo result would be range. | Yes | Deviation
4325 | 8NYA - Support Service - Failed connection of liquid material to pump tubing 705675 Buffer | alejandro.martinez2@sanofi.com | 8RM1350 Solution Prep suite | 2024-10-29 | 2024-10-29 | No leak should occur at connection from liquid weigh bag and pump tubing. | Yes | Deviation
4324 | 8NYA-Gemba Support Services- Autoclave used when T3 temp probe as founds not completed | william.lee@sanofi.com | 8NYA/1372 | 2024-10-13 | 2024-10-28 | As founds should have been completed by metrology after probes were run.
4323 | Sialic Acid Empower Custom Field Error on disposition of NGNA plate counts | nancy.watkins@sanofi.com | 45NYA, QC Chemistry | 2024-10-28 | 2024-10-29 | Standard Suitability Summary and Conclusion should have read PASSED for NGNA plate count
4322 | Cold Room 147 unplanned outage during scheduled compressor replacement on System B | nicholas.shoenfelt@sanofi.com | 51-147 | 2024-10-28 | 2024-10-29 | It was expected that System A would maintain temperature of cold room 147 at 2-8C throughout the work.
4321 | Non-Sterile Leak was observed during homogenization process on Tank V-3522 | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2024-10-29 | 2024-10-29 | No leaks should have been detected during this process | Yes | Deviation
4320 | Leak on V-3522 at the upper 4 inch sight glass on side port of the vessel. | Phillip.Mank@Sanofi.com | 76 NYA Room 3102 | 2024-10-29 | 2024-10-29 | It is expected that no leak would be observed.
4319 | 8NYA Downstream Fz SP CNC Record missing a manifold 692319 from the BOM | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1323 | 2024-10-29 | 2024-10-29 | To get a Phath forward | Yes | No Deviation
4318 | rooms 1323/ 1324 were not cleaned per AN 2024-I-060 | pavel.dobrovolskiy@sanofi.com | 1323 , 1324 | 2024-10-29 | 2024-10-29 | its expected that all rooms that are outlines in AN 2024-I-060 will be cleaned every night | Yes | No Deviation
4317 | Gemba-8NYA- Cz Downstream MMC skid CVed post-calibration but Zinc was not | aidan.mula@sanofi.com | 8NYA/RM1503 | 2024-10-29 | 2024-10-29 | Determination of impact at Gemba
4316 | Reagents Not Logged correctly for ID gram stains | kyle.lebherz@sanofi.com | 68 NYA 4700 QC Micro | 2024-10-04 | 2024-10-29 | for all reagents to be logged appropriately and for dates of log entries to be logged correctly | Yes | Deviation
4315 | Failed Weld on Sampling Line | ana.salichs@sanofi.com | RM1403 | 2024-10-29 | 2024-10-29 | Expectation is for weld to be integral and pass inspection.
4314 | 8NYA Gemba Downstream - Record not finished at time of activity | alex.tello@sanofi.com | CNC | 2024-10-24 | 2024-10-28 | For record to be executed until its finished | Yes | No Deviation
4313 | Cz MMC Eluate SUM Temp Excursion | nathaniel.barnes@sanofi.com | 8 NYA 1503 | 2024-10-29 | 2024-10-29 | Determine effect of temp excursion on product stored in SUM
4312 | Fz Zinc Eluate SUM Temp Excursion | nathaniel.barnes@sanofi.com | 8NYA 1323 | 2024-10-29 | 2024-10-29 | Determine effect of temp excursion on product stored in SUM
4311 | 8NYA - Support Services - Batch with expired timer | william.lee@sanofi.com | 8NYA/1350 | 2024-10-28 | 2024-10-28 | 10 hour batch timer expired for 705663 0.5M NaOH, 1M NaCl. | Yes | Deviation
4310 | X2250 DO controller HIHI alarm after innoculation | ralston.najarro@sanofi.com | 8NYA1403 | 2024-10-28 | 2024-10-28 | DO should be controlled to setpoint
4309 | TSH-45 NYA- Buffer prep - Onelims Lots generated from Athena are being cancelled and are not available for use on the day of the operation. | Jeremy.Wooden@Sanofi.com | 45 NYA Room 190 | 2024-10-28 | 2024-10-28 | It is expected that our samples would be available for collection in the Onelims system on the day of the operation. | Yes | No Deviation
4308 | EGF4B09 / H-11 MYCO and Vesi samples | kevin.devanna@sanofi.com | 51 New York Ave, RM 140 | 2024-10-07 | 2024-10-08 | Sample the correct lot | Yes | No Deviation
4307 | 8NYA - Support Service - PI Historian has not been receiving data since 15OCT24 for meters 2009226 PH and 2008025 Condo | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-10-15 | 2024-10-28 | Hold Gemba to see path forward.
4306 | Second page of QU-FOR-0012699 sample delivery start time, end time and performed by steps not filled out | leon.mamish@sanofi.com | 45 New York Ave | 2024-10-23 | 2024-10-28 | Both pages used should have the start, end and performed by portions filled out. |  | No Deviation
4305 | Pompe Seed Train SOP not aligned with the CST | kristine.caron@sanofi.com | 8NYA | 2024-09-19 | 2024-10-25 | Instructions should reflect the Pompe process |  | No Deviation
4304 | 8NYA - Upstream - EGHCE004 Butyl Eluate Weight Discrepancy After Tote Transport | wilfredo.lopez@sanofi.com | 8RM1333 | 2024-10-27 | 2024-10-27 | Expected that the weight remains the same before and after SUM transport
4303 | Cerezyme - RF639E010 - 8NYA - MCC UV Standardization Recipe required automation intervention | michael.knapp@sanofi.com | 8NYA 1403 | 2024-10-27 | 2024-10-27 | UVs are zeroed during recipe and remain within zero'ed tolerance. | Yes | Deviation
4302 | 8 NYA - Support Services - Leaking Tote Bag 760741 Media | nicholas.grote@sanofi.com | 08-1202 | 2024-10-26 | 2024-10-26 | Tote bag does not leak
4301 | Cerezyme 8NYA Misaligned registration plate on 3K growth media tote | michael.knapp@sanofi.com | 8NYA 1403 | 2024-10-27 | 2024-10-27 | Passing bags
4300 | GEMBA - 8 NYA - DOWNSTREAM - CNC Breakdown Step Not Signed for in Record | adam.keister@sanofi.com | 8 NYA 1503 | 2024-10-27 | 2024-10-27 | For step to be signed when performed
4299 | GEMBA - 8NYA - Downstream - Expired SUT used for F008-ED-ETP-0306. | kenny.leland@sanofi.com | 8 NYA 1503 | 2024-10-26 | 2024-10-26 | It is expected not to use expired SUT.
4298 | 8NYA - Support Services - Multiple Totes with Leaking Seals | wilfredo.lopez@sanofi.com | 8RM1210 | 2024-10-26 | 2024-10-26 | Expected that all seals remain integral | Yes | Deviation
4297 | GEMBA - 8 NYA - DOWNSTREAM - Wrong clamps are in FZ SOP instructions for installing nano filter | adam.keister@sanofi.com | 8 NYA 1323 | 2024-10-25 | 2024-10-25 | For SOP to have accurate instructions. | Yes | Deviation
4296 | Bent KLEENPAK Connection on 100L Buffer Bag GZ6581-01 | jeff.hill@sanofi.com | 45 NYA RM 190 | 2024-10-25 | 2024-10-25 | QU-SOP-0019608, 6.17
4295 | 8NYA- No Product- Potential SUB Bag Inspection failure for ATF Port Tubing Length | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2024-10-22 | 2024-10-22 | It is expected that the ports used to connect ATF10s for perfusion activities be of adequate length to allow for Amber Clamp placement without damaging the bag | Yes | Deviation
4294 | Fz media PN 705691 Cx value expiry assigned 60 days instead of 90 per MABC-QU-OPE-0455268 | Jacob.Roman@sanofi.com | 8RM1343 | 2024-10-25 | 2024-10-25 | The media is expected to have its Cx expiry value set at 90 days | Yes | No Deviation
4293 | Filter flow rates for SHMW lots are not being recorded as per QU-OPE-0309296 695L Sodium Hyaluronate Fermentation at 76 NYA | michael.purple@sanofi.com | 76NYA, Room 364 | 2024-10-25 | 2024-10-25 | Flow rates should have been captured per Step
4292 | 8NYA - Support Service - Target quantity of raw material changed from MBR V8 to MBR V9 for 705675 Batches | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-10-25 | 2024-10-25 | Hold Gemba to see path forward.
4291 | 8NYA Upstream - Cerezyme 500L Bioreactor Capto Butyl Eluate Collection Lot DGHBN012 Action Level Endotoxin | amy.roy@sanofi.com | 68 NYA 4700 QCM | 2023-09-27 | 2024-10-24 | It is expected that the correct specifications are utilized when determining valid results
4290 | X2200 SUB TCU low temp alarm | Yves.Antoine@sanofi.com | 8NYA, Room 1333 | 2024-10-23 | 2024-10-23 | It is expected SUB TCU to stay within spec thought the Fz run
4289 | UF/DF Collection sum mixing recipe not moving forward. | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1323 | 2024-10-23 | 2024-10-23 | To get a Phath forward | Yes | Deviation
4288 | 8NYA - Support Services - Fz Media stock creation takes uncorrected tote value to create stock | christopher.arnold@sanofi.com | 8RM1202 (solution pre CNC space) | 2024-10-24 | 2024-10-24 | That the EBR would create stock based upon the corrected tote weight (weight corrected above the stock creation step in the EBR)
4287 | Ineffective CAPA FRAD23A0389 | amanda.sorel@sanofi.com | 8 NYA | N/A | 2024-10-23 | Its is expected that the CAPA would reslove the problem.
4286 | 8NYA - Support services - DSP record RF1BRE038 is launching the incorrect recipe for the butyl load filter. | william.lee@sanofi.com | 8NYA/1350 | 2024-10-23 | 2024-10-23 | The correct recipe for butyl load should have been ran instead of butyl eluate. | Yes | Deviation
4285 | 51NYA Thyrogen APU-dept Purification-room 132-Delta V temperature indicator F-4236-012 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 51-132 | 2024-10-23 | 2024-10-23 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4284 | Custom field does not calculate the average of two injections for G_Recv% following the change to single injections per line | jacqueline.rapko@sanofi.com | N/A, affects QC Chem, for Cerezyme AA Oligo | 2023-2024 | 2024-10-22 | Average of two lines G-Recv%.
4283 | X2831 blank run 2 in hold, High pressure on column inlet | ralston.najarro@sanofi.com | 8NYA1403 | 2024-10-22 | 2024-10-22 | There should not be a high pressure event on the column
4282 | 8 Newyork Ave - Support Services - Incorrect Stock created for Media batch | princess.miller@sanofi.com | 8 Newyork Ave. Rm# 1202 | 2024-10-23 | 2024-10-23 | Two Mo's were being executed back and forth. Had the wrong Mo opened when pulling weight.
4281 | Onelims late entry for sampling event | david.loik@sanofi.com | 51 NYA | 2024-10-21 | 2024-10-21 | samples would be signed for in biolims at time of physical sampling
4280 | 76-EF1A/2A failure to start on swap over , lead/lag | john.beaule@sanofi.com | 76-224, 227 and 229 | 2024-10-22 | 2024-10-22 | FBS-SOP-000747, area pressures | Yes | Deviation
4279 | WFI/ Heat Exchanger Flushes not performed weekly | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2024-10-22 | 2024-10-22 | Per QU-SOP-0005873 step 6.4 it is expected that the operators perform weekly WFI/ Heat Exchange Flushes and document using Attachment #2
4278 | No product- No Lot- 8NYA- QA Contained Equipment form were removed without completion | Daniel.bati@sanofi.com | 8NYA | 2024-09-12 | 2024-10-22 | Forms to be completed on removal | Yes | Deviation
4277 | Upstream, CZ, RF639E007, TCU failure on SUB X2220 | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-10-21 | 2024-10-21 | TCU not in alarm state controlling the SUB temp
4276 | Cleaning Logbook QU-FOR-0042619, version 5.0, approved in QualiPSO without document control | Amy.McIntyre@sanofi.com | QualiPSO CMS | 2024-06-17 | 2024-10-21 | Expected results are that the workflow initiator adds all required signatories to the workflow.
4275 | Calculation Error in QU-OPE-0064308 | timothy.ryley@sanofi.com | 68NYA | 2024-10-21 | 2024-10-21 | Report should have correct calculations. | Yes | Deviation
4274 | 51NYA: Leak test failed on autoclave | adam.rowe@sanofi.com | 51NYA, rm129 | 2024-10-20 | 2024-10-20 | Leak test should pass. This is part of the fit for use performed any day an autoclave run is performed.
4273 | Wave B stopped overnight during media warming stage | nicholas.shoenfelt@sanofi.com | 51-152A | 2024-10-21 | 2024-10-21 | It is expected that the wave remaining powered and the recipe running. | Yes | Deviation
4272 | Upstream, CZ, RF639E010 Blank run 1 Hi Hi pressure alarm on MCC X2831 | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-10-21 | 2024-10-21 | system to maintain flow without a gradual increase in pressure | Yes | No Deviation
4271 | Gemba-8NYA-Downstream Fz-SP SUM bag failed inspection | lukus.palzkill@sanofi.com | 8NYA downstream suite 1 room 1323 | 2024-10-19 | 2024-10-19 | Inspect bag film for abnormalities such as deep scratches or pinholes that cling to the glove at touch.
4270 | 8NYA - Support service - DSP record RF1BRE037 is launching the incorrect recipe for the butyl load filter. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-10-20 | 2024-10-20 | hold gemba to see path forward
4269 | 8NYA - Support Service - MES system allowed for acknowledgment of cluster that was generated by the same operator. RF551E022 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-10-20 | 2024-10-20 | hold gemba to see path forward
4268 | Missing EOR SUB samples for EOR EGHBN006 | phoebe.karkos@gmail.com | 8NYA, QC at 68NYA | 2024-10-19 | 2024-10-19 | Eluate day samples were also taken from this sub yesterday for the same test as EOR, can the results from the eluate day sample tests be used for EOR?
4267 | Missing witness by signature of BIRF for TSK EGF2B03B | Gaelle.Nana@sanofi.com | 51NYA | 2024-08-09 | 2024-10-19 | Movement of product needs to be performed by and witness by two operators, QU-SOP-0074092 Attachmnt 1: BIRF
4266 | 8NYA GEMBA - Downstream. SUM bag leaking out of PH probe. | alex.tello@sanofi.com | 1323 | 2024-10-19 | 2024-10-19 | For PH probe not to leak and standardize correctly per EWI-10
4265 | 8NYA - Support Services - RF567E019 - 705675 refractometer value out of spec per MABC-QU-OPE-0224707 | christopher.arnold@sanofi.com | 8RM1350 | 2024-10-19 | 2024-10-19 | That the tested refractive index would be within the acceptable process range. |  | No Deviation
4264 | Women's Locker Room Badge Reader not Working at 8NYA | tanni.arif@sanofi.com | 8NYA Room 1124 (badge reader # 14A_07) | 2024-10-18 | 2024-10-18 | Badge readers expected to work to ensure secured access control | Yes | Deviation
4263 | GEMBA 8NYA Upstream FZ EGHCE004 Eluate 7 weight changed from 93kg at The Initial Collection point to 80kg When Moved to Butyl Parking Spot | gary.dodakian3@sanofi.com | 8NYA room 1333 | 2024-10-18 | 2024-10-18 | the eluate sum should read approximately the same weight in the buytl parking spot as it did in the collection spot
4262 | Quality compliance specialist did not document times on the Cryostorage Unit Key Usage Log at the time of use | timothy.pyne@sanofi.com | 51-1073 | 2024-10-18 | 2024-10-18 | It is expected that QU-FOR-0117767 is
4261 | GEMBA: 8NYA - CU- WFI Site 4-11, Room 1170 Mech Space leaking during sampling | judith.keller@sanofi.com | 8nya/room 1170 | 2024-10-18 | 2024-10-18 | Expected that site does not leak. | Yes | Deviation
4260 | 51 NYA rhTSH APT DS-Gemba: Document not routed to RM during Qualipso approval work flow | jacob.kallon@sanofi.com | 51 NYA | 2024-10-17 | 2024-10-17 | Per QU-WIN-0055911 step 6.3.5: "Verify the Final Authorizer box is checked and a member of the Doc Control Team is assigned"
4259 | 8NYA - Support service - RF567E018 - 705675 refractometer value out of spc per MABC-QU-OPE-0224707 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-10-18 | 2024-10-18 | hold gemba to see path forward
4258 | Gemba-8NYA-Downstream Fz UFDF Membrane Hydraulic Pressure lost overnight on cassette holder | aidan.mula@sanofi.com | 8RM1323 | 2024-10-18 | 2024-10-18 | Cassette holder should maintain pressure | Yes | Deviation
4257 | 8 NYA - Support Services - Bag Defect in 3K SUM Bag | christopher.brewster@sanofi.com | 8 NYA, Rm. 1350 | 2024-10-18 | 2024-10-18 | It is expected that there are no defects in the SUM bags. | Yes | Deviation
4256 | 8NYA-Fz-EGHCE004- Suspect Terumo weld while sampling X2210 for Eluate Day sampling activities | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2024-10-17 | 2024-10-17 | Weld Inspection yields passing welds
4255 | Atypically low LOD (Vac) result of 1.14% at 12 month stability time point for Sterile HA Powder Lot DBYHAS005 PN 5615-30 | yekaterina.astakhov@sanofi.com | QCBP Lab, Stability (NBO) | 2024-10-17 | 2024-10-17 | It is expected that LOD result is not atypical
4254 | Record stuck on "tare sub weight" after D0 1 point | joseph.gonzalez@sanofi.com | 8NYA | 2024-10-17 | 2024-10-17 | Sub tared correctly. | Yes | Deviation
4253 | Incorrect entry of HU for conductivity standard in Meter standardizing record | andrew.teixeira@sanofi.com | 8NYA/1703 | 2024-10-01 | 2024-10-17 | HU should have been entered instead of part number and manual entry should be verified by a second operator.
4252 | Fire Alarm sounded at 8NYA | justin.carpenter@sanofi.com | 8 NYA | 2024-10-17 | 2024-10-17 | Fire Alarm should not sound during fire pump testing | Yes | Deviation
4251 | 8 NYA Gemba - Support Services - 705683 Buffer pH out of range | Nicholas.grote@sanofi.com | 08-1350 | 2024-10-17 | 2024-10-17 | pH in range
4250 | 8NYA APU-Dept CC8-room 1403. flow indicators FB-4220-012 and FB-4220-013 not calibrated to the full range | robert.bauch@sanofi.com | 8-1403 | 2024-10-15 | 2024-10-17 | N/A
4249 | 8 NYA APU-CC* room 1403-flow indicators FB-4220-012 and FB-4220-013 not calibrated to the bottom end of the calibrated range | robert.bauch@sanofi.com | 8-1403 | 2024-10-16 | 2024-10-17 | Calibration should have been performed to the approved calibration range | Yes | Deviation
4248 | 8NYA-FZ-RF503E004 Multiple EBR steps executed without verification of manual data entry/ work performed | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2024-09-02 | 2024-10-16 | Manual data gets verified by a second operator per procedure | Yes | Deviation
4247 | 8NYA Upstream GEMBA PN 713983 Bag opened with Cap Off Tube, Part Needed for MCC Sip for Train 3 | Gary.dodakian3@sanofi.com | 8nya Framingham | 2024-10-16 | 2024-10-16 | The expectation is the cap should be on tubing when opened.
4246 | GEMBA - 8 NYA DOWNSTREAM-missing post nanofiltration retain sample | lukus.palzkill@sanofi.com | 8NYA downstream suite 1 post viral | 2024-10-14 | 2024-10-15 | MABC-QU-OPE-0028959 FZ sampling plan for 8NYA take 3x2ml samples aliquot into polypropylene tubes | Yes | Deviation
4245 | TSK EGF3B04A Temperature out of range during process at 51NYA | Sylvester.Kegbeh@sanofi.com | 51 NYA ROOMS 132 | 2024-10-15 | 2024-10-15 | It is expected that temperature stay in the range of 6-10C per SOP-0061107
4244 | 68 NYA Thermometers F-16515 and F-16517 Broke upon being Dropped in QC Micro Lab | Rebecca.Schmidt@sanofi.com | 68NYA/4700/QCM | 2024-10-12 | 2024-10-12 | QU-SOP-0069665 step 5.2.3.6 states the following: If unable to collect as-found data for GMP CRIT, an event must be recorded in the appropriate quality system.
4243 | 8 NYA Upstream Gemba For Atypical Trend On Cz X-2250 SUB, Batch EGHBN006 | todd.niziak@sanofi.com | 8 NYA Phase 2 | 2024-10-14 | 2024-10-14 | pH should be in criteria
4242 | Gemba-8NYA-Downstream Cz Phenyl-Floor Scale Standardization manual entry missing verifier | aidan.mula@sanofi.com | 8RM1503 | 2024-10-15 | 2024-10-15 | Manual entry requires verifier but also our PI-connected equipment should have traceable data. | No | Deviation
4241 | More than three (3) occurrences found in the lookback period for AV OOS | juan.ramos-torres@sanofi.com | N/A | 2024-10-15 | 2024-10-15 | CAPAs should have eliminated root cause
4240 | TSH- Buffer prep HEPA filter leaking clear fluid in 189B at 45 NYA | Jeremy.Wooden@Sanofi.com | 45 NYA room 189B | 2024-10-15 | 2024-10-15 | It is expected that there would be no clear fluid dripping from the HEPA grates.
4239 | 1425892 Negative Control 8 NYA Room: 1343 Site: FB-4529 EGHCG011 had 1(1) CFU mold | Stefanie.Erickson@sanofi.com | QC Micro counting | 2024-10-06 | 2024-01-27 | for the negative control to be 0 CFU
4238 | 45NYA Thyrogen APU-Room 169A-2 pressure indicators on Autoclave #6 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 45-169A | 2024-10-11 | 2024-10-15 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4237 | FZ manitol buffers were stored at an incorrect temp | pavel.dobrovolskiy@sanofi.com | 8nya 1323 | 2024-10-14 | 2024-10-14 | Its expected that the buffers are stored based on their temp requirement listed on the label | Yes | Deviation
4236 | X2240 DO not in range due to N2 sparge | ralston.najarro@sanofi.com | 8NYA1403 | 2024-10-14 | 2024-10-14 | DO should be controlled to setpoint
4235 | New Clamps utilized by SP for Nano filter do not align with DS EWI (cap confirmation) | pavel.dobrovolskiy@sanofi.com | 8NYA -1323 | 2024-10-15 | 2024-10-14 | Its expected that DS EWI will be edited to align with the new SP procedure of attaching plastic claps and leaving no gap | Yes | Deviation
4234 | CZ - EGHBN011 - 8NYA - Failed inspection of SUB bags | michael.knapp@sanofi.com | 8NYA 1403 | 2024-10-13 | 2024-10-13 | Passing bags | Yes | No Deviation
4233 | Insufficient process air flow during blowdown on Fz UFDF cleaning. | kenny.leland@sanofi.com | 8 NYA 1323 | 2024-10-13 | 2024-10-13 | It should be reading 10 psi while in blowdown at outlet of PCV-3017068
4232 | FZ - EGHCE004 - 8NYA - Leak on sample bag luerlock during sampling of SUB X2210 | michael.knapp@sanofi.com | 8NYA 1333 | 2024-10-13 | 2024-10-13 | Integral sample bag
4231 | 8NYA - Support Services, Broken Nano filter upon inspection | Jack.Laplante@sanofi.com | 08-1373 | 2024-10-13 | 2024-10-13 | Bleed valve is not broken, filter is in good condition.
4230 | 8NYA - Support Services, Broken Nano filter upon inspection. | JACK.LAPLANTE@SANOFI.COM | 8NYA 1373 | 2024-10-13 | 2024-10-13 | Path forward
4229 | 8NYA-Fz EGHCE004- Weight Excursion X2210 while monitoring for foam | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2024-10-11 | 2024-10-12 | It is expected the weight remain within allowable values per MABC-QU-OPE-0524531-Control Strategy for Fabrazyme Integrated Continuous Biomanufacturing (Fz-ICB) Process (95-105kg)
4228 | GEMBA - 8 NYA DOWNSTREAM - Multiple Automation Issues During FZ Zinc Production | adam.keister@sanofi.com | 8 NYA, 1323 | 2024-10-12 | 2024-10-13 | For recipes to run without automation assistance | No | Deviation
4227 | GEMBA- Downstream- Discolored UF PRODUCT PUMP INSERT | katherine.reardon@sanofi.com | 8 NYA 1323 | 2024-10-12 | 2024-10-12 | installing defect free SUT per QU-SOP-0078707 Step 6.4.1 Connect pump SUT (745389) to pressure switch
4226 | GEMBA - 8 NYA DOWNSTREAM - TCU alarm preventing zinc load from finishing | adam.keister@sanofi.com | 8 NYA 1323 | 2024-10-12 | 2024-10-12 | Resolution of TCU or advancement of recipe to allow load to proceed
4225 | 8NYA - Damaged SUT (3K SUM bag) Discovered by NTO Support Services | JACK.LAPLANTE@SANOFI.COM | 8NYA RM1350 | 2024-10-12 | 2024-10-12 | Path forward, add to SUT failure | Yes | No Deviation
4224 | 8 NYA GEMBA - Support Services - Bag defect noticed during SUM bag inspection | christopher.brewster@sanofi.com | 8 NYA, Rm. 1350 | 2024-10-12 | 2024-10-12 | It is expected that there are no defects in the SUM bags.
4223 | Incorrect Dilution Scheme used to calculate Thyrogen A280 Results, Lot EGF3B04B, not caught upon review | Kaylyn.Manoogian@sanofi.com | QC Chemistry; 45 New York Avenue | 2024-09-30 | 2024-10-11 | Per MABC-QU-OPE-0107274, "A280 of Thyrogen" step 6.5.5, a 1x / 2x dilution scheme is recommended for TSK Butyl Eluate. Testing Analyst, Anna Moriarty, did use the correct dilutions of 1x / 2x and appropriately documented on the assay worksheet. Upon entry of final results into Attachment 1: A280 of Thyrogen Worksheet, the incorrect dilutions of 10x / 15x were entered for rhTSH TSK Eluate, Lot: EGF3B04B, PN: GZ5153-00, QC ID: 24CS00033, causing the results to calculate incorrectly.
4222 | GEMBA - 8 NYA DOWNSTREAM - Corrections Needed For UFDF SOP | adam.keister@sanofi.com | 8 NYA, room 1323 | 2024-10-11 | 2024-10-11 | To be able to perform SOP as written | Yes | Deviation
4221 | A weekly managerial visit is required in QU-SOP-0075222 step 6.5 and was not completed by the required 04Oct2024 due date | randall.porcella@sanofi.com | 8NYA | 2024-10-04 | 2024-10-10 | A weekly managerial walkthrough would be completed and documented in WO 3260928 | No | Deviation
4220 | Leaking pH probe | katherine.reardon@sanofi.com | 8 NYA Room 1323 | 2024-10-10 | 2024-10-10 | FBL-EWI-000015 step 4 sub-step 4 pH probe inspection. | Yes | Deviation
4219 | X-2250 Antifoam Spinner Plate Found Off | brian.o'rourke@sanofi.com | 8 NYA Room 1403 Upstream Phase 2 | 2024-10-11 | 2024-10-11 | Per QU-SOP-0080429 step 6.10.12: Place spinner flask on spinner plate and initiate agitation at 70RPM. Agitation is initiated to maintain homogenous mixture of antifoam with no clumps and to avoid settling. | No | Deviation
4218 | 76 NYA - Bioburden WFI drop sample for V-3522 dropped during sample preparation | megan.botti-marino@sanofi.com | QC Micro lab/68 NYA | 2024-10-09 | 2024-10-09 | It is expected that bioburden samples be kept in tact and tested.
4217 | TSH - Buffer Prep - 2 x each GZ6581-01 bags have bent aseptic connections preventing pairing with the manifold | Jeremy.Wooden@Sanofi.com | 45 NYA Room 190 | 2024-10-10 | 2024-10-10 | It is expected that the connection would not be bent and would lock into place with the corresponding manifold. | Yes | Deviation
4216 | 8 NYA GEMBA - Support Services - Leaking 3K SUM bag | william.lee@sanofi.com | 8NYA/1350 | 2024-10-10 | 2024-10-10 | The SUM bag should not be leaking | Yes | Deviation
4215 | 8NYA- Support APU-Room 1170-Facilities-TOC FB-4002-005 and FB-4002-096 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1170 | 2024-10-07 | 2024-10-10 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
4214 | 8NYA APU-Dept CC8-Room 1403-Flow indicators FB-4220-047 and FB-4220-048 found out of tolerance during a unscheduled calibration | robert.bauch@sanofi.com | 8-1403 | 2024-08-12 | 2024-10-10 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4213 | 76 NYA Biosurgery-HAFM-Room 364-Agitator Indicators F-4274-022 and F-4274-024 went overdue for calibration | robert.bauch@sanofi.com | 76-364 | 2024-09-01 | 2024-10-09 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
4212 | mixer did not turn on after ammonium sulfate addition | Marc.Latta@sanofi.com | 1333 | 2024-10-09 | 2024-10-09 | PUMP MUST TURN ON AS SOON AS THE TIMER STARTS POST AMMONIUM SULFATE ADDITION | Yes | Deviation
4211 | 8NYA Downstream Fz SP column particle found on Slurry vessel after unpack | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1303 | 2024-10-09 | 2024-10-09 | To get a Phath forward | Yes | Deviation
4210 | 8NYA Support services BHT-2042006 scale not reading. | genesis.torres2@sanofi.com | 8nya CNC | 2024-10-09 | 2024-10-09 | Media install on trains 5&7 | Yes | No Deviation
4209 | 8NYA-Support Services- Sartorius Biosealer TC Usage in Non-Media EBRs | alejandro.martinez2@sanofi.com | Solution Prep 8RM1350 & CNC | 2024-10-08 | 2024-10-08 | ITSCR needed to continue referencing usage of Sartorius Biosealer TC until MBRs for all buffers get updated to allow for this equipment. | Yes | No Deviation
4208 | Q Chrom UV PM not created in INFOR | cindy.guerra@sanofi.com | 51 - 131 | 2024-10-08 | 2024-10-08 | A like for like PM code should have been established in INFOR for the FE UV Meter that was installed. | Yes | Deviation
4207 | zero filter didn't parse on 03SEP24 for 125mL flask prep operation | nicole.lufkin@sanofi.com | 5 1NYA | 2024-09-03 | 2024-10-09 | zero filter would parse with other data | Yes | Deviation
4206 | 8NYA Downstream Fz Sp Column failed Qualification | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1303 | 2024-10-09 | 2024-10-09 | To get a Phath forward
4205 | Multiple Issues with 760721 1.0M Citric Acid buffer | christopher.arnold@sanofi.com | 8RM1350 | 2024-10-09 | 2024-10-09 | That the buffer QS automation will dispense the correct amount of WFI. That the OSMO EWI would have a table or list of the standards needed for each buffer/intermediate that requires OSMO testing. That the buffer specification document details will align with the buffer EBR. | Yes | Deviation
4204 | QS WFI addition for 760721 1.0M Citric Acid triggered the 1% deviation limit (high) | christopher.arnold@sanofi.com | 8RM1350 | 2024-10-09 | 2024-10-09 | That the QS WFI addition will finish within tolerances and that the recipe will automatically end.
4203 | NeoGAA Aggregation Assay Worksheets not Completed Contemporaneously | Kaylyn.Manoogian@sanofi.com | QC Chemistry; 68NYA Laboratory 2700 | 2024-10-04 | 2024-10-07 | Assay Worksheet would've been completed contemporaneously on 04Oct24 during assay performance.
4202 | 8NYA QCAs within the EMS system loss of signal/communication | Gary.dodakian3@sanofi.com | 8nya Framingham | 2024-10-08 | 2024-10-08 | The expectation is there should be no loss of communication for ems causing qca alarms | Yes | No Deviation
4201 | EBYHAM009 QU-OPE-0202445 several issues QU-SOP-048453 missing verifier signature | dianne.moustafa@sanofi.com | 3102 76 NYA | 2024-09-24 | 2024-10-08 | 1. Record actual gauge number. 2. QU-OPE-0202445 has an RPM range of 145-155. 3. Expect Verifier signature.
4200 | Addtional Mini Eluate added to MCC skid 2731 End of the run prompt. | bhargav.patel@sanofi.com | 8 NYA / 1403 | 2024-10-07 | 2024-10-09 | Mini Eluate counter should be 126, it should not increase to 131.
4199 | EWI-000043 RoSS Fill Freeze Bag Update | katherine.reardon@sanofi.com | 8 NYA | 2024-10-08 | 2024-10-08 | update to EWI-000043 | Yes | No Deviation
4198 | 8 NYA, required amount of non-viable samples not collected for operations in collection batch 202408597-1, 26AUG24 | megan.botti-marino@sanofi.com | 8 NYA | 2024-08-26 | 2024-10-08 | Per step 6.7.5 of MABC-QU-SOP-0030277, at least 10 non-viable samples (1 cubic meter of air) needs to be collected during operations.
4197 | 8NYA-Fz-EGHCE004- Antifoam Replacement Step not documented at time of execution | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2024-10-05 | 2024-10-08 | Contemporaneous documentation of cGMP work performed | Yes | No Deviation
4196 | 8NYA-No Product- SUB Bag Inspection Failure P/N 753288 | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2024-10-07 | 2024-10-07 | It is expected for vendor supplied material pass internal inspection criteria
4195 | SP column pack failed qualification | zachary.rickard@sanofi.com | 8NYA/1303 | 2024-10-07 | 2024-10-07 | Asymmetry values should be within the defined range.
4194 | GEMBA - 8 NYA DOWNSTREAM - Path Forward Needed for SP Column Packing | adam.keister@sanofi.com | 8 NYA 1303 | 2024-10-07 | 2024-10-07 | Decision on how to proceed. | Yes | No Deviation
4193 | Navigo Product Loading PAR in the Control Strategy is not aligned with FMEA report | ilene.robinson@sanofi.com | 8NYA | 2024-04-10 | 2024-10-04 | Per Table 4 of QU-OPE-1945092 Navigo load density range is 4.0-17.0, the CST QU-OPE-0522556, Table 8-9 has a product loading PAR of 4-17g/L. The correct loading density PAR is 4.0-17.0 as document in the FMEA report. | Yes | Deviation
4192 | 8NYA-DOWNSTREAM-CZ CM-RF649E014/EGHBY0014-ELUATE SUM NOT CHILLING THE PRODUCT | Marc.Latta@sanofi.com | 1503 | 2024-10-06 | 2024-10-06 | PRODUCT SHOULD BE CHILLING AT A CONSTANT TEMPERATURE BTWEEN 2-8 DEGREE CELCIUS PER FBL-CST-000029
4191 | HBOD Waste tote SUT leak | leo.gillespie@sanofi.com | 1205 CNC East corridor - MCC skid tube raceway | 2024-10-06 | 2024-10-06 | possible SUT failure - bag should not leak
4190 | 8NYA - Support Service - RF615E009 760736 for bad seal on 1K tote | william.lee@sanofi.com | 8NYA/CNC | 2024-10-05 | 2024-10-05 | The seal should be intact after sealing with sartorius sealer. |  | No Deviation
4189 | 8NYA - Support Service - RF1AKE031 706210 for Q is missing from Material transport | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-10-05 | 2024-10-05 | hold gemba to see path forward | Yes | No Deviation
4188 | Slurry bag failed visual inspection during Column packing activities | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1303 | 2024-10-04 | 2024-10-04 | To get a Phath forward | No | Deviation
4187 | MCC X2510 in hold due to column 2 EM not aquired. | brendan.dickhaut@sanofi.com | 8NYA/Room 1333 | 2024-10-05 | 2024-10-05 | TCU operates as expected and maintains set point temperatures of bioreactor. | No | Deviation
4186 | Action Notice executed on printed copy | steven.erban2@sanofi.com | 8NYA 1323 | 2024-10-02 | 2024-10-03 | Action notice should have been picked up and signed for in log, and executed on approved form. Signage was also printed, signage issued by QA should have been used. | No | Deviation
4185 | 8NYA- Support Service - RF618E013 Mod enzyme transfer record points you to tote fill instead of 20L and 50L bag fill. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-10-04 | 2024-10-04 | hold gemba to see path forward
4184 | TSH - Buffer Prep - Audible alarm tied to Sash on BSC F-1890 not working and replacement parts may not available | Jeremy.Wooden@Sanofi.com | 45 NYA Room 190 | 2024-10-04 | 2024-10-04 | TBD at the Gemba. It is expected that the audible alarm would sound when Sash was raised or lowered above/below the demarcation line but unsure which SOP contains this requirement. | Yes | Deviation
4183 | GENCI number on new Thyrogen media hoses does not match PM/CoA | nicholas.shoenfelt@sanofi.com | 51-129 | 2024-08-06 | 2024-10-03 | It is expected that the hoses GENCI # matches the CoA of hoses. | Yes | Deviation
4182 | Buffer GZ1410-51(0.1 M Sodium Hydroxide) lot EGFB1300 leak observes in a barrel | Sylvester.Kegbeh@sanofi.com | 51 NYA ROOMS 132 | 2024-10-04 | 2024-10-04 | it is expected that the buffer remains integral prior to and during processing | Yes | Deviation
4181 | cold room 108/109 entry log revised, inocorrect version used 02OCT24-04OCT24 | nicole.lufkin@sanofi.com | 51 NYA 108/109 | 2024-10-02 | 2024-10-04 | correct version of logbook would be sued | Yes | Deviation
4180 | Waterbath pre and post use sanitization not documented for thaw | nicole.lufkin@sanofi.com | 51NYA 152 small scale | 2024-09-26 | 2024-10-04 | water bath would be cleaned pre and post use | Yes | Deviation
4179 | 8 NYA Gemba - Support Services - Black particulates found in SUM bag during inspection | Nicholas.grote@sanofi.com | 08-1350 | 2024-10-03 | 2024-10-03 | Bag passes bag inspection and is able to be used for production. | Yes | Deviation
4178 | 8NYA-No Product-P/N 692168 250L SUM Bag Inspection Failure | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2024-10-03 | 2024-10-03 | It is expected for vendor supplied material pass internal inspection criteria | No | Deviation
4177 | 8 NYA Gemba - Support Services - OOS Osmo Result 760744 Cz Seed Media - RF623E010 | nicholas.grote@sanofi.com | 08-1350 | 2024-10-03 | 2024-10-03 | Within range sample result. | Yes | Deviation
4176 | No training performed by QC on V 5 of MABC-QU-SOP-0073084 due to iLearn/publishing error | Sarah.Adams@sanofi.com | QC MABC | 2023-12-27 | 2024-10-03 | Trained on effective version in Veeva
4175 | Superseded and approved attachments of QU-SOP-0074634 executed instead of Effective version | amanda.lavalley@sanofi.com | 76 NYA Biosurgery | 2024-09-25 | 2024-10-02 | It is expected that the current effective revision v17.0 of QU-SOP-0074634 is executed for manufacturing activities. | Yes | Deviation
4174 | 8NYA-Fz-EGHCE004 Antifoam line has suspect marking TBD | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1333 | 2024-10-02 | 2024-10-02 | It is expected that SUT be free of foreign materials
4173 | GEMBA - 8 NYA DOWNSTREAM - Slurry Volume Percentage Over Target | adam.keister@sanofi.com | 8 NYA - 1303 | 2024-10-02 | 2024-10-02 | For slurry volume percentage to be within target range. | Yes | Deviation
4172 | EM Plates for FB EGF1B02 expired prior to incubation | Alyson.Belleville@sanofi.com | 51-87 | 2024-10-02 | 2024-10-02 | It is expected that the plates would be delivered and incubated prior to expiry.
4171 | Expired Metaflex Solution Packs | david.loik@sanofi.com | 51 NYA | 2024-10-02 | 2024-10-02 | unexpired packs would be available for use. | Yes | No Deviation
4170 | Analyst unable to document CZ SDS stain start time in OneLIMs | Sarah.Adams@sanofi.com | 68 NYA, QCC | 2024-09-30 | 2024-10-01 | Analyst to document start time.
4169 | Sample 24MS01410, Fz POU, tested for wrong incubation scheme | Alyssa.Barrett@sanofi.com | 68 NYA/Room 4700/QC Micro | 2024-09-21 | 2024-10-01 | Per QU-SOP-0073109 Attachment 2, Fz POU sample part number 705688 is to be incubated for 3-5 days at 30-35 degrees Celsius. | Yes | Deviation
4168 | Late Extended PM of Filter Integrity Tester at 51 NYA | kevin.southwick@sanofi.com | 51 NYA | 2024-10-01 | 2024-10-01 | It is expected that work orders are completed by their due date per QU-SOP-0081784, 'Preventive Maintenance Program', Step 6.1.9. | Yes | Deviation
4167 | 51NYA APU-APT SS-Room 124-BSC F-16381 no as found data taken due to hood being disposed of | robert.bauch@sanofi.com | 51-124 | 2024-10-01 | 2024-10-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have had as found data taken before disposal | Yes | Deviation
4166 | Overdue PM's at 76NYA. | justin.carpenter@sanofi.com | 76 NYA WFI Loop | 2024-09-30 | 2024-09-30 | PM's should be completed on time | Yes | Deviation
4165 | Lack of insight on which record was to be used for CZ seed train | ryan.devanna@sanofi.com | 8 NYA, innoc prep | 2024-09-30 | 2024-09-30 | n/a
4164 | QU-OPE-0202445 lot EBYHAM009 steps 7.12.17 – 7.14.9 performed on 26SEP24 had no verifier/witness signatures. | dianne.moustafa@sanofi.com | 76 NYA Purification | 2024-09-26 | 2024-09-27 | it is expected to sign verifier steps and witness steps at time of processing | Yes | Deviation
4163 | Operators discovered a leaking tote during buffer setup | pavel.dobrovolskiy@sanofi.com | 8NYA CNC space | 2024-09-30 | 2024-09-30 | Its expected that the correct size / type of clamp will be attached to the unused outlets of the buffer bag | No | Deviation
4162 | Personnel improperly gowned in room 176 at 45NYA | michelle.smolinsky@sanofi.com | 45NYA, Room 176 | 2024-09-30 | 2024-09-30 | Personnel would be properly gowned in all areas
4161 | 8 NYA Support Services | Jose.Lugo2@sanofi.com | 8 NYA CNC | 2024-09-30 | 2024-09-30 | Delta V prompt to be fixed. | Yes | Deviation
4160 | 8 NYA GEMBA - Support Services - Damaged SUT upon inspection | william.lee@sanofi.com | 8NYA/1350 | 2024-10-10 | 2024-10-10 | The SUM bag should not be leaking | Yes | Deviation
4159 | 8NYA-Gemba-Cz Downstream- Phenyl column bioburden above 0 CFU on NRSR and routine rinse | aidan.mula@sanofi.com | 8NYA/ RM1303 | 2024-09-30 | 2024-09-30 | Bioburden should read 0 CFU | Yes | Deviation
4158 | 8 NYA GEMBA - Support Services - Small Leak on SUM Bag of Batch EGHAV134 | christopher.brewster@sanofi.com | 8 NYA, Rm. 1350 | 2024-09-29 | 2024-09-29 | That our SUM bags would not have rips/tears and the bags should not leak. | Yes | Deviation
4157 | 8NYA Follow up gemba for automation intervention in column packing | zachary.rickard@sanofi.com | 8NYA/1303 | 2024-09-26 | 2024-09-26 | Operations would progress as intended and equipment damage will not occur
4156 | 8 NYA Gemba - Support Services - 760731 Buffer - WFI Fill Out Of Range | Nicholas.Grote@sanofi.com | 08-1350 | 2024-09-28 | 2024-09-28 | WFI Fill is completed per recipe, and it is within the acceptable range. | Yes | No Deviation
4155 | 8 NYA Gemba - Support Services - 760731 is over on weight tolerance following QS | christopher.brewster@sanofi.com | 8 NYA, Rm.1350 | 2024-09-28 | 2024-09-28 | The automation recipe should have added the correct amount of WFI during QS, as to not go over on its' weight tolerance.
4154 | 8 NYA Gemba - Support Services - Out of Spec pH RF542E005 - 705686 Buffer | nicholas.grote@sanofi.com | 1350 and 1503 | 2024-09-27 | 2024-09-27 | Passing pH | Yes | Deviation
4153 | 8 NYA Gemba - Support Services - Ph Meter Standardization | nicholas.grote@sanofi.com | 08-1350 | 2024-09-26 | 2024-09-27 | TBD
4152 | Upstream, CZ, TCU Alarms SUM 2916 and SUB 2250 | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-09-27 | 2024-09-27 | Alarms to occur when needed | Yes | Deviation
4151 | 8 NYA - Support Serices - FIT Status Change - RF621E126 | nicholas.grote@sanofi.com | 08-1350 | 2024-09-23 | 2024-09-26 | FIT is allocated to the record in a way that shows up on the equipment log in MES. | Yes | Deviation
4150 | 8NYA-Cerezyme-EGHAV126 LAL sample 1390021 MEDIA_POU760741_ENDOTOXIN sample tube cracked and leaking | larissa.fairweather@sanofi.com | 68 NYA / 4700 / QCM | 2024-09-25 | 2024-09-26 | Per QU-SOP-0070558 step 6.2.3, confirm sample containers are not cracked, leaking. The expectation is that samples are to be free of any cracks or leaks. | Yes | Deviation
4149 | Credo cooler max volume exceeded during sample transport | zachary.rickard@sanofi.com | QC sample receipt | 2024-09-27 | 2024-09-27 | Credo cooler volume will not be exceeded. | No | Deviation
4148 | Upstream, CZ, RF639E005, SUB Aber probe was inadvertently zeroed during clean pulse | jeffrey.boldizar@sanofi.com | 8RM 1403 | 2024-09-27 | 2024-09-27 | Clean Plus performed not Zeroed | Yes | Deviation
4147 | TSH - buffer prep - BSC F-1890 and cold room 174 were removed from the RQ schedule as part of GCB CCR 2020FRACC0006 | Jeremy.Wooden@Sanofi.com | 45 NYA rooms 174 and 190 | 2020-02-14 | 2024-09-26 | It is expected that only the GCB equipment would have been included in 2020FRACC0006 and that the Buffer equipment would have had RQ's at the expected intervals.
4146 | 76NYA - EBYMFL002 - HA fill - Multiple errors | Craig.stevenson@sanofi.com | 76NYA, Room 355 | 2024-09-10 | 2024-09-26 | It is expected that zero filters generate 0 results on the 5th count. It is also expected that if action NV results are genrated, the SOP is followed to ensure there is no increase in values, MFG and Quality be made aware.
4145 | 8NYA - Support Service - Debris found on the inside of a tote bag prior to use. | bailey.gleason@sanofi.com | 8NYA CNC | 2024-09-26 | 2024-09-26 | Hold Gemba to see path forward.
4144 | TSH-45 NYA-Buffer Prep - Approved version 20 of QU-SOP-0075222 was used in place of Effective version 19. | Jeremy.Wooden@Sanofi.com | 45 NYA | 2024-09-23 | 2024-09-25 | It is expected that the effective version of the attachment would be used to document the walk through.
4143 | Cerezyme - EGHBN006 – ATF B connected to the incorrect ports | amanda.sorel@sanofi.com | 8 NYA | 2024-09-25 | 2024-09-25 | It is expected that the ATF B be hooked up per QU-SOP-0080430 | Yes | Deviation
4142 | 8NYA Aseptic challenge samples delivered expired and incubated | cleiton.lopes@sanofi.com | 68NYA | 2024-09-04 | 2024-09-24 | The samples in concern are expected to be delivered and incubated within 24 hours from collection time for EM and 48 hours for aseptic bottle sample | Yes | Deviation
4141 | Antifoam pump turning on during growth day 2/Antifoam Pump Turning on During Growth Day 2 When Local Switch is Turned on | MARCELO.DOSSANTOS2@SANOFI.COM | 8NYA /1333 | 2024-09-25 | 2024-09-25 | Antifoam pump should first be turning on when instructed by Delta V for priming on day 4. |  | Deviation
4140 | ATF connected to wrong tube on X2250 | ralston.najarro@sanofi.com | 8NYA1403 | 2024-09-25 | 2024-09-25 | Line should be connected to middle column | Yes | Deviation
4139 | 8NYA - Incorrect gowning in MAL 1411 | Craig.stevenson@sanofi.com | 8 NYA Room 1411 | 2024-09-25 | 2024-09-25 | It is expected that correct gowning be used while in classified areas. | Yes | No Deviation
4138 | Incorrect format for pressure recording on witness step | karthik.sriram@sanofi.com | 76 NYA | 2024-09-04 | 2024-09-24 | Per witness step 7.12.10.6 in QU-OPE-0309296, it is expected that the adjusted filtration pressure is ≤ 10.0.
4137 | Temperature spike during EDC reaction CMC 76 NYA | dianne.moustafa@sanofi.com | 76 NYA | 2024-09-25 | 2024-09-25 | Temperature stay in range of 23-27C
4136 | 45NYA - OOS Dewpoint , Compressed Air (45-CA-1) | Craig.stevenson@sanofi.com | 45NYA Room 228 | 2024-09-24 | 2024-09-24 | It is expected that Compressed Air generates a Dewpoint result of >-40C. | Yes | Deviation
4135 | 8NYA-Cz-X2230 Antifoam Stir Plate found disconnected from power supply | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2024-09-24 | 2024-09-24 | It is expected that the Antifoam be agitated at 70rpm using the electric spinner plate | Yes | No Deviation
4134 | 8NYA-No Product- SUM Bag Inspection Failure (P/N 692168) | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2024-09-24 | 2024-09-24 | It is expected to have single use containers pass inspection criteria set forth in EWI | Yes | No Deviation
4133 | 8NYA APU-80NYA room 122-EMS FB-5105-005 for -80 C Freezer reading at bottom of EMS range of -100 C | robert.bauch@sanofi.com | 80-122 | 2024-09-24 | 2024-09-24 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
4132 | 51NYA Gemba: WCB Fabrazyme at 45 not in locations listed in logbook | adam.rowe@sanofi.com | 45NYA, rm101 | 2024-09-24 | 2024-09-24 | Logbook vial locations should match physical locations
4131 | 8NYA-Cz-X2811 MCC hold for Buffer Line Pressure High High | JOHN.LYTLE@SANOFI.COM | Rm 1403, 8NYA | 2024-09-24 | 2024-09-24 | It is expected that buffer line is kept within allowable pressure ranges while under recipe
4130 | Fz Media Leak SUB 1 Support Services | robert.morawski@sanofi.com | 8NYA CNC RM 1210 | 2024-09-23 | 2024-09-23 | Amber clamps installed properly to not allow liquid to leak by. | Yes | Deviation
4129 | 8NYA Upstream - 705691 Media Tote with Leaking Aseptic Connection | wilfredo.lopez@sanofi.com | 8RM1202 | 2024-09-23 | 2024-09-23 | Expected that no liquid is present downstream of pinch clamps once the container is filled
4128 | 51 NYA - Thyrogen APU cold room 87 BSC FB-2074 found Out of Tolerance during scheduled calibration | glen.ciano@sanofi.com | 51-87 | 2024-09-19 | 2024-09-24 | per QU-SOP-0051601 unit should have produced passing results | Yes | Deviation
4127 | Upstream, FZ, RF503E006 media ROC low low alarm | jeffrey.boldizar@sanofi.com | 8RM 1333 | 2024-09-23 | 2024-09-23 | Rate of change to be around 100
4126 | 76NYA - HA - EBYHAM008 - Missing Non-Viable Data | Craig.stevenson@sanofi.com | 76NYA, Room 355 | 2024-09-03 | 2024-09-23 | At the end of the Fill or extra filling operations, turn off the continuous non-viable particulate counting cycle of all non-viable counters. Ensure that the full volume of air has been collected prior to shutting off the Climet otherwise collected data may be lost. | Yes | Deviation
4125 | 8NYA-Cerezyme-EGHBN005 LAL sample 1319543-CZ2_500_761217_ATF_ENDO submitted to QCM in polypropylene tube. | larissa.fairweather@sanofi.com | 68NYA/4700/QCM | 2024-09-19 | 2024-09-23 | Per QU-SOP-0014245 step 6.1.2, store samples at 2-8 °C in polystyrene containers. The expectation is that samples are not submitted to QCM in polypropylene tubes. | Yes | Deviation
4124 | 24 Hour Expiration After Media Charge on X2200(Upstream Operation)/ Prompt for 100% DO probe in media not available due to 24 hour media charge pH expiration. | MARCELO.DOSSANTOS2@SANOFI.COM | 8NYA /1333 | 2024-09-21 | 2024-09-22 | DO Probe stabilization should have occured 16 hours form the end of media load. And pH control should have occured 24 hours from the end of media load.
4123 | 8NYA Gemba-Downstream- Automation intervention required for Fz Zinc Blank Run | aidan.mula@sanofi.com | 8NYA/RM1323 | 2024-09-21 | 2024-09-21 | Recipe should run as intended without automation intervention.
4122 | 8NYA- Support service - Filter bleed valve cracked prior to transfer | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-09-21 | 2024-09-21 | swap out filter transfer line and tote bag | Yes | Deviation
4121 | TCU failed on SUB 2200 | brendan.dickhaut@sanofi.com | 8NYA/Room 1333 | 2024-09-20 | 2024-09-20 | TCU operates as expected and maintains set point temperatures of bioreactor. | Yes | Deviation
4120 | Upstream GEMBA Peristaltic Media Pump Rotating in the Opposite Direction on SUB skid X-2200 | Gary.dodakian3@sanofi.com | 8nya / 1333 | 2024-09-20 | 2024-09-20 | The media pump on X-2200 should be rotating clockwise
4119 | High top boots not downed over street shoes in CNC | jesse.welch@sanofi.com | 8NYA | 2024-09-20 | 2024-09-20 | high top boots worn | Yes | Deviation
4118 | Missing signature on probe set up record/ pH Autoclave Record Missing Signature | MARCELO.DOSSANTOS2@SANOFI.COM | 133 | 2024-09-19 | 2024-09-20 | Step for preparing the probe should have been signed
4117 | Expiry Issue of P/N#705691 Lot#EGHMD018 SU#204725 | yunbo.lou@sanofi.com | 8 NYA CNC | 2024-09-19 | 2024-09-19 | per FBL-EWI-000019, chill temperature should be selected when manufacturing 705691. Tote transportation to cold room should be recorded in MES.
4116 | During change over to LIMS in athena, master data recipe was updated per MABC-QU-FOR-0113845. During the update a master data change was made to the TSH Fractions (591161, 591166, and 591167). Due to this change QA is not able to cancel associated inspection lots per MABC-QU-SOP-0059507-Framingham Biologics Paper Batch Production Review Process for Manufacturing and (v30.0) during release of fractions. | Jason.Brodrick@Sanofi.com | 51 NYA | 2024-09-14 | 2024-09-19 | Expected results are that each fraction has an individual inspection lot with Athena SAP that can be released or canceled independently.
4115 | Misalignment of annual requalification of Thyrogen PRS for Bioassay with filed Thyrogen Dossier | Michael.Estepp@sanofi.com | 68 NYA, QC Microbiology | 2024-09-18 | 2024-09-18 | Alignment of Thyrogen dossier and QU-SOP-0055177 for annual PRS requalification for Bioassay.
4114 | 68 NYA HA Cracked EM Plates | karen.ramos2@sanofi.com | 68 QCM | 2024-09-18 | 2024-09-18 | Plates to be intact and incubated 20-25C | No | Deviation
4113 | Leaking 1K SUM Bag - 820323 2% MANNITOL,2%GLY,0.01%PS80,10MM L-HIS | Nicholas.Grote@sanofi.com | 08-1350 | 2024-09-19 | 2024-09-19 | SUM bag is integral, no leaks occur. | No | No Deviation
4112 | Media tote 2042001 not reading correct volume | daniel.sciotti@sanofi.com | 8NYA CNC | 2024-09-18 | 2024-09-18 | it is expected that the load cells on the media tote will read accurately
4111 | 8NYA Upstream - 706930 Fabrazyme Butyl Eluate SUM Outlet Line in Condensate | wilfredo.lopez@sanofi.com | 8RM1333 | 2024-09-18 | 2024-09-18 | Expected that no condensation buildup occur on the equipment. The line should be located in a way that any condensate does not impact the associated SUT.
4110 | 51 NYA APU-Dept APU SS-Room 129-Conductivity Indicator F-17195-025 found out of tolerance during an unscheduled calibration. | robert.bauch@sanofi.com | 51-129 | 2024-09-18 | 2024-09-19 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
4109 | 705675 Buffer Conductivity sample result out of Specification. | Bradd.Charles@sanofi.com | 8 NY AVE Room 1350 | 2024-09-18 | 2024-09-18 | Conductivity Measures within specification. |  | No Deviation
4108 | ATF 10 Filter Wetted Out | brian.o'rourke@sanofi.com | 8NYA Room 1403 | 2024-09-18 | 2024-09-18 | Stop the pump prior to liquid entering the Permeate line. | Yes | Deviation
4107 | L-Glutamine P/N 726192 was mislabeled with HU and Batch and consumed to process in solution prep at 8NYA. | michael.montiverdi@sanofi.com | Solution prep 8NYA | 2024-09-17 | 2024-09-17 | HU Label applied with corresponding material information as the affixed NBO label. | Yes | Deviation
4106 | Wash cycle conductivity on 51NYA parts washer found to be reading a constant 120mS. | nicholas.shoenfelt@sanofi.com | 51-129 | 2024-09-09 | 2024-09-17 | It is expected that the value of the conductivity be > 73mS/cm during CIP100 washes, and > 4mS/cm during CIP200 washes, and that the value not be stuck at one value. | Yes | Deviation
4105 | Insufficient 3X Contact Time 80 NYA Room 101 and 102 | jeff.hill@sanofi.com | 80NYA RM 101-102 | 2024-09-13 | 2024-09-18 | Contact time should be greater or equal to 5 minutes QU-SOP-0064997 step 6.6.1 and step 8. Glossary | Yes | Deviation
4104 | Temperature excursion per QU-OPE-0520874 during reaction of HA-CMC lot EBYCMC151 | nicholas.macdonald@sanofi.com | 76-229 | 2024-09-18 | 2024-09-18 | It is expected that the temperature remain within the range of 25C +/- 2C per QU-OPE-0520874.
4103 | EWI instructions missing from DO probe DLI reset | jeffrey.boldizar@sanofi.com | 8NYA 1333 | 2024-09-18 | 2024-09-18 | Instructions to reset DLI counter
4102 | 8NYA-Gemba-Downstream-No verifier for manual entry on load take out. | andrew.teixeira@sanofi.com | 8NYA/1703 | 2024-09-07 | 2024-09-13 | Manual entry should be verified by a second operator, also storage unit should auto populate as intended. | No | No Deviation
4101 | Leaking diaphragm on Clean Steam Generator #2 | justin.carpenter@sanofi.com | 45-181 | 2024-09-18 | 2024-09-18 | There should not be a leaks on the critical utility system | Yes | No Deviation
4100 | No Products – No Lots – 68NYA – Plate not submitted for ID for biolims sample 1270879. | megan.botti-marino@sanofi.com | 68 NYA, QC Microbiology Lab | 2024-07-11 | 2024-09-17 | It is expected that all EM/CU mold recoveries be submitted for identification per step 6.5.1 of MABC-QU-SOP-0019490-Counting of Microbial Plates in Quality Control (v4.0).
4099 | 8NYA- Support Services- Bag rejection due to debris found. | daniela.medina2@sanofi.com | 8NYA/ 1350 | 2024-09-17 | 2024-09-17 | Per FBL-EWI-00015 the bag should not have any abnormalities, particles or debris.
4098 | Cerezyme DS P/N 781143 (Lot: EGHDS014) Appearance Bag Cannot be Opened for Analysis | abel.agnes@sanofi.com | 68NYA 2700 | 2024-09-17 | 2024-09-17 | The stopper should be removable.
4097 | 250L SUM bags failed visual inspection | ralston.najarro@sanofi.com | 8NYA1403 | 2024-09-17 | 2024-09-17 | SUM bag SUT should be free of defects | Yes | Deviation
4096 | TSK EGF2B04A product loading capacity (0.2-0.5 g/L-resin) out of range (0.6g/L) discovered during prep at 51NYA. | Sylvester.Kegbeh@sanofi.com | 51 NYA ROOMS 132 | 2024-09-17 | 2024-09-17 | it is expected that the TSK EGF2B04A product loading capacity is within range of 0.2-0.5 g/L-resin
4095 | 76NYA-Biosurgery APU-Dept HABP-Room 3105-Viscometer F-21213 no as found data | robert.bauch@sanofi.com | 76-3105 | 2024-08-23 | 2024-09-17 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
4094 | 8NYA- Support service - Hard goods load was performed in autoclave while the equipment was OOS | bailey.gleason@sanofi.com | 8NYA RM1373 | 2024-09-16 | 2024-09-16 | hold gemba to see path forward
4093 | 8NYA APU-Dept PF8-room 1503-Pressure indicator FB-4778-004 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1503 | 2024-09-11 | 2024-09-17 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
4092 | Pompe ETP- Sample Volume Issue | nathaniel.barnes@sanofi.com | 8 NYA 1703 | 2024-09-16 | 2024-09-16 | Path forward to samples and QC testing
4091 | 8 NYA 8-WFI2-3 Daily WFI Site sample 1382377 Leaking from Cap | Rebecca.Schmidt@sanofi.com | 68NYA/4700/QCM | 2024-09-16 | 2024-09-16 | 8-WFI2-3 is a site that is collected daily for LAL, BIOB, COND, and TOC per QU-FOR-0118296. Sample should not be leaking after collection in order to be put on test in a timely manner.
4090 | Sweco screen noticed to have tear during disassembly room 3103 at 76 NYA EBYHAS002 | dianne.moustafa@sanofi.com | 76 NYA room 3103 | 2024-09-16 | 2024-09-16 | expected to have no tears Per QU-SOP-0008047 Section 6.7.4 | Yes | Deviation
4089 | 8NYA - Support Service - Leaked bag during initial WFI fill RF621E112 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-09-16 | 2024-09-16 | Hold a gemba, change out the SUM set up SUT | Yes | Deviation
4088 | 8NYA - Support Service - PG was consumed in entirely prior to transacting new PG in SAP. Causing the incorrect batch of PG attached to record RF606E010 | bailey.gleason@sanofi.com | 8NYA | 2024-09-16 | 2024-09-16 | Hold Gemba to see path forward. | Yes | Deviation
4087 | 8 NYA - Support Services - Small leak observed on SUM 1500 during agitator check | christopher.brewster@sanofi.com | 8 NYA, Rm. 1350 | 2024-09-15 | 2024-09-15 | That the SUM bag would not leak.
4086 | Cerezyme - EGHBN004 - 8NYA - Eluate liquid between clamp and aseptiquik on outlet line | michael.knapp@sanofi.com | 8NYA 1403 | 2024-09-13 | 2024-09-14 | No liquid beyond pinch clamp; aseptiquik connection is dry | No
4085 | Cerezyme - EGHBN004 - 8NYA - Eluate liquid between clamp and aseptiquik on outlet line | michael.knapp@sanofi.com | 8NYA 1403 | 2024-09-13 | 2024-09-14 | No liquid beyond pinch clamp; aseptiquik connection is dry
4084 | Cz MCC Buffer changeout MES Issue Upstream | robert.morawski@sanofi.com | 8NYA CNC RM 1210 | 2024-09-14 | 2024-09-14 | EBR advanced properly to the next step to complete execution.
4083 | Total mix time of eluate during TSK EGF2B03A out of range. | Alyson.Belleville@sanofi.com | 51-132 | 2024-09-11 | 2024-09-13 | It is expected that the start and end times recorded would add up to a total time within the range of 10-14 minutes.
4082 | 24L Wave should be gassing at 5%, but it is gassing at 10%. | jonathan.lucas@sanofi.com | 51 NYA, rm 152A | 2024-09-13 | 2024-09-14 | 24L should be gassing at 5%. MABC-QU-OPE-0254812-Provisional Process Control Strategy for Thyrogen Wave and BSB Processes (v1.0), p.24; MABC-QU-OPE-0515888-rhTSH Suspension Culture using Wave Bags to Seed the 160L Bioreactor at 51 NYA ( (v11.0), step 4.22.12 and step 4.29.
4081 | 8NYA Upstream - EGHCE004 - Fabrazyme SUB 2210 Aber Probe 2 Disabled | wilfredo.lopez@sanofi.com | 8RM1333 | 2024-09-13 | 2024-09-13 | It is expected that the probes retain consistent capacitance readings with minor delta readings between both. Capacitance readings are for information only at this point in the SUB recipe as it is in the growth period and bleed control has not begun. | Yes | No Deviation
4080 | 8 NYA - Support Services - Welder not qualified to perform 5 ft Extension to Manifold Weld | nicholas.grote@sanofi.com | 08-1202 | 2024-09-12 | 2024-09-12 | When the 832327 buffer is used in the upstream MCC process the buffer is shared. In order to share the buffer a 5 ft extension and manifold are welded onto the transfer line. The 5 ft extension line is c flex and the manifold is advantflex. Although the sartorius welder is capable of performing this weld the capability to do it was not captured in the qualification process.
4079 | N2 cool down PSI out of specification EBYHAS002 QU-OPE-0275728 | dianne.moustafa@sanofi.com | 76NYA purification | 2024-09-13 | 2024-09-13 | 5-20psi is maintained during the nitrogen cool down and continue the nitrogen purge overnight per QU-SOP-0048453 step 6.9.15 | Yes | Deviation
4078 | Remaining precipitate out of range. EBYHAS002 76 NYA Purification | dianne.moustafa@sanofi.com | 76NYA purification | 2024-09-12 | 2024-09-12 | specification 280-320KG QU-OPE-0202445 step 7.12.30
4077 | GEMBA for 8 NYA Dynamic Monitoring Plates Discarded in Error | jessica.kychler@sanofi.com | QC Micro 68 New York Avenue, floor 4, room 4700 | 2024-08-24 | 2024-09-12 | It is expected that plates are incubated for the appropriate amount of time and counted to provide data
4076 | MCC skid X-2831 in hold for BFLOWFAILURE on buffer pump P2831002 | dana.hopkins@sanofi.com | 8NYA 1403 | 2024-09-12 | 2024-09-12 | DeltaV recipe should not hold | Yes | No Deviation
4075 | 8NYA - Support service - Employee walked out of PAL RM1360 and into solution prep without proper gowning | bailey.gleason@sanofi.com | 8NYA 1360, CNC, and 1350 | 2024-09-12 | 2024-09-12 | Hold Gemba to see path forward.
4074 | 8NYA- Support service - RF1QJE014 832327 - manifold got stuck in welder. Line was cut. | bailey.gleason@sanofi.com | 8NYA | 2024-09-12 | 2024-09-12 | hold gemba to see path forward | Yes | No Deviation
4073 | I/O Error on Level Indicators on All TSH Bioreactors | james.letourneux@sanofi.com | 51 NYA Room 140 | 2024-09-12 | 2024-09-12 | Level indicators are critical components to the bioreactors, and manufacturing cannot proceed without Level Indicator functioning. It is expected that Level Indicators function and display accurate values through DeltaV. It is expected that critical process components not be decommissioned. | Yes | Deviation
4072 | low low alarm for media flow on bioreactors X2240 and X 2250 | brendan.dickhaut@sanofi.com | 8NYA/Room 1403 | 2024-09-11 | 2024-09-11 | installed media totes are free from obstructions such as clamps. | Yes | Deviation
4071 | ATF10 Failed visual inspection | ralston.najarro@sanofi.com | 8NYA1403 | 2024-09-11 | 2024-09-11 | ATF10 should be free of defects | Yes | Deviation
4070 | 51 NYA Purification-alpha numeric A280 result reported in OneLIMS for Q fraction 8 lot EGF1B02. | jaccob.kallon@sanofi.com | 45 NYA, QC Chemistry lab | 2024-09-10 | 2024-09-11 | a28 results are only reported in numeric values | No | Deviation
4069 | RO #2 storage tank low alarm causing breach when loop was used by cleaners | paul.mclaughlinjr@sanofi.com | 51NYA | 2024-09-10 | 2024-09-10 | Loop to remain integral
4068 | Logbook 2023LBF0127 missing last page (page 78) and back cover | Amy.McIntyre@sanofi.com | 51 New York Avenue | 2024-09-10 | 2024-09-10 | Expected results are that the issued logbook is issued with all its pages and returned with all pages accounted for per QU-SOP-0026193
4067 | 8NYA - Support Service - Freezer 1421 went into alarm on 10SEP24 | bailey.gleason@sanofi.com | 8NYA RM1421 | 2024-09-10 | 2024-09-11 | hold gemba to see path forward
4066 | 8NYA SS / upstream/ trains 3 +4 media install support | dennis.bazinet@sanofi.com | 8 NYA CNC space | 2024-09-11 | 2024-09-11 | Need support from MES to advance record
4065 | 2 cracked EM plates delivered to QC Micro lab | valeria.moreno@sanofi.com | 68 NYA- QC Micro lab | 2024-09-10 | 2024-09-10 | Incubate samples until Gemba is presented
4064 | 51 NYA RO2 Tank low low level | justin.carpenter@sanofi.com | 51 NYA | 2024-09-10 | 2024-09-10 | RO2 should continue to feed tank, low low tank level should alarm out
4063 | 51 NYA Clean steam use before proper sampling occurred | justin.carpenter@sanofi.com | 51 NYA | 2024-09-04 | 2024-09-10 | Clean steam needs to be qualified before being released
4062 | 8NYA-No Product-P/N 692168 250L SUM Bag Failed Visual Inspection for Free Floating Debris | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-09-10 | 2024-09-10 | Bags meet defined criteria found in procedure for passing inspection | Yes | Deviation
4061 | HA Fermentation EBYFRM044 Chart left in past 24 hours | Kenneth.morgan@sanofi.com | 76 NYA | 2024-09-04 | 2024-09-06 | It is expected that the process takes less than 24 hours
4060 | 45NYA Thyrogen APU-Room 190-Dept SP45- Weight indicators F-7093-008 and F-7093-009 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 45-190 | 2024-09-10 | 2024-09-10 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | No | Deviation
4059 | Late PM for 51NYA CSG (Original Due Date - No Risk) not Completed 'On-Time' | tanni.arif@sanofi.com | 51NYA mech room | 2024-09-01 | 2024-09-10 | all GMP PMs must be completed (<FTC'ed> in EAM) or extended by 31Aug24 | Yes | Deviation
4058 | FZ - EGHCE004 - 8NYA - Antifoam line did not prime with automated prime step | michael.knapp@sanofi.com | 8NYA 1333 | 2024-09-08 | 2024-09-08 | Line is primed when prompt is answered. | Yes | Deviation
4057 | Filter test for Q pooling GZ5155-00 lot EGF1B02A had printout with incorrect units of mbar instead of psig | Alyson.Belleville@sanofi.com | 51-140 | 2024-09-09 | 2024-09-09 | It is expected the test and results would be run and reported using the proceduralized units of psig, not mbar.
4056 | A centipede was found in common corridor inside the Manufacturing suite | valeria.moreno@sanofi.com | 76-363 | 2024-09-09 | 2024-09-09 | Follow SOP for pest detection
4055 | Cz Media EBR Issue During Installation Upstream | robert.morawski@sanofi.com | 8NYA CNC RM 1210 | 2024-09-04 | 2024-09-09 | The EBR would have populate all the BF necessary to perform the Media Change Out. | Yes | No Deviation
4054 | FZ - RF505E010 - 8NYA - Rinse water HBOD failed TOC | michael.knapp@sanofi.com | 8NYA 1333 | 2024-09-08 | 2024-09-09 | Passing result | Yes | No Deviation
4053 | 8NYA APU-downstream-Fabrazyme-Room 1323-two UV EasyCals FB-3970 and FB-3991 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1323 | 2024-08-22 | 2024-09-09 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | No | Deviation
4052 | 51NYA: QC unable to receive sample, it was not entered into BIOLims correctly. | adam.rowe@sanofi.com | 51NYA, 129 | 2024-09-07 | 2024-09-07 | Sample needs to be entered when it is pulled in order to capture the time | No | Deviation
4051 | EMS Differential Pressure Room 80 NYA RM 101-102 Less Than 0.04 For JIT Buffer Dispensing | jeff.hill@sanofi.com | 80NYA RM 101-102 | 2024-09-09 | 2024-09-09 | DP must be above 0.04.
4050 | 8NYA- Support Services- 8NYA. 760741 leak during initial WFI fill | maria.palmar@sanofi.com | 8 New York Ave, room 1350. | 2024-09-08 | 2024-09-08 | The bag to now leak | Yes | Deviation
4049 | 8NYA support services/upstream: Profibus cable damaged on bht-2041002 (EGHAV099) | phoebe.karkos@sanofi.com | 8NYA CNC of P2 (1403) | 2024-09-08 | 2024-09-08 | Profibus cable repaired. Work order number 3255742 placed | Yes | No Deviation
4048 | 8NYA support services/upstream: Leaking Media Tote EGHAV100 | phoebe.karkos@sanofi.com | 8NYA CNC of P2 (1403) | 2024-09-08 | 2024-09-08 | media should not be leaking
4047 | 8NYA support services/Upstream Media install on Trains 3&4 | genesis.torres2@sanofi.com | 8nya CNC | 2024-09-08 | 2024-09-08 | Media install on trains 3&4
4046 | Incorrect PG amount dispensed | maria.palmar@sanofi.com | 8NYA / 1350. | 2024-09-08 | 2024-09-08 | PG needs to meet the qt required.
4045 | Train 1,2,3&4 Bleed tote removal and install not moving forward in EBR | genesis.torres2@sanofi.com | 8nya CNC | 2024-09-07 | 2024-09-07 | Fix bleed tote setup on all 4 records so that they function properly. | Yes | No Deviation
4044 | Shared media disposed in one record instead of two. | jose.lugo2@sanofi.com | 8 NYA CNC | 2024-09-07 | 2024-09-07 | Record back to normal function. | Yes | Deviation
4043 | 8NYA - CZ - Shipping bag transfer recipe incorrect volume bag 2 | amanda.sorel@sanofi.com | 8 NYA 1703 | 2024-09-07 | 2024-09-07 | It is expected that MES should generate the correct recipe for shipping bag transfer | Yes | Deviation
4042 | 8NYA GEMBA Downstream. Cerezyme, Incorrect instructions in SOP for Shipping bag transfer | alex.tello@sanofi.com | 1703 | 2024-09-07 | 2024-09-07 | We are not supposed to change SUT after filter.
4041 | SUB CNC EBR stuck. new problem not referenced in job aid. | genesis.torres2@sanofi.com | 8nya | 2024-09-06 | 2024-09-06 | Job aid instructions to work | Yes | Deviation
4040 | non uniformity found on surface of installed eluate collection bag, not yet in use | brendan.dickhaut@sanofi.com | 8NYA/Room 1403 | 2024-09-07 | 2024-09-07 | raw materials are free of defect | Yes | Deviation
4039 | Cz Nano- Leak on Nanofilter inlet during Load | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2024-09-07 | 2024-09-07 | Determine path forward for Nano processing | Yes | Deviation
4038 | Fabrazyme: Harvest pumps not turn on after Guard filter change out FBL-EWI-000011. | amanda.sorel@sanofi.com | 8 NYA | 2024-09-06 | 2024-09-06 | It is expected that once the guard filter is changed out that the pump would resume running | Yes | Deviation
4037 | 8NYA GEMBA Cerezyme Downstream, Osmo out of range after formulation | alex.tello@sanofi.com | 1503 | 2024-09-06 | 2024-09-06 | Osmo result to be in range | Yes | Deviation
4036 | Failed Visual Inspection of 250L SUM | ana.salichs@sanofi.com | 8NYA 1403 | 2024-09-01 | 2024-09-01 | Bags pass visual inspection per QU-WIN-0019199 | No | Deviation
4035 | X2230 DO SEL DEV alarm | ralston.najarro@sanofi.com | 8NYA1403 | 2024-09-06 | 2024-09-06 | Deviation alarm put recipe in hold | Yes | Deviation
4034 | X2220 modbad alarm on bleed pump | ralston.najarro@sanofi.com | 8NYA1403 | 2024-09-06 | 2024-09-06 | Modbad put recipe in hold | Yes | No Deviation
4033 | Blue Eluate EGF1B02A was removed from cold 147 at 51NYA without signing the BIRF | sylvester.kegbeh@sanofi.com | 51 NYA ROOMS 132 | 2024-07-22 | 2024-08-01 | it is expected that the Eluate be stored per QUSOP-0074092 | Yes | No Deviation
4032 | FZ - EGHCE004 - 8NYA - Leak on ATF B Guard Filter Bleed Valve | michael.knapp@sanofi.com | 8NYA 1333 | 2024-09-05 | 2024-09-05 | No leak on guard filter | Yes | Deviation
4031 | GEMBA - 8 NYA - DOWNSTREAM - CM SOP Has Wrong PH Spec Range for pH Adjust | adam.keister@sanofi.com | 8 NYA, room 1503 | 2024-09-05 | 2024-09-05 | For SOP to have correct pH target range
4030 | 8NYA-Cz-EGHBN006- X2250 Bleed EM failed to switch from Fixed Rate Bleed to Variable Bleed | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-09-03 | 2024-09-05 | It is expected for the Bleed EM to switch from "fixed rate' bleed to "variable" bleed on Day 10 of Cz-ICB. Automation intervention requested to force state to "Variable" bleed | Yes | Deviation
4029 | 45NYA- Buffer Lot: CGFB2037 PN: 1402- Sample 2200051178 disposed of before 6 month hold time | lauren.ruffing@sanofi.com | 45 NYA | 2024-09-05 | 2024-09-05 | Samples retains are held for the minimum hold time before being disposed per QU-SOP-0010783 | Yes | Deviation
4028 | 8NYA-Cerezyme CM - EGHBV013 - Recipe went into hold due to temp alarm. | stanley.durand@sanofi.com | 8NYA rm1503 | 2024-09-05 | 2024-09-05 | It's expected for recipe to be completed without holds due to alarms | Yes | Deviation
4027 | MMC Equil Endotoxin Sample Missing | zachary.rickard@sanofi.com | 8NYA1220/8NYA1503 | 2024-09-03 | 2024-09-04 | Samples will be correctly handled, meaning dropped off where they are documented to be and delivered to QC on time.
4026 | 8NYA- Cz-EGHBN007- X2560 Blank Run 1 holding for High Pressure | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-09-04 | 2024-09-04 | It is expected that buffer flow freely through the column during Blank Run operations | Yes | No Deviation
4025 | Several PM Ext Forms [MABC-QU-FOR-0093877] (>70 WOs) (Effective from 01May2024) Processed Without Date Stamp in Veeva (GDP Issue) | MA-BioCampus_F&E-Compliance@sanofi.com | sitewide | 2024-05-01 | 2024-09-04 | all documents downloaded from Veeva shall have a date/time stamp to ensure latest/original copy | Yes | Deviation
4024 | Cz CNC SUB EBR Issue during Media Changeout Support Services | robert.morawski@sanofi.com | 8NYA CNC RM 1210 | 2024-09-04 | 2024-09-04 | The EBR would have populate all the BF necessary to perform the Media Change Out. | Yes | No Deviation
4023 | Gemba/8NYA/Downstream/Shipping bag failed pre-use inspection | andrew.teixeira@sanofi.com | 8NYA/1703 | 2024-09-04 | 2024-09-04 | Bag to pass inspection per QU-OPE-0492708
4022 | 8NYA APU-CC8-Room 1403-UV EasyCal FB-4890 and FB-4891 found out of tolerance during a scheduled Calibration | robert.bauch@sanofi.com | 8-1403 | 2024-08-14 | 2024-09-05 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | No | Deviation
4021 | 8 NYA - Support Services - Small leak on tote bag during transfer of 760727 | christopher.brewster@sanofi.com | 8 NYA, Rm. 1202 | 2024-09-04 | 2024-09-04 | That the tote bag would not leak during transfer. | Yes | Deviation
4020 | Missing entries and review on QU-OPE-0202442 Carbohydrate Analysis of rhTSH assay | aimee.park@sanofi.com | 45NYA, QC Chemistry | 2024-08-18 | 2024-09-04 | Per QU-OPE-0202442, all data entries and review would have been performed at the time of execution.
4019 | 8NYA-Cz-EGHBN007 X2220 Aber probe 2 divergence past EWI criteria for disabling and trending down, but recipe logic re-enables it each Aber Pulse operation (every 8hrs) | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-09-04 | 2024-09-04 | It is expected to use the average of the 2 capacitance probes to calculate average capacitance with probes within allowable limits of each other set forth by EWI-000010 | Yes | Deviation
4018 | 51 NYA discovered MECCO WFI still not discharging to WFI tank after shutdown | justin.carpenter@sanofi.com | 51 NYA | 2024-09-03 | 2024-09-04 | MECCO WFI still should fill the tank | Yes | Deviation
4017 | No products - No Lots - 68 NYA - No reportable non-viable result for weekly EM site 134-1 in 51NYA | megan.botti-marino@sanofi.com | 51 NYA, Room 134 | 2024-04-10 | 2024-09-04 | It is expected that grade C site 341-1 be sampled weekly for non-viables, viable air, RODAC A and RODAC B
4016 | GEMBA - 8NYA - Downstream - No part number listed for filter bleed tubing | danny.martinez@sanofi.com | 8NYA/RM 1503 | 2024-09-04 | 2024-09-04 | Per QU-SOP-0078709 step 6.16.11, Add a part number for the tubing to bleed the filter.
4015 | Missing data on chart for Sweco pressure test and SIP EBYHAM008 76 NYA | dianne.moustafa@sanofi.com | 76 NYA room 3103 | 2024-08-29 | 2024-09-04 | Chart should display all data
4014 | FZ- EGHCE004- 8NYA - SUB 2210 unable to complete DO Stabilization | michael.knapp@sanofi.com | 8NYA 1333 | 2024-09-04 | 2024-09-04 | DO Probes stabilize | Yes | Deviation
4013 | No LFH EUR to track Hood Cleaning | ryan.devanna@sanofi.com | 1343 8 New York Ave | 2024-08-21 | 2024-09-03 | Creating MES EUR or Logbook to log and follow monthly and daily cleanings.
4012 | Record RF600E030 answered too early | dana.hopkins@sanofi.com | 8NYA 1403 | 2024-09-03 | 2024-09-03 | DeltaV recipe should end before reocrd is terminated
4011 | QC Flying Visit not performed for the Month of August | kyle.lebherz@sanofi.com | QC Micro. 4700 68 New York Ave | 2024-09-01 | 2024-09-03 | QC Flying Visit to be performed in August | Yes | No Deviation
4010 | Sealer was out of service, MES batch record turned green instead of red | dennis.bazinet@sanofi.com | 8 NYA CNC space | 2024-09-03 | 2024-09-03 | Record didn't stop us from advancing in the record, turned green instead of red
4009 | Poor vacuum on the Gemco Dryer Vacuum system caused delay in processing. | jan.stepanik@sanofi.com | 76-229 | 2024-08-31 | 2024-08-31 | Expected vacuum should have been within specification.
4008 | AUG Late PMs - 2 Total - 45NYA WFI Vent Filters Pre-Testing - WOs 3105258 and 3105259 | kevin.southwick@sanofi.com | 45NYA Mech Rm | 2024-09-01 | 2024-09-03 | No Late PMs | Yes | Deviation
4007 | Media pump malfunctioning/ Media pump on X-2210 malfunctioning | MARCELO.DOSSANTOS2@SANOFI.COM | 1333 | 2024-09-02 | 2024-09-02 | Media load time of transfer should not exceed two and a half hours | Yes | Deviation
4006 | Determine the best way to process RF639E004 CNC SUB record | dennis.bazinet@sanofi.com | 8 NYA / CNC space | 2024-09-01 | 2024-09-01 | Answer the prompt in record and the record will advance to the next prompt | Yes | Deviation
4005 | The Phenyl SOP and MES record do not match and leave gaps in instruction for load pooling. | patrick.mitchell@sanofi.com | 8NYA/1503 | 2024-09-02 | 2024-09-02 | The SOP should instruct on how to pool foru all the necessary material | Yes | No Deviation
4004 | 8 NYA 8-WFI3A-1 Daily Site sample 1356564 not collected - tube found dry upon pre-testing | Rebecca.Schmidt@sanofi.com | 68NYA/4700/QCM | 2024-09-01 | 2024-09-01 | 8-WFI3A-1 is a site that is collected daily for LAL, BIOB, COND, and TOC per QU-FOR-0118296. Sample should have been collected during the day in order to be tested later.
4003 | Cz CNC SUB EBR Issue | robert.morawski@sanofi.com | 8NYA CNC RM 1210 | 2024-09-01 | 2024-09-01 | The EBR would have populate all the BF necessary to perform the Media Change Out.
4002 | CZ, RF639E004, SUM bag failed inspection | jeffrey.boldizar@sanofi.com | 8NYA1403 | 2024-09-01 | 2024-09-01 | a clean and obvious integral seal
4001 | 8 NYA Gemba - Support Services - Damaged Air Filter 717323 | nicholas.grote@sanofi.com | 08-1350 | 2024-09-01 | 2024-09-01 | Filter is integral
4000 | 8NYA - 761219 - Cerezyme Zinc Incorrect Load Filtration Setup | wilfredo.lopez@sanofi.com | 8RM1503 | 2024-09-01 | 2024-09-01 | It is expected that the operations are completed as instructed per the SOP. Expected that all SUT is used during the operation as described. | Yes | Deviation
3999 | 8NYA - 813910 SUM Bag Integrity Compromised During Probe Standardization | Kelvin.Glean@sanofi.com | Room 1503 | 2024-09-01 | 2024-09-01 | Buffer to go into the attached waste bag.
3998 | "8 NYA Loose Cz Media Biob Sample Cap - Leaking Sample Batch: EGHAV090". | yanitza.montalvo@sanofi.com | 68NYA, QC Micro | 2024-08-31 | 2024-08-31 | Per QU-SOP-0329679 3. Responsibilities, QC Personnel are supposed to note if there are any issues with samples and notify management. Samples should not be leaking prior to testing.
3997 | 8 NYA Gemba - Support Services - Out of range liquid weigh out during Reagent Prep for 760741 Media | nicholas.grote@sanofi.com | 08-1350 | 2024-08-31 | 2024-08-31 | Liquid weights and reagent prep is within range. | Yes | Deviation
3996 | 8 NYA Gemba - Support Services - Leaking Seal 760741 Media Transfer Line | Nicholas.Grote@sanofi.com | 08-1202 | 2024-08-31 | 2024-08-31 | Good seal, no leaking from seal. | Yes | Deviation
3995 | A re-gemba is needed for: "8 NYA Gemba needed for issue with flow sensor". | kenny.leland@sanofi.com | 8 NYA 1503 | 2024-08-30 | 2024-08-30 | Recipe should see flow during recirculation. Restarted recipe and called Gemba
3994 | 8 NYA - Support Services - Tolerance discrepancy on the HCl during DTT formulation for 760741 | christopher.brewster@sanofi.com | Rm. 1350 at 8 NYA. | 2024-08-30 | 2024-08-30 | If in fact, they weighed below the target amount, the step would have turned red and would not have allowed the record to proceed. It seems there may be a discrepancy with the weighing tolerance of this material. Also, the air flow/sash in the LFH seems to heavily influence the reading on the scale. This is what most likely lead to the operators seeing. 0.154, but only 0.148 being captured by the batch record.
3993 | CZ, RF600E023, UV Low sensor not responding during CIP setup | jeffrey.boldizar@sanofi.com | 8NYA1403 | 2024-08-30 | 2024-08-30 | Low filter show a value for UV check
3992 | 8 NYA Gemba needed for issue with flow sensor | kenny.leland@sanofi.com | 8 NYA 1503 | 2024-08-30 | 2024-08-30 | Recipe should see flow during recirculation. Restarted recipe and called Gemba
3991 | Contemporaneous documentation of GMP training event for GPS3 Investigator 26-27Aug2024 | melissa.goldesberry@sanofi.com | 68NYA 1410 | 2024-08-27 | 2024-08-30 | To provide documentation of response to an FDA 483, Sanofi will provide documentation of training evidence.
3990 | 51NYA - Pest Discovered in room 140 | Craig.stevenson@sanofi.com | 51 NYA Room 140 | 2024-08-30 | 2024-08-30 | It is expected that the controled clean rooms be pest free
3989 | 8NYA-Cz 761217- EGHBN006- MCC X2811 Column 2 pH probe non responsive during Blank Run 2 recipe | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-08-30 | 2024-08-30 | functioning pH probe |  | No Deviation
3988 | 8NYA-Cz-EGHBN005- Bleed EM failed to switch from Fixed Rate Bleed to Variable Bleed | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-08-28 | 2024-08-28 | It is expected for the Bleed EM to switch from "fixed rate' bleed to "variable" bleed on Day 10 of Cz-ICB. Automation intervention requested to force state to "Variable" bleed
3987 | No Products - No Lots - 8 NYA - MABC-QU-SOP-0038066 "8 NYA Cell Culture Contamination Response" is not a part of any curricula | amanda.sorel@sanofi.com | 8 NYA | 2024-08-23 | 2024-08-23 | It is expected that if a SOP is required for a job it would be in a training curricula. | Yes | Deviation
3986 | 2250 Bleed Pressure I/O Failure | brian.o'rourke@sanofi.com | 8NYA Room1403 | 2024-08-30 | 2024-08-30 | Pressure indicator should give continuous accurate representation of bleed line pressure. | Yes | Deviation
3985 | Power outage report received - DP alarms 45 and 51 NYA | nicholas.shoenfelt@sanofi.com | 45/51NYA | 2024-08-29 | 2024-08-29 | It is expected that DP will be maintained in the manufacturing areas.
3984 | 8NYA-Cz-Multiple Lots- No Dispase Inventory to complete Daily Sampling activities for Cz Trains | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-08-29 | 2024-08-29 | Sufficient QTY of dispase to perform daily sampling activities for Cz trains | Yes | No Deviation
3983 | CAMPUS POWER INTERUPTION . WO# 3253140 | MACJONES.OSEIKUFFOUR@SANOFI.COM | 8NYA, ROOM 1170 MECHANICAL ROOM LOST ONE LOOP DISTRIBUTION PUMP TO LOOPS 1,2,3A,3B AND 4, BUILDING 45 PENTHOUSE VACUUM, MAHU 1 AND MAHU 2.AND 51 RTU 3 AND RTU 8 | 2024-07-29 | 2024-07-29 | MABC-QU-SOP-0001515
3982 | X2250 went into hold for low media flow | ralston.najarro@sanofi.com | 8NYA1403 | 2024-08-29 | 2024-08-29 | media should switch over and have no obstructions
3981 | Max Prime Time timeout on MCC skid X2811 | brendan.dickhaut@sanofi.com | 8NYA/Room 1403 | 2024-08-28 | 2024-08-28 | Recipe runs without max timers expiring and going into hold. | Yes | No Deviation
3980 | 8 NYA APU-Dept CC8-room 1333-Temperature Indicator FB-2771-034 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 8-1333 | 2024-08-28 | 2024-08-29 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3979 | X2220 ABER probe not reading on transmitter | ralston.najarro@sanofi.com | 8NYA1403 | 2024-08-28 | 2024-08-28 | FBL-EWI-000010 Zero Aber probe
3978 | 68 NYA - No Product - No Lot. TSA lot 24GN01313 plate had mold grown on it prior to use | judith.keller@sanofi.com | 68 NYA 4700 QCM | 2024-08-27 | 2024-08-27 | It is expected that all media plates are sterile and have no growth when in use. | Yes | Deviation
3977 | EMS Readouts Did Not Match Balancing | tanni.arif@sanofi.com | 51NYA | 2024-08-26 | 2024-08-26 | EMS readouts are expected to be accurate, ITSCR required to drive this change
3976 | 68 NYA - No Product - No Lot. TSA lot 24GN01313 plate had mold grown on it prior to use | joseph.napoli@sanofi.com | 68 NYA 4700 QCM | 2024-08-27 | 2024-08-27 | It is expected that all media plates are sterile and have no growth when in use.
3975 | No Athena Goods Receipt of 761217 EGHBN00306 | joshua.brooks@sanofi.com | 8NYA Upstream | 2024-08-11 | unknown | Stock created in Athena via MES goods receipt
3974 | 51NYA Wave Bioreactors need reboot coming out of facility shutdown | adriana.alban@sanofi.com | 51NYA/152 | 2024-08-23 | 2024-08-28 | It is expected that equipment reconnects to DeltaV when network restored
3973 | Re-Gemba for Sartorius BioWelder seized during operation to weld a manifold to glycol tote 832327 | joel.henchar@sanofi.com | CNC area at 8NYA. | 2023-08-01 | 2023-08-01 | Work order 3212393 was created for this occurrence. | Yes | Deviation
3972 | QE-068461 additions for events 23AUG2024, Wrong label was placed on sample | Phoebe.karkos@sanofi.com | 8NYA phase 1 (1333) | 2024-08-23 | 2024-08-24 | events to be added to QE-068461, , correct label was printed and the sample was relabeled. QC is now able to process sample. Per QC no change is needed in biolims as correct sample was processed in biolims, but Judith Keller will mark the sample with the QE number. Per QA, no comment is needed in record as nothing ties the incorrect label to the EBR for RF500E004 | Yes | Deviation
3971 | 8NYA-Gemba-Downstream DP Excursion in RM1503 | aidan.mula@sanofi.com | 8NYA/ RM1503 | 2024-08-27 | 2024-08-28 | Doors should not get stuck open | Yes | Deviation
3970 | Slight leak on Zinc Skid X-3311 | charles.salgadinho@sanofi.com | 8NYA, room 1503 | 2024-08-28 | 2024-08-28 | Skid needs to run regen/ store without leaks. QU-SOP-0080419
3969 | 8NYA-RF639E004- CZ - X2240 SUB Hold due to ATF A&B harvest line pressure event. System in hold | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-08-27 | 2024-08-27 | It is expected that the system continues perfusion uninterrupted
3968 | QE-068933 missing sample issue needs closure for Lot Shipment | robert.morawski@sanofi.com | 8NYA CNC RM 1210 | 2024-07-22 | 2024-07-22 | Resampling of EGHMD051 to be submitted under the correct Batch ID.
3967 | Automation intervention will be required to reset WFI drop Amber light | dom.arone@sanofi.com | 8 NYA, RM 1373 | 2023-08-28 | 2024-08-27 | QU-OPE-0521196, Attachment 1 and 2 decision tree
3966 | Bioburden hit of b. cereus for 51CV-0018, resulting in failed RQ for sensor array X-3350. | Alyson.Belleville@sanofi.com | 51 NYA, 138 | 2024-07-25 | 2024-07-31 | It is expected that the bioburden sample would come back with no growth and passing criteria.
3965 | Broken sample port on 200 L SUM bag | christopher.brewster@sanofi.com | Rm. 1350 at 8 NYA. | 2024-08-26 | 2024-08-26 | It is expected that the SUM bags arrive undamaged from the vendor.
3964 | Broken AseptiQuik Connection - Discovered prior to transfer on 760741 XVF Media with DTT | nicholas.grote@sanofi.com | 1220 | 2024-08-26 | 2024-08-26 | Connection is in good condition when it is opened and it is able to be used for the transfer.
3963 | CAPA-005242 was not implemented as expected and the record is closed | kristine.caron@sanofi.com | 8NYA | 2024-06-19 | 2024-08-26 | Instructions should have been added to the Eluate Pooling Setup (FZ) in EWI-33
3962 | GEMBA 8NYA - 714043; unnecessary SUT for liquid addition. | JACK.LAPLANTE@SANOFI.COM | 8NYA RM1350 | 2024-08-25 | 2024-08-25 | path forward for the future
3961 | Wood identified in warm room door frame exposed to MFG area room 51-148 | nicholas.shoenfelt@sanofi.com | 51-148 | 2024-08-26 | 2024-08-26 | Wood is not expected to be in the MFG area.
3960 | 45 NYA RO drop testing failed on conductivity | justin.carpenter@sanofi.com | 45/rm181 | 2024-08-24 | 2024-08-24 | Sample should pass Quality inspection
3959 | TCU F008-XA-2835005 (F3771-006) was found not calibrated but still in use | amanda.sorel@sanofi.com | 8 NYA Mezziane | 2024-08-23 | 2024-08-23 | It is expected that all equipment is use is calibrated | Yes | Deviation
3958 | C-3110030 CV Rinse for Fz RF533E004_003 Bioburden Recovery | zachary.rickard@sanofi.com | QC | 2024-08-09 | 2024-08-09 | Bioburden will be within the range defined.
3957 | 8NYA-Cz-EGHBN005- ATF A alarming for low flow rate after perfusion rate shift | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-08-24 | 2024-08-26 | It is expected that ATF A and ATF B meet the flow rate criteria specified under the recipe and CST | Yes | No Deviation
3956 | GEMBA 8NYA - 2x Black particulates embedded on 3k SUM bags | JACK.LAPLANTE@SANOFI.COM | 8NYA 1350 | 2024-08-25 | 2024-08-25 | Path forward
3955 | 8NYA-DOWNSTREAM-CZ MMC-RF641E011-A280 SAMPLE | Marc.Latta@sanofi.com | 1503 | 2024-08-25 | 2024-08-25 | QU-SOP-0079722
3954 | Missed biolims Day 0 samples and sampling loop on X2250 | ralston.najarro@sanofi.com | 8NYA1403 | 2024-08-24 | 2024-08-25 | Samples should've been taken and sent to QC
3953 | Multiple holds on X2240 for bleed pressure | ralston.najarro@sanofi.com | 8NYA1403 | 2024-08-25 | 2024-08-25 | Bleed should not go into hold
3952 | Unable to locate 3 samples for SHIPMENT of Pompe EGHDS004(sample ID 1237822, 1273550 and 1273551) | jose.antunes@sanofi.com | 8NYA | 2024-07-06 | 2024-08-24 | It is expected that all buffers and media would have associated samples per FBL-SOP-000696 attachment 12. | Yes | Deviation
3951 | Gemba for DeltaV at 51 for WFI-2 | nathaniel.milde@sanofi.com | 51 NYA room 140 | 2024-08-24 | 2024-08-24 | Manufacturing should be able to open the valve
3950 | 2240 is delayed in getting to variable bleed state | ralston.najarro@sanofi.com | 8NYA1403 | 2024-08-24 | 2024-08-24 | Bleed should be in variable bleed state
3949 | Media clamps not squeezed completely together media leaked by into lines | dennis.bazinet@sanofi.com | 8 NYA / CNC space | 2024-08-24 | 2024-08-24 | Needed to use the correct clamp
3948 | Column PI was found to be out of calibration | pavel.dobrovolskiy@sanofi.com | 8nya 1503 | 2024-08-23 | 2024-08-23 | Ite expected that all equipment remain within calibration during GMP operations
3947 | No documentation of training for Analyst performing Determination of Aggregates in NeoGAA | sarah.adams@sanofi.com | 68 NYA, QCC | 2023-11-03 | 2024-08-22 | Analyst complete a Read and Understood on method prior to executing.
3946 | Weight discrepancy during due to a failed PI pull | pavel.dobrovolskiy@sanofi.com | 8NYA 1503 | 2024-07-20 | 2024-08-22 | Its expected that once the operator manually enters the weight, they will re-calculate the following step to reflect new value. | No | Deviation
3945 | Cracked EM plate for FZ Flask Sampling | Phoebe.karkos@sanofi.com | 8NYA innoc prep | 2024-08-23 | 2024-08-23 | discuss with QC on best steps forward on compromised plate | Yes | Deviation
3944 | RO/WFI not available for NTO cleaning 45NYA TSH | nicholas.shoenfelt@sanofi.com | 45NYA all rooms | 2024-08-22 | 2024-08-22 | It is expected that RO and/or WFI will be available for daily NTO cleanings after the area is return to a normal state.
3943 | Gemba 8NYA 20L 705691 media leak into sample bag | william.lee@sanofi.com | 8NYA/1350 | 2024-08-22 | 2024-08-22 | Pinch clamp should have been closed. Media should not have flowed into sample bag. | Yes | Deviation
3942 | Assess risks of new amber clamps on the floor--not enough available for schedule | leo.gillespie@sanofi.com | 1350 SP/ CNC Tote filling hallway | 2024-08-22 | 2024-08-22 | Clamps should not fail | Yes | No Deviation
3941 | CZ, EGHBN004, first antifoam sample in BioLims not performed in correct LOT | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-08-16 | 2024-08-22 | samples to be performed in correct LOT in BioLims | Yes | Deviation
3940 | 8 NYA Loose CZ Media Bioburden Sample Cap - Leaking Sample Batch: EGHAV078 | Rebecca.Schmidt@sanofi.com | 68NYA/4700/QCM | 2024-08-20 | 2024-08-20 | Per QU-SOP-0329679 3. Responsibilities, QC Personnel are supposed to note if there are any issues with samples and notify management. Samples should not be leaking prior to testing.
3939 | 68 NYA - No Product - No Lot - Positive Turbidity Identified in Aseptic Qualifications | amy.roy@sanofi.com | 68 NYA 4700 QCM | 2024-07-13 | 2024-08-20 | It is expected that there is no positive turbidity in aseptic qualifications.
3938 | rhTSH Unformulated Bulk Bioburden Tested at Incorrect Volume | jessica.kychler@sanofi.com | 68 NYA QC Micro 4700 | 2024-08-06 | 2024-08-21 | It is expected that the specified volume be tested per the applicable SOP and that errors in testing parameters will be caught on verification prior to testing
3937 | After initial media charge, TCU for X-2250 would not turn on | dana.hopkins@sanofi.com | 8NYA 1403, interstitial space | 2024-08-20 | 2024-08-20 | TCU should turn on after media charge
3936 | The monthly maintenance of water bath FB-2324 was not performed | craig.stevenson@sanofi.com | 68NYA/4700/QC Micro | 2024-08-20 | 2024-08-19 | The monthly maintenance of the water bath is to be performed and documented per QU-SOP-0024808, step 6.3
3935 | Controlled Print - Form Execution used, and in duplicate | abel.agnes@sanofi.com | 68NYA QCM | 2024-07-03 | 2024-08-19 | MABC-QU-SOP-0034503 Step 6.5 instructs that controlled prints be pulled for production execution. Same step also states that each Controlled Print should only be printed once.
3934 | pH out of range step 7.2.9.2 EBYFRM042 at 76NYA room 364 Fermenation | dianne.moustafa@sanofi.com | 76 NYA Room 364 Fermentation | 2024-08-20 | 2024-08-20 | pH should be within range of 6.40 to 7.40 per QU-OPE-0079507 step 7.2.9.2 | Yes | Deviation
3933 | CMC restriction for MX Thyrogen | katarzyna.macko@sanofi.com | NBO | 2024-08-15 | 2024-08-15 | Change needs to be justified in order to release Batch EY0168
3932 | Record did not execute as expected with interim control. | brendan.dickhaut@sanofi.com | 8NYA/Room 1403 | 2024-08-20 | 2024-08-20 | Record runs as expected after following interim controls and EWI/SOPs
3931 | 68 NYA QCM - No Product - No Lots- Training pH Standardization under incorrect login | judith.keller@sanofi.com | 68 NYA 4700 QCM | 2024-07-02 | 2024-08-16 | It is expected of the analyst to be performing an assay under their own account.
3930 | 8 NYA GEMBA - Settled resin drip tube above gross weight | alex.tello@sanofi.com | 1303 | 2024-08-19 | 2024-08-19 | drip tube weight to be within range of 11g - 13g | Yes | Deviation
3929 | Cold Room 1220 out of temp night of 18AUG24 | genesis.torres2@sanofi.com | 8NYA/ Cold Room 1220 | 2024-08-18 | 2024-08-18 | cold room to stay in temp range while there is gmp material stores inside. | Yes | Deviation
3928 | TSH - 45 NYA Action notice 2024-T-020 gowning clarification/review needed and paint chips throughout rooms 172 and rooms 190 as well as no DP for 174 mag guage. | Jeremy.Wooden@Sanofi.com | 45 NYA - rooms 172, 174 and 190 | 2024-08-19 | 2024-08-19 | It is expected that the Action notice would call for the correct gowning to be used during the 3x facility and manufacturing cleaning. It is also expected that paint chips would be vacuumed/remediated prior to cleanings and that cold room 174 would have the appropriate DP.
3927 | 8NYA APU-Room 1403-product impact unknown-5 UV EasyCal units found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1403 | 2024-07-24 | 2024-08-19 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
3926 | 68NYA Support APU-Metrology-Room 1830-no potential product impact-Pressure Calibrator S-1660 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2024-08-05 | 2024-08-19 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3925 | No offset being applied for Train 3 BGA sample for last 2 sample loops. | joseph.gonzalez@sanofi.com | 1403, 8NYA | 2024-08-18 | 2024-08-18 | Record should be giving adjustment to make the recall. | No | Deviation
3924 | SUB X-2250 PI-2250021 Left in Simulation State of 0.0PSI | brian.o'rourke@sanofi.com | 8NYA Room1403 | 2024-08-17 | 2024-08-17 | Initial inflation set point of 0.25PSI is referenced in MABC-QU-SOP-0080429 step 6.1.100. This setpoint when met would stop the MFC from inflating the SUB bag. Because the indicator remained at 0.0PSI the bag would have been at risk of compromise if not corrected. | Yes | Deviation
3923 | Gemba 8NYA different record version of 705688 used | william.lee@sanofi.com | 8NYA/1350 | 2024-08-17 | 2024-08-17 | Correct version of record should be used.
3922 | Seprafilm Casting machine stopped working during processing. | jan.stepanik@sanofi.com | 76-238 | 2024-08-16 | 2024-08-16 | Casting machine should have kept working. | Yes | Deviation
3921 | 8NYA-Tote Failed to switchover from Growth to Production | steven.rampello@sanofi.com | 1403 | 2024-08-17 | 2024-08-17 | It is expected that the tote will automatically switchover from the appropriate valves | No | Deviation
3920 | Fz Butyl Column C-3110020 2yr Hose PM | zachary.rickard@sanofi.com | 8NYA/1323 | 2024-08-15 | 2024-08-15 | A 2yr Hose replacement should exist for all columns used for production. | Yes | Deviation
3919 | FZ - EGHSD003 - 8NYA - Incubator door not fully closed during day | michael.knapp@sanofi.com | 8NYA 1343 | 2024-08-15 | 2024-08-15 | Door is closed properly | Yes | Deviation
3918 | low incubator temperature from open incubator door | Phoebe.karkos@sanofi.com | 8NYA innoc prep | 2024-08-15 | 2024-08-15 | viability sampling is scheduled to occur 16aug2024, may show any affect to culture
3917 | Power Loss Impacting 8NYA and 78NYA | justin.carpenter@sanofi.com | 8 and 78NYA | 2024-08-15 | 2024-08-15 | Power stays on | Yes | Deviation
3916 | SUB 2230, Production recipe stuck to begin temperature control | bhargav.patel@sanofi.com | 8 NYA / 1403 | 2024-08-15 | 2024-08-15 | Recipe progress for temperature control
3915 | MES failed to launch Media Load product recipe/ Open the next step | kurt.marsocci@sanofi.com | 8 NYA Phase 2 | 2024-08-14 | 2024-08-14 | Fix the automation error. | Yes | No Deviation
3914 | 8NYA-Gemba-Downstream Pompe UF Membrane Uninstall Path Forward on CNC SFO closure | aidan.mula@sanofi.com | 8NYA/RM1503 | 2024-08-15 | 2024-08-15 | Determine path forward on termination of SFO record in the best way possible with MES support.
3913 | X2230 Agitator not turning on after Media Load completed. | sean.gormley@sanofi.com | 8RM1403 | 2024-08-14 | 2024-08-14 | Agitator should turn on when commanded by Bioreactor Recipe. | Yes | No Deviation
3912 | Wood beam exposed during 51NYA shutdown in material airlock 51-127 | nicholas.shoenfelt@sanofi.com | 51-127 | 2024-08-13 | 2024-08-13 | It is expected that wood will not be present in the MFG area. A discussion in needed around whether the beam can be resealed within FRP or needs to be removed.
3911 | 8NYA - Equipment not cleaned correctly in MAL | Craig.stevenson@sanofi.com | 8 NYA Room 1204 | 2024-08-13 | 2024-08-13 | It is expected that all equipment being transfered from an uncontrolled space to a CNC space be wiped with IPA.
3910 | Supplier lot on material does not match supplier paperwork and Athena label | vincent.palumbo@sanofi.com | 11 Forbes | 2024-08-13 | 2024-08-13 | It is expected that the supplier lot is consistent on the material, labels, and documentation.
3909 | Follow up of QE-065454 | zachary.rickard@sanofi.com | 8NYA/1503 | 2024-08-12 | 2024-08-12 | Material will be processed before expiration.
3908 | Multiple Equipment Being Used Before Action Notice Release in 8 NYA | yunbo.lou@sanofi.com | 8 NYA 1503 and 1403 | 2024-08-12 | 2024-08-12 | Equipment restricted by action notice should not be used for GxP processing prior to action notice release. | Yes | Deviation
3907 | 8NYA- X2831Leak observed at Column 3 Steam Thru connection during Buffer Prime operations | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-08-12 | 2024-08-12 | integral MCC system, no leaks | Yes | Deviation
3906 | Cerezyme - EGHBN004 - 8NYA - DeltaV Recipe launched with CCR in incorrect state | michael.knapp@sanofi.com | 8NYA 1403 | 2024-08-08 | 2024-08-11 | CCR is in correct state prior to use. | Yes | Deviation
3905 | 51NYA APU-Purification Dept-Room 87-Bench Scale FB-2373 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 51-87 | 2024-08-09 | 2024-08-12 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3904 | pH out of specification | brendan.dickhaut@sanofi.com | 8NYA Room, Room 1350 | 2024-08-11 | 2024-08-11 | Batch is within acceptable pH parameters
3903 | 8NYA - RF525E027 and RF525E028 leaked from outlet clamp prior to 40 minute cool down | bailey.gleason567@gmail.com | 8NYA RM1373 | 2024-08-11 | 2024-08-11 | N/A
3902 | 8NYA GEMBA- Pompe UFDF-A Hihi Alarm | alex.tello@sanofi.com | 1503 | 2024-08-11 | 2024-08-11 | Recipe to run without any alarms
3901 | Two Cz 500L SUB Bags failed visual inspection per new SOP | kurt.marsocci@sanofi.com | 8 NYA Phase 2 | 2024-08-10 | 2024-08-10 | Per MABC-QU-SOP-0080429 reject the bags | Yes | No Deviation
3900 | Pompe Nanofilter- Premature stock creation | nathaniel.barnes@sanofi.com | 8 NYA 1373 | 2024-08-10 | 2024-08-11 | Determine path forward and usability of Nanofilter. See if permanent fix can be implemented to have MES reflect the accurate 24hr expiry time
3899 | 8NYA-DOWNSTREAM-POMPE-NANOFILTRATION-NANOFILTER LEAK FROM INLET OF FILTER DURING EQUILIBRATION | Marc.Latta@sanofi.com | 1503 | 2024-08-10 | 2024-08-10 | EQUILIBRATION SHOULD PROCEED WITHOUT A LEAK OF NANOFILTER
3898 | PI pull incorrect transmitter value for PH and CO2 probes in MES. | Yves.Antoine@sanofi.com | Room 1403 | 2024-08-10 | 1300 | We expect MES to pill the correct transmitter value in MES
3897 | Mfg at 8 NYA was trying to use Clean steam but they couldn't. F008-PIc-9101003 was found in manual mode and also in alarm. | jesus.acosta2@sanofi.com | 08-1170 | 2024-08-10 | 2024-08-10 | MABC-QU-SOP-0018877, F008-PIc-9101003 should be set to auto mode. | No | Deviation
3896 | Solution transfer setup performed incorrectly for 705664 buffer | christopher.arnold@sanofi.com | 8RM1202 (CNC space outside of Solution Prep) | 2024-08-09 | 2024-08-09 | That the aseptic connection between the extension tubing and the tote bag wye is performed per FBL-EWI-000019. | Yes | Deviation
3895 | Wrong aseptic connection during buffer transfer. | daniela.medina2@sanofi.com | 8NYA/ 1202 CNC | 2024-08-09 | 2024-08-09 | The transfer line has to be on the correct line.
3894 | Multiple Controlled Prints unable to be reconciled per MABC-QU-SOP-0034503 | Kaylyn.Manoogian@sanofi.com | Across all QC Laboratories: QC Chemistry, QC Microbiology at 68NYA, Raw Materials Lab at 11 Forbes | 2024-03-25 | 2024-07-29 | Any Controlled Print generated in error, would be recalled real time per procedure.
3893 | 8NYA-Gemba- Fabrazyme Downstream UFDF skid does not have dirty hold time configured in MES | aidan.mula@sanofi.com | 8NYA/RM1323 | 2024-08-09 | 2024-08-09 | MES should have a configured state diagram for the UF skids for all applicable states prior to validation. | Yes | Deviation
3892 | 8NYA - cover/cap on inside of agitator was dislodged and was noticed in the transfer line prior to filter | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-08-09 | 2024-08-09 | Hold gemba to see path forward | Yes | Deviation
3891 | Defective Amber clamps - assess and determine next steps | leo.gillespie@sanofi.com | 1350 | 2024-08-09 | 2024-08-09 | pivot pins should not move out of position and amber clamp integrity should not be compromised | No | Deviation
3890 | 8NYA-No Product- 500L Cz SUB Bag P/N 753288 failed visual inspection (x6) | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-08-08 | 2024-08-08 | It is expected that SUB bags pass inspection per EWI to be utilized in production activities | Yes | Deviation
3889 | Email Received from Repligen Stating to Return ATF6 devices to Repligen | sheila.pandolfo@sanofi.com | Upstream Suite 1 | 2024-08-08 | 2024-08-08 | It is expected that Sanofi inventory of ATF6 would not be defective. It is expected that inventory be recalled through a formal customer notfication and not via email. (Need to understand if there was a recall from Repligen issued to Sanofi.) | Yes | No Deviation
3888 | Work Order 3182019 Not Contemporaneously Updated During Execution (FDA CAPA) | tanni.arif@sanofi.com | 8NYA room 1350 | 2024-06-20 | 2024-08-08 | contemporaneous updates (ALCOA principal) | No | Deviation
3887 | FZ - EGHCE003 - 8NYA - ATF Failure | michael.knapp@sanofi.com | 8NYA 1333 | 2024-08-08 | 2024-08-08 | ATF remains operational | Yes | Deviation
3886 | Failed Shipping Bag Inspection | katherine.reardon@sanofi.com | 8 NYA 1703 | 2024-08-08 | 2024-08-08 | per FBL-EWI-000029 Shipping Bag Installation step 1. Inspect Bag & Shipping Barrel Sub steps 1-2 Operators are instructed to Open bag and visually inspect for defects. Management should be contacted for all failed inspections. | Yes | Deviation
3885 | Fz Concentration Results for EGHCE003 - 2 OD Out | abel.agnes@sanofi.com | 68NYA QCC | 2024-08-08 | 2024-08-08 | SUB Samples day 9 through 60 are to follow a 9x dilution and be within 20-157.5 ug.
3884 | EGHDT008, Cz2G Thaw time exceeded 38-41 hours | Dang.tran@sanofi.com | 8NYA, CNC | 2024-03-17 | 2024-06-11 | Expected thaw time to occur between 38-41 hours.
3883 | GEMBA - 8 NYA - SP Column HETP out of Range on Qualification | adam.keister@sanofi.com | 8 NYA 1303 | 2024-08-07 | 2024-08-07 | HETP should be within acceptable range
3882 | CZ, RF600E025, CIP, WFI not reaching flowrate during CV rinse samples | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-08-07 | 2024-08-07 | reach flow rate long enough to progress to next sample |  | No Deviation
3881 | 8NYA CZ 2220 Weight was Incorrectly Tared | Jared.Watson@Sanofi.com | 8RM1403 | 2024-08-07 | 2024-08-07 | WT-2220032 was expected to be tared
3880 | 8NYA - WFI valve incorrectly charging for SUM 1520 FV-9257025 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-08-07 | 2024-08-07 | N/A
3879 | Missing witnessed by signature on step 7.17.11 on QU-OPE-0202445 | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2024-08-02 | 2024-08-06 | QU-SOP-0074684. All signatures and entries are to be made sequentially at time of execution
3878 | CO2 Bulk 02AUG24 | NOAH.ZAHN@SANOFI.COM | 68 NYA, 2700 | 2024-08-07 | 2024-08-07 | QU-OPE-0513594: Supplier location is 2 Van Buren BLVD Guilderland Center
3877 | 8NYA-EGHNC010-Pompe Copper blank run ran out of elution buffer. | stanley.durand@sanofi.com | 8NYA rm1503 | 2024-08-07 | 2024-08-07 | It's expected to have enough buffer to complete operations. | Yes | Deviation
3876 | 8NYA - RF623E007 WFI drop for SUM 1520 is not charging WFI correctly | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-08-07 | 2024-08-07 | Hold Gemba to see path forward.
3875 | 2 MFG WFI drops failed weekly flush for third consecutive time. | brendan.dickhaut@sanofi.com | 8NYA Room, US Ballroom phase 1, DS Ballroom Phase 1 | 2024-08-07 | 2024-08-07 | WFI drops flush successfully each week | Yes | No Deviation
3874 | Eluate SUM PBM-2547001 is not currently configured to MCC skid X-2510. | kurt.marsocci@sanofi.com | 8 NYA phase 1 | 2024-08-06 | 2024-08-06 | Configure the new eluate SUM.
3873 | MCC skid X-2831 Buffer flow low low alarm on column 3 during CIP | Yves.Antoine@sanofi.com | Room 1403 | 2024-03-28 | 1201 | We expect the recipe to work properly after FBL-EWI-000004 was followed
3872 | 8NYA - RF1NSE005 FIT failed third time after protocol | bailey.gleason@sanofi.com | CNC | 2024-08-06 | 2024-08-06 | hold gemba to see path forward
3871 | 8NYA - RF1NSE005 FIT failed third time after protocol | bailey.gleason@sanofi.com | CNC | 2024-08-06 | 2024-08-06 | hold gemba to see path forward
3870 | 8NYA - RF621E075 tote required 2 additional seals | bailey.gleason@sanofi.com | CNC | 2024-08-06 | 2024-08-06 | Hold gemba to see path forward |  | No Deviation
3869 | 51NYA July QCA Trend Review PM overdue | adriana.alban@sanofi.com | 51NYA/45NYA/80NYA | 2024-08-01 | 2024-08-06 | It is expected to complete PM's by the end of the month they are due. | Yes | Deviation
3868 | Eluate SUM Bags Coming In Without Foam Frame | kurt.marsocci@sanofi.com | 8 NYA Phase 1 | 2024-08-06 | 2024-08-06 | N/A
3867 | Record calls for wrong part | jose.lugo2@sanofi.com | 8 NYA CNC | 2024-08-06 | 2024-08-06 | Correct SUT asked in record |  | No Deviation
3866 | 8NYA - ANX filter RF1Y6E005 failed twice | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-08-06 | 2024-08-06 | hold gemba to see path forward | Yes | No Deviation
3865 | Asset # 2009351 VACUUM, HEPA, NON-CLEANROOM #2 is missing. | joel.henchar@sanofi.com | 08-1342 | 2023-08-10 | 2023-08-10 | I would like to update the location in HxGN to wherever Facilities stores this equipment, and ensure that the PM is not assigned to manufacturing. LOCATION CHANGE TO 08-1342 PER FACILITIES REQUEST. Created: 21-DEC-2021 07:52, unknown reason.
3864 | 8NYA-Gemba-Downstream Pompe Copper Skid Disassembly step was performed but not signed in EBR | aidan.mula@sanofi.com | 8NYA/RM1503 | 2024-08-06 | 2024-08-06 | Steps should be signed at time of performance | Yes | Deviation
3863 | 8NYA8NYA-DOWNSTREAM-RF1NQE007-UF/DF-A CONJUGATED POOL STORAGE CONDITION DISCREPENCY | Marc.Latta@sanofi.com | 1503 | 2024-08-05 | 2024-08-05 | UF/DF-A CONJUGATED POOL MUST BE STORED AT AMBIANT TEMPERATURE PER FBL-CST-000031.The conjugated pool label must indicate a storage condition referenced in the CST | Yes | Deviation
3862 | SUB record is stuck in sampling loop. Cannot move forward. | brendan.dickhaut@sanofi.com | 8NYA Room 1403, US Ballroom | 2024-08-05 | 2024-08-06 | Operators execute in correct record, corrected elements in EUR do not cause record to become stuck in loop | Yes | No Deviation
3861 | a rupture of a chilled water line at 78NYA. This resulted in a loss of chilled water flow impacting 68, 74, 76, 78 and 80. | paul.mclaughlinjr@sanofi.com | 78NYA | 2024-08-03 | 2024-08-03 | Loop to remain integral
3860 | 8NYA-Gemba-Downstream Oxidation automation intervention needed to restart priming of periodate buffer | aidan.mula@sanofi.com | 8NYA/RM1503 | 2024-08-05 | 2024-08-05 | Pump cover pin holder orientation should be correct. | Yes | Deviation
3859 | 8NYA - Freezer 1421 alarm occurred on 03AUG24 | bailey.gleason@sanofi.com | 8NYA 1421 | 2024-08-03 | 2024-08-05 | Hold gemba to see path forward |  | No Deviation
3858 | 8NYA - 808909 - ANX Eluate a280 sample storage discrepancy | charles.salgadinho@sanofi.com | 8NYA, 1503 | 2024-08-04 | 2024-08-04 | update needed on either sampling plan or the sample labels to have a consistent and accurate storage description
3857 | 8NYA - EGHAE005 - 2000L SUM Bag Failed Pre-use Inspection | wilfredo.lopez@sanofi.com | 8RM1503 | 2024-08-04 | 2024-08-04 | That the SUT is found without foreign materials and in sterile conditions | No | Deviation
3856 | CZ, RF658E007, Gas drop out on X2151 wave during equil | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-08-04 | 2024-08-04 | Gases to maintain during wave stages |  | No Deviation
3855 | Slurry concentration too low | zachary.rickard@sanofi.com | 8NYA/1303 | 2024-08-04 | 2024-08-04 | Slurry SUM weight would be >20L more than the required slurry volume. |  | No Deviation
3854 | 8NYA - 2510 - pressure spike on column 1 | Jared.Watson@Sanofi.com | 8NYA room 1333 | 2024-08-04 | 2024-08-04 | no pressure spike
3853 | EF-12 Fan Down Room 189B and 190 Ceiling Breached | yunbo.lou@sanofi.com | 45 NYA / 189B & 190 | 2024-08-03 | 2024-08-03 | Exhausted Fan is expected to work as intended.
3852 | 8NYA - EGHAE005, 808929 Pump Tubing Rupture during WFI Dilution | wilfredo.lopez@sanofi.com | 8RM1503 | 2024-08-03 | 2024-08-03 | That the SUT in use remains integral throughout the operation
3851 | 8NYA-Cz-X2220 hold due to slip/ trip/fall incident near SUB skid wiring/ Recovery with MSAT and Automation intervention | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-08-03 | 2024-08-03 | It is expected that the wiring remain integral to control the associated equipment during production activities | Yes | Deviation
3850 | 51NYA: rusty water observed coming through the ceiling in rm129 at multiple locations. | adam.rowe@sanofi.com | 51NYA, rm129 | 2024-08-03 | 2024-08-03 | There should be no water leaking through the cieling
3849 | Gemba-8NYA-double amount of Polysorbate called for 705675 batch | william.lee@sanofi.com | 8NYA/1350 | 2024-08-03 | 2024-08-03 | The polysorbate P/N 667330 should only be listed once in the BOM. Only 5.104 L is required for this batch. | Yes | Deviation
3848 | 8NYA-Fz-X2200 pCO2 Probe divergence | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2024-08-02 | 2024-08-02 | it is expected that the controlling probes remain within allowable limits during production activities | Yes | No Deviation
3847 | 8NYA-Fz-Eluate SUM2 clamp not fully engaged at start of mini eluate collection | JOHN.LYTLE@SANOFI.COM | Rm 1333, 8NYA | 2024-08-03 | 2024-08-03 | it is expected that the bottom outlet clamp be properly deployed on the SUM bag, preventing product from entering the line leading to the Aseptiqwik connection | Yes | Deviation
3846 | GEMBA - 8 NYA - Copper Load GPT slightly out of upper range | adam.keister@sanofi.com | 8 NYA Room 1503 | 2024-08-02 | 2024-08-02 | To evaluate GPT and have production continue
3845 | CZ particulate observed during SUB bag inspection | Jared.Watson@Sanofi.com | 8NYA room 1403 | 2024-08-02 | 2024-08-02 | No particulates within the bag
3844 | 8NYA-Cz-X2220 pH Probe 1 Atypical Trend/ Resample requested | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-08-02 | 2024-08-02 | it is expected that a successful 1point calibration continue to control the system as expected via air sparge | Yes | No Deviation
3843 | During a routine Calibration, HEPA Filter F-2082 Failed As-Found in Drying room 76-234 | jan.stepanik@sanofi.com | 76 NYA/ 234 Drying Room | 2024-08-02 | 2024-08-02 | Should have passed As-Found
3842 | Gemba-8NYA-WFI valve not operating on SUM 1505 during batching | william.lee@sanofi.com | 8NYA/1350 | 2024-08-02 | 2024-08-02 | WFI should fill at 2300L and recipe should proceed with batching. | Yes | No Deviation
3841 | 8NYA - Welder 2012069 stuck on line while performing weld on RF1QJE007 | bailey.gleason@sanofi.com | CNC | 2024-08-01 | 2024-08-01 | Have engineering assess | Yes | No Deviation
3840 | Gemba: 45/51 Bacillus Cereus Identified in CV | amy.roy@sanofi.com | 68 NYA 4700 QCM | 2024-07-25 | 2024-08-01 | Expected bioburden CV to not obtain Bacillus Cereus
3839 | GEMBA-8NYA-Downstream-Copper set-up BOM issues | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 room 1503 | 2024-08-01 | 2024-08-01 | The record for copper set-up should be adjusted to include 2 10ft extensions 714010 in the eluate SUM take-outs and the take-out section of the skid set-up should be adjusted to scan in a bubble trap filter 717322
3838 | Cold room 138 high temperature (15.7C) alarm at 51NYA | sylvester.kegbeh@sanofi.com | 51 NYA ROOMS 138 | 2024-08-01 | 2024-08-01 | it is expected that the cold temp stay with the range of 6-10C
3837 | Fz Butyl Load Bag inspection failed Material\692299 per FBL-EWI-000015 | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1323 | 2024-08-01 | 2024-08-01 | To get a Phath forward
3836 | Seprafilm - EBYSEP050B - HA-CMC lot EBYCMC097B with OOS used in Seprafilm manufacturing | juan.ramos-torres@sanofi.com | 76 NYA 237 | 2024-07-28 | 2024-08-01 | HA-CMC lot with OOS should not be use in manufacturing activities
3835 | Late Extended PM at 76 NYA | kevin.southwick@sanofi.com | 76 NYA | 2024-08-01 | 2024-08-01 | It is expected that work orders are completed by their due date per QU-SOP-0072346. | Yes | Deviation
3834 | 76NYA Biosurgery APU-Room 3102 Pressure Gauge F-18957 was found out of tolerance during an unscheduled calibration. | robert.bauch@sanofi.com | 76-3102 | 2024-07-31 | 2024-08-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3833 | Gemba: 45/51-TSH Bioburden UF/DF Load Leaking Sample | amy.roy@sanofi.com | 68 NYA 4700 QCM | 2024-07-31 | 2024-07-31 | Expected bioburden sample to be intact, sterile, and not have leaks.
3832 | X2200 Low Viability | sean.gormley@sanofi.com | 8RM1334 | 2024-08-01 | 2024-08-01 | Viability should be above 80 percent
3831 | Expired sterile 50L carboys used for JIT | kevin.ornell@sanofi.com | 80 NYA room 102 | 2024-07-31 | 2024-07-31 | QU-OPE-0273744 step 4.0 | Yes | Deviation
3830 | CBYHAS005 24M Time point missed for Appearance and Endotoxin | kristina.erickson@sanofi.com | QCBP 76 NYA RM 107 | 2024-05-29 | 2024-07-31 | Endotoxin Weighed out and tested with Micro and Appearance Testing completed within window of 25May24-24Jun24
3829 | 8 NYA - Cerezyme - EGHDS002 - Mixing not performed at Validated rate | catherine.kuzdzal@sanofi.com | 8 NYA Suite 2 DOwnstream | 2024-01-23 | 2024-07-29 | It is expected that mixing is performed using qualified parameters per PQR – F008-MIX-200-8NYA-CQ-0009
3828 | Missing PMs on LFH F008-INC-100 | megan.perkins@sanofi.com | Room 1343 at 8NYA | 2020-06-08 | 2024-07-30 | Preventative maintenance is scheduled for this LFH per the new equipment form. | Yes | Deviation
3827 | Tote's calibration date was extended already and tote was due 31Jul24 but is in use | dennis.bazinet@sanofi.com | 8 NYA / CNC space | 2024-07-31 | 2024-07-31 | QU-SOP-0069490 calibration program section 5.8.6.3 "Only one calibration extension can be granted per instrument for each calibration interval." |  | No Deviation
3826 | Personnel reviewed step that they performed | karthik.sriram@sanofi.com | 76 NYA - 231 | 2024-07-15 | 2024-07-30 | Per QU-SOP-0074684-Good Documentation Practices, step 6.11.1, Personnel will not review, verify or witness a step they have performed.
3825 | LOOPS 3A ,3B AND 4 SHUT DOWN AT 1930 , 30 JUL 2024 | MACJONES.OSEIKUFFOUR@SANOFI.COM | 8NYA, ROOM 1170 MECHANICAL ROOM | 2024-07-30 | 2024-07-30 | MABC-QU-SOP-0001515
3824 | Recount not initiated for Cycle Count at 51NYA | Joshua.brooks@sanofi.com | 51NYA Bullpen | 2024-07-28 | 2024-07-29 | When the entered counted quantities do not align with the quantities in the ERP system, it is expected that a recount is initiated within the ERP prior to clearing differences.
3823 | Gemba: R&D LAL sample incorrect standard curve used | Pushti.Shah@sanofi.com | 4700 68 NYA QC Micro | 2024-07-16 | 2024-07-30 | QU-SOP-0014245 6.7.3
3822 | GEMBA Upstream 8 NYA VI Skid X-2825 PH Probe Temperature Sensor Ti-2825011 Not Scaled Correctly in DeltaV, Scaled to 0-100 Verses 0-80 | GARY.DODAKIAN3@SANOFI.COM | 8NYA / Upstream / 1403 | 2024-07-29 | 2024-07-29 | Scaling should be correct in DeltaV for temperature sensors |  | Deviation
3821 | QC PM5 analytical worksheet obtained outside controlled print | cleiton.lopes@sanofi.com | 68NYA | 2024-07-26 | 2024-07-30 | Analytical worksheets and ID obtained via controlled issuance
3820 | 8NYA- Three Eluate Pool Collection Bags (PN 692168) failed Pre-Use Inspection | timothy.ryley@sanofi.com | 1403 | 2024-07-30 | 2024-07-30 | Bag film should not contain anomalies
3819 | GEMBA-8NYA-Tote bag failed inspection | danny.martinez@sanofi.com | 8NYA/ RM 1323 | 2024-07-30 | 2024-07-30 | Bag should pass inspection
3818 | 51NYA general use valve PM overdue 3061512 | adriana.alban@sanofi.com | 51-143 | 2024-06-30 | 2024-07-30 | It is expected that equipment is available for PM
3817 | 2250x stuck in "restarting" phase after recipe went into hold- automation intervention. | joseph.gonzalez@sanofi.com | 8NYA room 1403 | 2024-07-30 | 2024-07-30 | recipe should be able to resume without automatipon
3816 | Unable to complete Seprafilm Lot EBYSEP050B, Part 768276. | jan.stepanik@sanofi.com | 76-237 | 2024-07-28 | 2024-07-28 | The container of HCl should have been poured in to the tank
3815 | Operators with Supervisor User Rights in MES Without Delegate Form | sabrina.waring@sanofi.com | 8NYA | 2024-07-25 | 2024-07-25 | It is expected that a delegate form would be filled out when granting personnel delegate authority to perform tasks outside of their role |  | No Deviation
3814 | Protective membrane from aseptic connection not cleanly removed during installation | christopher.arnold@sanofi.com | 8RM1202, the CNC space outside of Solution Prep | 2024-07-26 | 2024-07-26 | That the removal of the membrane will result in two integral barrier pieces removed, leaving no remaining barrier inside the Aseptiquik connection. | Yes | No Deviation
3813 | Delta V EM stuck and needs automation intervention | robert.morawski@sanofi.com | 8NYA RM 1210 | 2024-07-25 | 2024-07-25 | EM would perform tote switchover properly. |  | No Deviation
3812 | 8NYA-No Product- SUT Inspection Failure | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1333 | 2024-07-25 | 2024-07-25 | It is expected that SUT pass visual inspection prior to use
3811 | Discuss Tote Transport Step in Solution Prep EBRs | robert.morawski@sanofi.com | 8NYA CNC RM 1210 | 2024-07-25 | 2024-07-25 | Path forward on Tote Transport | Yes | No Deviation
3810 | GEMBA - UFDF-B Retentate Filter Failed Twice | JACK.LAPLANTE@SANOFI.COM | 8NYA 1373 | 2024-07-25 | 2024-07-25 | Looking for path forward.
3809 | 8NYA-Pompe- Train 4 End of Run will occur during NTO on 29JUL25. Meeting to discuss extending run to collect full eluate | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 1403 | 2024-07-25 | 2024-07-25 | EOR occurs after all eluates have been collected |  | No Deviation
3808 | 51NYA Thyrogen APU-Room 140-Filter Integrity tester F-15855 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 51-140 | 2024-07-17 | 2024-07-25 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | No | Deviation
3807 | 8NYA-Fz-Eluate SUM 1 outlet Aseptiquik in contact with excessive condensate | JOHN.LYTLE@SANOFI.COM | 8NYA Rm 13333 | 2024-07-25 | 2024-07-25 | It is expected that environmental controls utilized for Mfg are effective at maintaining specific temperature and humidity ranges
3806 | Multiple Fz buffers with double Acetic Acid BOM positions | christopher.arnold@sanofi.com | 8NYA 8RM1350 | 2024-07-23 | 2024-07-23 | That all released for mfg execution EBRs will be correct.
3805 | Conductivity Standard Failing to Meet Acceptance Criteria for Testing | jessica.kychler@sanofi.com | QC Micro 68 New York Avenue, floor 4, room 4700 | 2024-07-24 | 2024-07-24 | Per step 6.3.3 of QU-OPE-0115243, the acceptance criteria for the reference read is 81.48 – 86.52 µS/cm
3804 | Look into Vendor Qualifications to perform chamber maintenance on stability incubator N-0487 in room 1318 | kathy.tourville@sanofi.com | 11 Forbes Road, Northborough | 2024-07-17 | 2024-07-24 | Vendor would only perform work authorized for, instrument left in validated state.
3803 | Gemba-8NYA Damaged SUT Aseptiquik connection RF1YIE053 | william.lee@sanofi.com | 8NYA/1350 | 2024-07-24 | 2024-07-24 | Seal should not have been broken on aseptiquik x connection | Yes | Deviation
3802 | X-2220 Bleed Pump did not restart after pressure Hold | Yves.Antoine@sanofi.com | 1403 | 2024-03-28 | 1201 | We expect the recipe to work properly after pressure alarm was cleared and system restart
3801 | 8NYA - RF1NQE003 UF A filter untested after QE-031706. | bailey.gleason@sanofi.com | 8NYA 1420 | 2024-07-23 | 2024-07-23 | Hold gemba to see path forward
3800 | sample 2400008922 entered into lims during validation | Stefanie.Erickson@sanofi.com | QC Micro counting | 2024-01-09 | 2024-01-27 | for the iLIMS system to fully go through the validation process before entering results.
3799 | 8NYA - EGHBJ019 Samples not in Biolims. No point of batch documented | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-05-18 | 2024-07-23 | N/A
3798 | Foreign material detected in 60 g bottle | valeria.moreno@sanofi.com | 76-355 NYA | 2024-07-01 | 2024-07-19 | QU-OPE-0039925, Step 7.9.1
3797 | 8NYA-Gemba-Downstream 8NYA SP Column Pack Leak from Bottom Mobile Phase Piping During Column Prime | aidan.mula@sanofi.com | 8NYA/RM1303 | 2024-07-23 | 2024-07-23 | Determination on path forward at gemba | Yes | No Deviation
3796 | 8NYA - EMS alarm for out-of-range temperature on -20 freezer RM1421 | bailey.gleason@sanofi.com | 8NYA RM1421 | 2024-07-23 | 2024-07-23 | N/A | Yes | No Deviation
3795 | 8NYA-MANUFACTURING -DOWNSTREAM-POMPE-UF/DFA CIP-RF1NQE006-NWP FAILED on High recovery | Marc.Latta@sanofi.com | 1503 | 2024-07-23 | 2024-07-23 | RECIPE SHOULD PASS NWP RECOVERY
3794 | Freezer Excursion | jacqueline.rapko@sanofi.com | Room 3700, side room at the end of Bay 10 | 2024-07-22 | 2024-07-19 | QU-SOP-0068547, Attachment 1 -50C to -90C
3793 | 8NYA Pompe downstream UF\DF-A CIP recipe ran without Rupture disk sensor | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1703 | 2024-07-22 | 2024-07-22 | To get a Phath forward | Yes | No Deviation
3792 | Unrecorded contact time prior to entering BSCs | Craig.stevenson@sanofi.com | 76 NYA Room 386B and 412 | 2024-07-14 | 2024-07-22 | It is expected that the start and end time be recorded in QU-FOR-0111567 prior to BSC entry and work.
3791 | CV samples taken at incorrect location during UFDF A CIP | pavel.dobrovolskiy@sanofi.com | 8NYA 11503 | 2024-07-22 | 2024-07-22 | Its expected that the samples are taken at the correct location | Yes | No Deviation
3790 | Cerezyme - RF639E002 - 8NYA - Automation Intervention Required for MCC X-2731 End of Run | michael.knapp@sanofi.com | 8NYA 1403 | 2024-07-21 | 2024-07-22 | Correct number of elutions are collected and end of run prompt appears at completion of 6th eluate pool. | Yes
3789 | SP1 EGF2B07B pH is out of range during Chromatogram Review | sylvester.kegbeh@sanofi.com | 51 NYA ROOMS 138 | 2024-07-10 | 2024-07-20 | it is expected that the SP1 pH is within the range of 4.8-5.6 during Equil, Load, wash , and elution per QU-SOP-0061107 attachment 3
3788 | Fabrazyme - EGHCE003 - 8NYA - Harvest Bag Replacement Automation Requirements with Harvest Weight Alarm | michael.knapp@sanofi.com | 8NYA 1333 | 2024-07-22 | 2024-07-22 | A clear path with automation for the replacement of the harvest bag.
3787 | Gemba 8NYA -Sealer stuck in close position during batching of Mannitol buffer 820323 | william.lee@sanofi.com | 8NYA/1350 | 2024-07-21 | 2024-07-21 | Sealer should be able to open up. Transfer to continue to barrel.
3786 | GEMBA - 8 NYA - Damaged 1K SUM bag | william.lee@sanofi.com | 8NYA/1350 | 2024-07-21 | 2024-07-21 | Agitator should not have moved out of place and bent.
3785 | 8NYA #1404 -Pest Control | jeffrey.boldizar@sanofi.com | 8NYA 1404 | 2024-07-21 | 2024-07-21 | No pest in controlled environment
3784 | 8NYA-Fz-X2200 SUB TCU troubleshooting | JOHN.LYTLE@SANOFI.COM | Rm 13333 8NYA | 2024-07-20 | 2024-07-20 | It is expected that the TCU control the SUB temperature and maintain a steady operational state
3783 | Freezer temperature excursion | jacqueline.rapko@sanofi.com | Room 3700, side room at the end of Bay 10 | 2024-07-19 | 2024-07-19 | QU-SOP-0068547, Attachment 1 -50C to -90C
3782 | 8NYA - Copper filter #1 RF1Y5E004 failed twice | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-07-18 | 2024-07-18 | Hold gemba to see path forward
3781 | Late RQ for Stability INCUBATOR 1 | kathy.tourville@sanofi.com | Room 1318A | 2024-07-01 | 2024-07-18 | RQ would be completed by 30Jun24
3780 | Re-GEMBA Cerezyme - EGHBN003 - 8NYA - SUB X2220 DO Low Low | ralston.najarro@sanofi.com | 8NYA1403 | 2024-07-18 | 2024-07-18 | DO should be controlled to setpoint
3779 | Temp controller failure on SUB X-2200 | brendan.dickhaut@sanofi.com | 8NYA Room 1333, US Ballroom | 2024-07-17 | 2024-07-17 | TCU operates as expected and maintains temperature of bioreactor.
3778 | 8NYA - Continuation of EMS alarms in cold room 1421 | bailey.gleason567@gmail.com | 8NYA RM1421 | 2024-07-17 | 2024-07-17 | Hold Gemba to document the alarms after QE-057564
3777 | Gemba - 8NYA- Support Services -Continuous Improvement CAPA request | alejandro.martinez2@sanofi.com | 1350/CNC space | 2024-07-17 | 2024-07-17 | A CAPA to reference these continous improvements to avoid getting kickbacks from QA when reviewing.
3776 | 8NYA Gemba/Pompe copper pooling EWI/BOM discrepancy | lukus.palzkill@sanofi.com | 8NYA downstream suite 2 room 1503 | 2024-07-17 | 2024-07-17 | wyes will be used instead of manifold
3775 | GEMBA - 8 NYA - Pompe Nano Record Cannot Consume Load Material | adam.keister@sanofi.com | 8 NYA 1503 | 2024-07-11 | 2024-07-11 | For nano load material to be consumed into record as intended. |  | No Deviation
3774 | Shipping Bag Failure | katherine.reardon@sanofi.com | 8 NYA 1703 | 2024-07-16 | 2024-07-16 | per FBL-EWI-000029 Shipping Bag Installation step 1. Inspect Bag & Shipping Barrel Sub steps 1-2 Operators are instructed to Open bag and visually inspect for defects. Management should be contacted for all failed inspections. |  | No Deviation
3773 | GEMBA - 8 NYA - SP Column Failed Qualification for Third Time | adam.keister@sanofi.com | 8 NYA 1303 | 2024-07-16 | 2024-07-16 | For Asymmetry and HETP values to be in range after column pack
3772 | Suprvisor did section review and final review of OPEs they verified in QU-OPE-0520874 and QU-OPE-0511911 | dianne.moustafa@sanofi.com | 76 NYA CMC | 2024-01-31 | 2024-07-16 | QU-SOP-0074684 step 6.11.9 Reviewd by must not be same person who preformed activity.
3771 | Seprafilm- no lots- 76NYA- RM233- Metrology personal not wearing a mask in Grade C area | nicole.fragaalvarez@sanofi.com | 233 | 2024-07-15 | 2024-07-15 | mask must be worn in Grade C areas
3770 | 8NYA-Pompe- Condensate build up in X2250 SUB Vent Filter Line | JOHN.LYTLE@SANOFI.COM | Rm 1434 | 2024-07-16 | 2024-07-16 | It is expected that the heated jacket allow the system to moderate condensate levels
3769 | -20 Cold Room 1421 out of range | robert.morawski@sanofi.com | 8 NYA RM 1421 | 2024-07-16 | 2024-07-16 | Cold room maintains temperature
3768 | 51NYA Thyrogen APU-cold room 138, F-3933-300 EMS found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 51-138 | 2024-07-16 | 2024-07-16 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3767 | 8NYA-No Associated Product- SUTx5 Failed Mfg. Visual Inspection | JOHN.LYTLE@SANOFI.COM | 8NYA Rm1403 | 2024-07-15 | 2024-07-15 | Passing visual Inspection/ No signs of damage or defect
3766 | Process Testing pH result high for 805009 | christopher.arnold@sanofi.com | 8RM1350 | 2024-07-16 | 2024-07-16 | That all process control testing will result in a sample within process parameters. Additionally, that the process range is set accordingly so that normal process testing isn't consistently riding the very top of a process range.
3765 | Gemba - 8NYA - Automation needed to end DeltaV recipe. | kenny.leland@sanofi.com | 1703 | 2024-07-15 | 2024-07-15 | Recipe should end after 2 or more shipping bags have transfered
3764 | NBO - Supplier Deviation Required. Contaminated RODAC plates | Craig.stevenson@sanofi.com | NBO | 2024-07-15 | 2024-07-15 | It is expected that the RODACs are not contaminated when opened from the Supplier.
3763 | 8 NYA - WFI Loop 3 breach | Craig.stevenson@sanofi.com | 8 NYA Room 1170 | 2024-07-15 | 2024-07-15 | It is expected that the WFI system not leak during QC sampling.
3762 | GEMBA - 8 NYA - SP Column Failed Qualification Twice | adam.keister@sanofi.com | 8 NYA 1303 | 2024-07-14 | 2024-07-14 | For SP column to be packed and have passing qualification results.
3761 | 68NYA QCM support APU-Room 4700-Dew point Hygrometer FB-2679 found non-operational no as found data collected | robert.bauch@sanofi.com | 68-4700 | 2024-07-11 | 2024-07-15 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3760 | 68NYA support APU-room 4700-Balance FB-5088 no as found data collected | robert.bauch@sanofi.com | 68-4700 | 2024-07-11 | 2024-07-15 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3759 | Cerezyme - EGHBN003 - 8NYA - SUB X2220 DO Low Low | michael.knapp@sanofi.com | 8NYA 1403 | 2024-07-15 | 2024-07-15 | System maintains DO control | Yes | Deviation
3758 | Pressure event led to automation intervention on VI skid. | brendan.dickhaut@sanofi.com | 8NYA Room 1403, US Ballroom | 2024-07-14 | 2024-07-14 | Transfers occur without incident, holds, or alarms
3757 | TSH - Buffer Prep - GZ6574-01 filter failed post test 3 times | Jeremy.Wooden@Sanofi.com | 45 NYA Room 190, 51 NYA room 140 | 2024-07-13 | 2024-07-13 | It is expected that the filter would not fail 3 post tests. | No | Deviation
3756 | 8 NYA downstream shipping barrels recipe RF1E4E020#103 Aborted no additional barrels needed. | Christian.Venegas@Sanofi.com | 8NYA Cell Culture Phase 2 Room 1703 | 2024-07-13 | 2024-07-13 | To get a Phath forward
3755 | 8 NYA downstream shipping barrels recipe failed to lunch on MES record | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1703 | 2024-07-13 | 2024-07-13 | Need Phath forward to continue process |  | No Deviation
3754 | DO Lo-Lo alarm on X2220 | ralston.najarro@sanofi.com | 8NYA1403 | 2024-06-30 | 2024-06-30 | DO should be controlled to setpoint |  | No Deviation
3753 | Automation intervention required to avoid 10 hour batch timer for Fz Media | christopher.arnold@sanofi.com | 8RM1202, the CNC space outside of Solution Prep | 2024-07-13 | 2024-07-13 | That recipe prompts would be clear about what they were asking. That the EWI would have guidance for non-routine tasks like seed media.
3752 | failure of membrane while mating asteptiquick G and S connections on a filter to manifold being set up for eluate transfer. | brendan.dickhaut@sanofi.com | 8NYA Room 1403, US Ballroom | 2024-07-12 | 2024-07-12 | Membrane does not fail/tare while mating aseptiquick connections |  | No Deviation
3751 | Third Party Vendor Event required for Alcohol Bulk Delivery | sarah.adams@sanofi.com | 68 NYA, QCC | 2024-07-12 | 2024-07-12 | Tucks seals match provided paperwork.
3750 | TSH / Buffer cold room 174 out of range for temperature | kevin.ornell@sanofi.com | 45-174 | 2024-07-10 | 2024-07-12 | Expected temperature would stay in range 2-8 DegC
3749 | Leaking 3K tote PT# 705665 MO RF5444E007 for Pompe Copper line in Phase 2 | dennis.bazinet@sanofi.com | 8 NYA /1212 Phase 2 area | 2024-07-12 | 2024-07-12 | Tote not leaking
3748 | 1K SUM bags rejected during pre-use inspection | christopher.arnold@sanofi.com | 8RM1350 | 2024-07-11 | 2024-07-11 | That no defects would be detected during pre use inspection.
3747 | Critical Utility EM settle plate unverifiable open and closed time | yanitza.montalvo@sanofi.com | 68NYA, QC Micro | 2024-07-03 | 2024-07-11 | QU-SOP-0057388 Step 6.2.2. Ensure agar surface is not exposed for greater than the time specified below: IRRADIATED TSA (TRIPLE WRAPPED)(DEEP FILL 4 HOUR MAX EXPOSURE); GZ95000585 (BBL CATALOG # 292652 OR 215343) – MAX EXPOSURE TIME 4 HOURS.
3746 | Critical Utility EM settle plate unverifiable open and closed time | yanitza.montalvo@sanofi.com | 68NYA, QC Micro | 2024-07-03 | 2024-07-11 | QU-SOP-0057388 Step 6.2.2. Ensure agar surface is not exposed for greater than the time specified below: •IRRADIATED TSA (TRIPLE WRAPPED)(DEEP FILL 4 HOUR MAX EXPOSURE); GZ95000585 (BBL CATALOG # 292652 OR 215343) – MAX EXPOSURE TIME 4 HOURS.
3745 | FZ, MCC 2510, SIP filter broken during trouble shooting pressure test | jeffrey.boldizar@sanofi.com | 8NYA 1333 | 2024-07-11 | 2024-07-11 | SUT not to break
3744 | Unverifiable CSE Solution | lauren.ruffing@sanofi.com | 68NYA 4740 QC Micro | 2024-07-03 | 2024-07-11 | The CSE solution that corresponds to the assay type is written on assay paperwork
3743 | 8NYA-No Associated Product-Expired material found in GMP area during walkthrough activities | john.lytle@sanofi.com | Rm 1343 8NYA | 2024-07-11 | 2024-07-11 | It is expected that material be discarded if the expiry has been exceeded, and while investigating the issue it was found that the EWI instruction does not direct the operator to the correct instruction
3742 | GEMBA - 8 NYA - Colum Pack Failed Qualification Recipe | adam.keister@sanofi.com | 8 NYA, room 1303 | 2024-07-11 | 2024-07-11 | For asymmetry and hetp results to be within expected range in MES.
3741 | High Humidity and Heat in 8NYA, Specifically Upstream Suite 2 Room 1403 | gary.dodakian3@sanofi.com | 1403 | 2024-07-10 | 2024-07-10 | A comfortable working environment in the suite.
3740 | High Humidity and Heat in 8NYA, Specifically Upstream Suite 2 Room 1403 | gary.dodakian3@sanofi.com | 1403 | 2024-07-10 | 2024-07-10 | A comfortable working environment in the suite.
3739 | SUM bag particulate found during inspection | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-07-10 | 2024-07-10 | no particulates present
3738 | Tyvek Holder P/N GZ6374-01 perforation issue. Lot 24MN00126 vendor: Oliver | nicole.fragaalvarez@sanofi.com | room 233 | 2024-07-10 | 2024-07-10 | tyvek holders meet specs with no major gaps
3737 | Incorrect dispense/draw speed parameters in QU-OPE-0202442 instrument method | aimee.park@sanofi.com | 45NYA, QC Chemistry | 2024-07-10 | 2024-07-10 | Default Instrument method parameters would have been set correctly and used for each assay occasion on 70-QC-0623.
3736 | QU-OPE-0220031 revised and sample volume inadvertently removed | aimee.park@sanofi.com | 45NYA, QC Chemistry | 2024-06-24 | 2024-07-10 | Sample volume for the pre-dilution step would have been documented correctly.
3735 | 8NYA-Cz- X2220 DO probe 1 intermittently dropping below processing requirements | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-07-10 | 2024-07-10 | no alarms
3734 | 8NYA-Pompe- X2831 Post Use CIP Low Flow on Base Line needs additional troubleshooting to determine root cause | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-07-09 | 2024-07-09 | it is expected that the base line and associated equipment provide adequate flow to perform cleaning operations
3733 | UFDF-B Retentate filter failed first FIT | christopher.arnold@sanofi.com | 8RM1290 | 2024-07-09 | 2024-07-09 | That a filter will pass FIT testing after WFI wetting.
3732 | Gemba-8NYA- Pompe UF-B RETENTATE filter | alejandro.martinez2@sanofi.com | 8RM1290 | 2024-07-09 | 2024-07-09 | Written filter label and email should reference accurate part number.
3731 | 2X Shipping bags rejected during pre-inspection for EGHDS003 | Christian.Venegas@Sanofi.com | 8NYA Cell Culture Phase 2 Room 1703 | 2024-07-09 | 2024-07-09 | To get a Phath forward
3730 | 51NYA Room 152B - QC Micro - Dewpoint Failure in BSC F-16386 | Craig.stevenson@sanofi.com | 51 NYA Room 152B | 2024-07-09 | 2024-07-09 | It is expected that dewpoint testing generate a result of >-40C
3729 | rhTSH Harvest in the 500 L TSH Purification Tank is not between 6 C to 10 C. | Luis.dominguez@sanofi.com | 51NYA room 139 | 2023-11-23 | 2023-11-23 | per QU-OPE-0351657 temperature must be between 6-10C at step 4.7.4
3728 | 68 NYA - QCM - Insufficient Volume of Pompe Drug Substance for Bioburden Testing | joseph.napoli@sanofi.com | 68 NYA 4700 QCM | 2024-07-08 | 2024-07-09 | for there to be at least 100mL of sample submitted for bioburden testing. only 37mL were in the bag.
3727 | 76 NYA - HA Fill - Incomplete Non-Viable Site CP001 | Craig.stevenson@sanofi.com | 76NYA, Room 355 | 2024-06-26 | 2024-07-09 | It is expected that all climets be aligned and have the same number of continuous readings.
3726 | 8NYA - Discussion around pompe media SUM went into alarm which tared the media tote. delta V generated 2099 as value. Automation could not enter in correct value | bailey.gleason@sanofi.com | 8NYA RM1350 and CNC | 2024-07-08 | 2024-07-08 | Hold gemba to see path forward
3725 | Manufacturing Staff not trained on current version of OJT, MABC-QU-FOR-0005105-Welding and Sealing Operations of Single Use Tubing at 8 NYA, V4.0 | danielle.woods@sanofi.com | 8 NYA | 2024-03-21 | 2024-07-08 | Operators that are performing sealing and welding operations should be trained on the current version of the OJT
3724 | 8NYA suite 2 Downstream incorrect Pompe Nano formulation set up was performed | Christian.Venegas@Sanofi.com | 8NYA Cell Culture Phase 2 Room 1503 | 2024-07-08 | 2024-07-08 | To get a Phath forward |  | No Deviation
3723 | FAC rejected WO for QC Sampling | Craig.stevenson@sanofi.com | 8NYA Room 1412 | 2024-07-08 | 2024-07-08 | It is expected that WOs used as justification to reject GxP samples not be rejected.
3722 | POMPE, RF1NAE008, MCC Resin failure on column 3 | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-07-07 | 2024-07-07 | Resin to capture product
3721 | GEMBA - 8 NYA - Nano Filter Leaking on Outlet Clamp | adam.keister@sanofi.com | 8 NYA, 1503 | 2024-07-07 | 2024-07-07 | Equil should run and finish with no leaks |  | Deviation
3720 | RF639E003, CZ, SUT SUM filter failed visual inspection before use | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-07-06 | 2024-07-06 | SUT not be damaged | Yes | Deviation
3719 | POMPE, RF1NAE008, VI skid PH Probe recalibration PI pull pulls SUB probe value | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-07-06 | 2024-07-06 | PI to pull the VI skid mixer value
3718 | Bioburden raw material samples tested with wrong rinse fluid | roxanne.vacha@sanofi.com | 68 NYA, 4700, QC Micro | 2024-06-26 | 2024-07-06 | Per QU-SOP-0056015, samples 1272279 and 1272237 should have been rinsed with 3X100 mL DPBS.
3717 | 8NYA Gemba-leaking sample | william.lee@sanofi.com | 8NYA/CNC | 2024-07-06 | 2024-07-06 | Sample should not have leaked
3716 | MCC 2831, CIP, flow rate reaching but not maintaining enough flow for cleaning | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-07-06 | 2024-07-06 | Flow to maintain for a proper clean of the MCC skid | Yes | No Deviation
3715 | 8NYA - EGHMD045, 805013 Pompe Media Buffer Leak on Transfer/Fill Line | wilfredo.lopez@sanofi.com | 8RM1212 | 2024-07-05 | 2024-07-05 | It is expected that the seals remain integral once the totes are disconnected after the transfer. | Yes | Deviation
3714 | 8NYA - EGHAV055, 760741 Cerezyme Media Buffer Leak on Transfer/Fill Line | wilfredo.lopez@sanofi.com | 8RM1212 | 2024-07-05 | 2024-07-05 | It is expected that the seals remain integral once the totes are disconnected after the transfer. | Yes | Deviation
3713 | Fabrazyme - RF503E003 - 8NYA - Leak on ATF6 during ATF wetting | michael.knapp@sanofi.com | 8NYA 1333 | 2024-07-06 | 2024-07-05 | No leak on ATF |  | Deviation
3712 | CZ - RF639E003 - 8NYA - Capacitance High High Alarm | michael.knapp@sanofi.com | 8NYA 1403 | 2024-07-05 | 2024-07-05 | It is expected that the bleed controls capacitance | Yes | No Deviation
3711 | Thyrogen product placed on restricted freezer shelf | jacob.kallon@sanofi.com | 80 NYA Room 150 | 2024-07-05 | 2024-07-05 | Sign on freezer states bagged product should not be stored on shelf 3
3710 | MCC Stock creation EUR Harvest Day step not execute Automatically as intended in record | bhargav.patel@sanofi.com | 8 NYA / 1403 | 2024-06-24 | 2024-06-24 | Harvest Day step auto execute
3709 | pNp SUB sample not diluted for CZ eluate 2 Batch EGHBN002 | elaine.jansen@sanofi.com | 8NYA 1403 | 2024-06-27 | 2024-07-02 | pNp SUB sample needs to be diluted per QU-WIN-0046060 |  | Deviation
3708 | MCC Production recipe send VI tag for Trian 3 2865 instead of 2825 | bhargav.patel@sanofi.com | 8 NYA / 1403 | 2024-03-30 | 2024-03-30 | Automation adjustment should not require to adjust pump speed
3707 | Superseded MBR used for production (SP2 EGF3B01A) | jacob.kallon@sanofi.com | 51 NYA Room 132 | 2024-06-25 | 2024-07-05 | Per QU-SOP-0071158 step 6.5.1:Once a Production Record Package (i.e. PR documents and any applicable worksheets) is issued, the owning Department MUST verify the effective date as current prior to use if operations commence after the issuance date.
3706 | Missing one dry material bag | maria.palmar@sanofi.com | 8 New York Ave, room 1350. | 2024-07-03 | 2024-07-03 | To not miss any dry material bag
3705 | 76NYA - HA Fill/Records Management - Batch information not printed on paperwork | Craig.stevenson@sanofi.com | 76NYA, Room 355 | 2024-06-03 | 2024-07-03 | It is expected that the batch information be printed on each page
3704 | 8NYA - 808210 Pompe UF-A Load Filter Failed Integrity Testing Twice | wilfredo.lopez@sanofi.com | 8RM1350 | 2024-07-03 | 2024-07-03 | Expected that the filters pass integrity testing within the first two allowed attempts.
3703 | A GEMBA is needed for Discovered broken Retentate vessel waste valve Bracket | Juan.Bonilla@sanofi.com | 8NYA/1503 | 2024-07-02 | 2024-07-02 | Retentate vessel waste valve bracket not to be broken
3702 | Pompe - EGHUA005 - 8 NYA Downstream - UFDF skid 3431 went into hold with a retentate vessel weight low low alarm | adam.keister@sanofi.com | 8 NYA 1503 | 2024-07-02 | 2024-07-02 | To know why alarm happened, to clear alarm, and to progress with recipe. | Yes | Deviation
3701 | 8NYA Gemba-805007 batch not proceeding due to wfi acquiring issue | william.lee@sanofi.com | 8NYA/1350 | 2024-07-02 | 2024-07-02 | Recipe should have acquired the wfi drop and started delta V recipe.
3700 | Copper column regen/store not ran | zachary.rickard@sanofi.com | 8NYA/1503 | 2024-06-30 | 2024-07-02 | Upon completion of the copper production recipe a regen store recipe will be ran
3699 | Requesting Gemba for event QE-052053 for pompe batch record. | christian.galvez-tello@sanofi.com | 8NYA RM1360 | 2024-07-01 | 2024-07-01 | fbl-ewi-000011 Media Change Out
3698 | Requesting Gemba for event QE-052053 for pompe batch record. | christian.galvez-tello@sanofi.com | 8NYA RM1360 | 2024-07-01 | 2024-07-01 | fbl-ewi-000011 Media Change Out
3697 | Ethanol waste tank failure while decanting supernatant | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2024-07-01 | 2024-07-01 | Per QU-SOP-0057391, step 6.3.6 and QU-OPE-0434886 step 7.13.27, decanting should be happening through the side diptube to the ethanol waste line.
3696 | Pompe A280 Result out of Spec | brian.o'rourke@sanofi.com | 8NYA Room1403 | 2024-07-01 | 2024-07-01 | Batch Record requires a value of 4.90g/l or greater. Sample result was 0.58g/l |  | No Deviation
3695 | One non-viable unit not aligned to wall clock | valeria.moreno@sanofi.com | 76-NYA-355 | 2024-07-01 | 2024-07-01 | It is expected that all climets are set +/- 1 minute from the wall clock.
3694 | Vacuum pump broke down during fill operations | valeria.moreno@sanofi.com | 76-NYA 355 | 2024-07-01 | 2024-07-01 | Refer to QU-SOP-0074306, step 6.5.1 for HA fill sampling. Refer to QU-SOP-0006133, for SMA operations
3693 | gemba be held to determine there is a deviation needed to correct this situation and compliantly move this material to MTS. | jan.stepanik@sanofi.com | 74-164 | 2024-06-03 | 2024-06-28 | A procedure should be used to transfer the powder to MTS.
3692 | 2 Late PMs for Thyrogen | kevin.southwick@sanofi.com | 51 NYA | 2024-07-01 | 2024-07-01 | It is expected that work orders are completed by their due date per QU-SOP-0072346. | Yes | Deviation
3691 | Unable to locate HA-CMC from cold room at 74 NYA | jan.stepanik@sanofi.com | 74-164 | 2024-06-16 | 2024-06-24 | Should have located the powder
3690 | 8NYA LAL Sample 1238072 leaked and resubmitted on manual test request form | michael.aiello@sanofi.com | 4700 QCM | 2024-06-29 | 2024-06-30 | sample containers to pass visual check at receipt and resamples to have different id #'s from original
3689 | Condensate Observed in Vent line/filter of Pompe Harvest Collection Tanks PBM-2816001, 2916001 | elon.elzam@sanofi.com | Upstream suite 2, room 1403 | 2024-07-01 | 2024-07-01 | there should be no condensate in the vent line
3688 | Operator signed steps that he is not signed off on. 76 NYA Fermentation | dianne.moustafa@sanofi.com | 76 NYA/364 | 2024-06-26 | 2024-07-01 | It is expected the operator be signed off on QU-FOR-0045619 before completing transferring of ctab
3687 | EM Personal Monitoring Settle Plates and Rodac's moved on day 8 instead of day 7 | kathryn.lentoni@sanofi.com | 68NYA, 4700, QCM | 2024-06-29 | 2024-07-01 | For plates to be moved on days 5-7 for 20-25 incubation
3686 | A Gemba is needed to discuss the loading of the Dry Ice Shipper | Juan.Bonilla@sanofi.com | 8NYA/1503 | 2024-06-26 | 2024-06-26 | A safe path forward.
3685 | Seprafilm room 233-232-231 - | nicole.fragaalvarez@sanofi.com | room 231 | 2024-06-28 | 2024-06-29 | MABC-QU-SOP-0015815 step 6.3.1.
3684 | 8NYA Gemba-805013 exceeded mix timer | william.lee@sanofi.com | 8NYA/1350 | 2024-07-01 | 2024-07-01 | Mix timer should not have exceeded.
3683 | 45NYA Support APU-QCC45- Room 1136-FB-3465 UV Standard overdue for Calibration | robert.bauch@sanofi.com | 45-1136 | 2024-07-01 | 2024-07-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
3682 | 76NYA Biosurgery APU-Room 352-Dept HAFL-F-1310 Climet Particle Counter overdue for calibration | robert.bauch@sanofi.com | 76-352 | 2024-07-01 | 2024-07-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
3681 | SUT spill/failure on MCC tubing X2731 | brendan.dickhaut@sanofi.com | 8NYA 1210 CNC space, by viewing windows into 1403/skid X2731. | 2024-07-01 | 2024-07-01 | unused sterile connections are properly clamped off
3680 | Spill inside of Solution preparation | maria.palmar@sanofi.com | 8 New York Ave, room 1350. | 2024-06-30 | 2024-06-30 | We need to if we will be able to transfer our Media Buffer Pompe RF1YIE043 / EGHMD043 / 805013 and also our RF1N0E003 EGHBG003 P001 804998 .
3679 | Second acid addition prompt did not come up on for VI, mini eluate lost | ralston.najarro@sanofi.com | 8NYA1403 | 2024-06-30 | 2024-06-30 | Mini eluate should have gone to eluate SUM
3678 | Gemba-8NYA-Downstream Column Packing Sartorius Welder jammed during WFI tote changeout | aidan.mula@sanofi.com | 8RM1303 | 2024-06-30 | 2024-06-30 | Determination on forward processing of column packing activities as a result of PM and engineering assessment.
3677 | 8 NYA Downstream phase 2 Pompe UF\DF-B CIP Recipe Max stabilization timer expired for temp. | Christian.Venegas@Sanofi.com | 8NYA Cell Culture Phase 2 Room 1503 | 2024-06-29 | 2024-06-29 | To get a Phath forward
3676 | pH probe AI1 of VI skid 2825 primary mixer displaying input of secondary mixer ph probes leading to hold | brendan.dickhaut@sanofi.com | 8NYA Room 1403, US Ballroom | 2024-06-30 | 2024-06-30 | Correct probe input values are displayed on Delta V and locally.
3675 | Hold on MCC X-2811 train 4 | brendan.dickhaut@sanofi.com | 8NYA Room 1403, US Ballroom | 2024-06-29 | 2024-06-29 | Expected results: equipment is properly identified in the MFG recipe prior to launch, system properly loads columns and concentration is not too high, automation intervention is not needed.
3674 | Multiple 100L SUB Bags Are failing Inspection | kurt.marsocci@sanofi.com | 1333 | 2024-06-29 | 2024-06-29 | Order new bags, or find a way to use the exisiting ports.
3673 | 805010 Buffer was not completed within 10 hour time limit | christopher.arnold@sanofi.com | 8RM1350 | 2024-06-29 | 2024-06-29 | That the pressure sensor automation will not throw modbad alarms or hi/hi with no triggering pressure spike.
3672 | UF\DF-A CIP missing prompt to install nova septum bags | Christian.Venegas@Sanofi.com | 8NYA Cell Culture Phase 2 Room 1503 | 2024-06-28 | 2024-06-28 | To get a Phath forward | Yes | Deviation
3671 | Pompe - RF1NAE007 - 8NYA - VI Mixer Weight alarm callout during drain of eluate | michael.knapp@sanofi.com | 8NYA 1403 | 2024-06-28 | 2024-06-28 | Only critical alarm callouts are sent out.
3670 | 68 NYA - QCM -Controlled Copy Utilized for Analytical Testing Rather Than Controlled Print | joseph.napoli@sanofi.com | 68 NYA 4700 QCM | 2024-06-20 | 2024-06-27 | to control print all assay execution analytical paperwork
3669 | 8NYA - RF1N4E006 PH was out of range, residual material left in bottom of bag. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-06-26 | 2024-06-26 | Hold gemba to see path forward | Yes | No Deviation
3668 | 8NYA - 3k SUM bag failed pre inspection | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-06-28 | 2024-06-28 | N/A | No | Deviation
3667 | SUB 2220 Bleed Pressure MODBAD alarm | christopher.arnold@sanofi.com | 8RM1403 | 2024-06-28 | 2024-06-28 | That the pressure sensor automation will not throw modbad alarms or hi/hi with no triggering pressure spike. | Yes | Deviation
3666 | 8NYA- RF639E03 FAA Sample not taken per Per MABC-QU-WIN-0046060-Sampling & MABC-QU-OPE-0044689-Cz-ICB Sampling Plan for 8 NYA | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-06-25 | 2024-06-27 | It is expected that sampling occur according to Per MABC-QU-WIN-0046060-Sampling & MABC-QU-OPE-0044689-Cz-ICB Sampling Plan for 8 NYA | Yes | Deviation
3665 | Evidence of Tier 1 initiated and completed on 08APR2024 missing | Theresa.samanya@sanofi.com | 76 NYA/ R&D | 2024-04-08 | 2024-06-27 | Remove Tier 1 signs upon completion of work and return to QA. When the logsheet is completed will be reviewed by QA.
3664 | GEMBA - 8 NYA - Automation Intervention Needed for Column Clean Recipe | adam.keister@sanofi.com | 8 NYA. 1303 | 2024-06-27 | 2024-06-27 | Recipe brough back to allow for column priming
3663 | 8NYA-Fz-100L SUB Bag P/N 3202164 Failed Pre-Use Inspection x3 | john.lytle@sanofi.com | Rm 1333, 8NYA | 2024-06-27 | 2024-06-27 | During Inspections, the SUT should show no signs of damage/ defect and are suitable for use in cGMP processing
3662 | Automation intervention needed for UFA at 8 NYA | kenny.leland@sanofi.com | 1503 | 2024-06-27 | 2024-06-27 | Loop 4 needs to own all WFI droops |  | No Deviation
3661 | Unused GMBR/PVL not blocked in MES | peter.jansen@sanofi.com | MES | 2024-06-26 | 2024-06-26 | QU-SOP-0041372 step 6.5.6 | Yes | No Deviation
3660 | 805003 POU Sample already submitted i BioLIMS prior to POU sampling | robert.morawski@sanofi.com | 8NYA CNC | 2024-06-26 | 2024-06-27 | POU Endotoxin sample would not be signed for in BioLIMS. |  | No Deviation
3659 | 8NYA - 805009 / EGHBP006 Buffer In-process pH Result Out of Range | alejandro.martinez2@sanofi.com | 8NYA / RM1350 | 2024-06-26 | 2024-06-26 | pH value should have been within range 5.46-5.66. Gemba to discuss with MSAT the target range of Acetic Acid to charge so a more adequate amount is added during weighing/dispensing to achieve a median pH result
3658 | Gemba needed for 8 NYA SP column leak. | kenny.leland@sanofi.com | 1303 | 2024-06-26 | 2024-06-26 | No water should come out of the top inlet assembly.
3657 | Gemba - 8NYA - Pompe - UFDF CIP SUT taken out and was not used. | stanley.durand@sanofi.com | 8NYA ROOM 1503 | 2024-06-25 | 2024-06-26 | TBD
3656 | Bubble trap on qual skid goes into alarm for high high level. | kenny.leland@sanofi.com | 1303 | 2024-06-26 | 2024-06-26 | Bubble trap level should not reach the high high level.
3655 | Pompe Medai 805013 NRSR not delivered in time for bioburden testing | robert.morawski@sanofi.com | 8NYA CNC | 2024-06-26 | 2024-06-26 | Sample and manual test request for should have been put in QC sample bin and delivered with other daily samples.
3654 | A Gemba is required for failed shipping bags due to unacceptable welds. | Juan.Bonilla@sanofi.com | 8NYA/1703 | 2024-06-26 | 2024-06-26 | Less or no Striations on shipping bags
3653 | Spider found in 8NYA Womens Locker Room | stephanie.aceto@sanofi.com | 8NYA, Room 1126 | 2024-06-26 | 2024-06-26 | No pests would be observed in womens locker room
3652 | 8NYA Gemba For -20 Cold room alarm | william.lee@sanofi.com | 8NYA/1421 | 2024-06-24 | 2024-06-26 | Temperature shoudl remain at -20 degrees celsius
3651 | MLC identified side entrance to 74 cryo storage (RM 166) not secured. This area is badge access only. | michael.montiverdi@sanofi.com | 74NYA | 2024-06-25 | 2024-06-25 | Area to be secure
3650 | EWI is Missing Instructions for Shared Bleed Tote Setup | ana.salichs@sanofi.com | 8 NYA CNC | 2024-06-25 | 2024-06-25 | EWI would include correct instructions for shared bleed setup |  | No Deviation
3649 | 8NYA-Gemba-Downstream-Fz Column Packing, rouging found on column during vendor PM | aidan.mula@sanofi.com | 8RM1303 | 2024-06-25 | 2024-06-25 | Determination on forward processing of column packing activities as a result of PM and engineering assessment. | Yes | No Deviation
3648 | Warm room 148 at 51 NYA Low Temp | james.letourneux@sanofi.com | 51RM148 | 2024-06-25 | 2024-06-25 | Temperature should stay between 36-38 F.
3647 | Return Pumps not Returning CIP Fluids Back to CIP Skid CP600 | james.letourneux@sanofi.com | 51RM140 | 2024-06-24 | 2024-06-25 | Return pump should pump CIP fluids back to CIP skid.
3646 | Pompe - 8NYA - Cedex PI pull failure for Glucose out of range results during Pompe seed train | michael.knapp@sanofi.com | 8NYA 1403 | 2024-06-25 | 2024-06-25 | It is expected that the system can pull PI data into MES without issue.
3645 | RF1YIE038 - 805013 - small amount of liquid under SUM/SUT | leo.gillespie@sanofi.com | 1350 | 2024-06-24 | 2024-06-24 | no liquid should be evident under the SUM during or after formulation
3644 | ATF flow meter communication error causing delay in perfusion start | ralston.najarro@sanofi.com | 8NYA1403 | 2024-06-24 | 2024-06-24 | Prompt would pop up
3643 | Column 1 pH probe vibrating in CIP on Skid 2560 (pre-use CIP for CZ) due to missing O ring | phoebe.karkos@sanofi.com | 1403 in 8NYA | 2024-06-24 | 2024-06-24 | Re-standardized with new probe and restart CIP
3642 | Control Copy was used instead of Control Print for 76 NYA CTM Area Equipment Worksheet and Acceptance Criteria | kathryn.lentoni@sanofi.com | 68NYA, 4700, QCM | 2024-06-16 | 2024-06-24 | Expected results should be control print being used instead of control copy
3641 | Controlled Copy utilized for performing Assay execution instead of Controlled Print | joseph.napoli@sanofi.com | 68 NYA 4700 QCM | 2024-06-16 | 2024-06-24 | to control print all assay execution analytical paperwork
3640 | 8NYA - Incorrect labels generated for samples in biolims 805013 batch on 23JUN24 | bailey.gleason@sanofi.com | 1350 | 2024-06-23 | 2024-06-24 | N/A | Yes | Deviation
3639 | Tier 1 Action Notice initiated 20Dec2023 not returned to QA | Theresa.samanya@sanofi.com | 74 NYA/ 1017 | 2023-12-22 | 2024-06-24 | Remove Tier 1 signs upon completion of work and return to QA. When the logsheet is completed will be reviewed by QA.
3638 | Puncture found on the SWECO screen mesh after product collection | munish.kumar2@sanofi.com | 76 NYA/Room#3103/Purification | 2024-06-24 | 2024-06-24 | Per QU-SOP-0008047, the screen was inspected for any dents, creases, or tears and a puncture was notices. Per step 6.7.4, this is an abnormality found and supervisor was notified.
3637 | 8 NYA GEMBA - Skid setup not open before launching Recipe | alex.tello@sanofi.com | 1503 | 2024-06-24 | 2024-06-24 | For Setup step to be available in MES before launching recipe | Yes | No Deviation
3636 | 8NYA-Pompe 807198- Prompt not acknowledged during pH adjustment of VI skid. Eluate discarded to drain | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-06-22 | 2024-06-23 | It is expected that the DeltaV prompt to continue acid additions is acknowledged appropriately per MABC-QU-WIN-0025265-Viral Inactivation Operations
3635 | 8NYA-Cz-X2220- Intervention requested to disable high side pH control | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-06-22 | 2024-06-23 | It is expected that pH control is adequate to meet processing requirements per CST
3634 | 8NYA - Mice Found in Break Area Room 1120 | wilfredo.lopez@sanofi.com | 8RM1120 | 2024-06-21 | 2024-06-21 | Expected that area remains rodent-free
3633 | 8NYA - EGHWG, 706210 WFI Batch SUM Bag Inflation Setup Step Not Signed | wilfredo.lopez@sanofi.com | 8RM1350 | 2024-06-22 | 2024-06-23 | It is expected that all batch record steps are executed and signed contemporaneously.
3632 | 8NYA - EGHAV047, 760741 Media Leak Observed on Filter Inlet During Transfer Operation | wilfredo.lopez@sanofi.com | 8RM1350 | 2024-06-23 | 2024-06-23 | Expected that all connections remain integral and no leaks observed
3631 | 8NYA - EGHAV047, 760741 Media Reagent Setup MES Steps and EWI Not Aligned | wilfredo.lopez@sanofi.com | 8RM1350 | 2024-06-22 | 2024-06-22 | Operators expect the proper flow of MES steps, correct EWI references contained to accurately execute and document the process. | Yes | No Deviation
3630 | Gemba-8NYA - Shipping bag failed visual inspection not recorded on the shipping bag protocol. | stanley.durand@sanofi.com | TBD | TBD | 2024-06-21 | All bags going through visual inspection should be recorded on the Protocol For Shipping Bag Visual Inspection at 8NYA QU-OPE-0492708.
3629 | 8NYA-Pompe- ATF x2 failed visual Inspection during setup for production activities | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-06-22 | 2024-06-22 | It is expected that ATF 10s be inspected for possible damage/ defects prior to use in production activities. Inspection occurs according to MABC-QU-WIN-0046472-Single Use Bioreactor (SUB), ATF10 SetUp
3628 | Gemba – 8NYA –Create stock issue for 804995 buffer | william.lee@sanofi.com | 8NYA/1350 | 2024-06-22 | 2024-06-22 | Stock creation should be 50L
3627 | 8NYA - EGHBF027 - 705665 Filter Tested with Incorrect Recipe | wilfredo.lopez@sanofi.com | 8RM1350 | 2024-06-22 | 2024-06-22 | It is expected that the correct filter testing recipe launches through MES to DeltaV based on the CX values pulled in from the parent MO. After the selecting wetting not required, the record should not change the testing parameters.
3626 | Seprafilm - EBYSEP034 - 76 NYA - Room 237 - Formatting error in batch record | karthik.sriram@sanofi.com | 76 NYA - 237 | 2024-06-06 | 2024-06-21 | Batch record should meet all formatting requirements per record management procedures
3625 | Disabling Aber probe AI2 on SUB X-2220 | elon.elzam@sanofi.com | Upstream suite 2 room 1403 | 2024-06-21 | 2024-06-21 | Capacitance probe AI2 for X-2220 to be disabled by manufacturing per FBL-EWI-000010
3624 | QC Micro - 68 NYA - Cerezyme - No Reportable Result for LAL | Craig.stevenson@sanofi.com | 68NYA Room 4700 | 2024-06-20 | 2024-06-20 | Its expected that QCM is obtain reportable results.
3623 | 51NYA: equipment prep. Multiple errors across several logbooks | adam.rowe@sanofi.com | 51NYA, 129 | 2023-10-15 | 2024-06-20 | Steps should be verified in real time. And the racks used should match the SOP
3622 | No Reportable Result for LAL Sample | joseph.napoli@sanofi.com | 68NYA QCM | 2024-06-20 | 2024-06-20 | reportable result for LAL sample
3621 | No Reportable Result for LAL Sample | joseph.napoli@sanofi.com | 68NYA QCM | 2024-06-20 | 2024-06-20 | reportable results for LAL Sample
3620 | No Reportable Result for LAL Sample | joseph.napoli@sanofi.com | 68NYA QCM | 2024-06-20 | 2024-06-20 | reportable results for LAL Sample
3619 | Warm room 148 in alarm for high and low side of range. | adam.hildebrand@sanofi.com | 51-148 | 2024-06-20 | 2024-06-20 | It is expected the warm room will keep temp within the validated range
3618 | Re Enabling pCO2 diverging alarm SEL_DEV on SUB X-2240 | elon.elzam@sanofi.com | Upstream suite 2, room 1403 | 2024-06-20 | 2024-06-20 | pCO2 Diverging alarm to be re instated to monitor if probes
3617 | QC Micro - 51/8NYA - EM sampler not qualified to perform routine EM | Craig.stevenson@sanofi.com | 51 NYA and 8 NYA | 2024-06-17 | 2024-06-18 | It is expected that all EM sampling analysts are qualified to perform EM
3616 | 8NYA - GEMBA CZ Thaw Filter FIT launched to the incorrect FIT unit. | alex.tello@sanofi.com | 1290 | 2024-06-15 | 2024-06-15 | for FIT to launch in room 1290 on unit 2009075 where EBR says it was launched | Yes | Deviation
3615 | 8NYA - EGHBJ016 805003 missing samples in Biolims for in process and POU | bailey.gleason@sanofi.com | CNC | 2024-06-19 | 2024-06-19 | Hold gemba to see path forward
3614 | Temp controller not working on SUB X-2220 | brendan.dickhaut@sanofi.com | 8NYA Room 1403, US Ballroom | 2024-06-19 | 2024-06-19 | Temp controller will control temperature setpoints of the reactor by either cooling or warming the reactor jacket.
3613 | ATF10 found to have particles | ralston.najarro@sanofi.com | 8NYA1403 | 2024-06-19 | 2024-06-19 | Missing clarity on whether the particles are permitted
3612 | 8NYA - 805003 POU samples Auto generated in Biolims. Not required per MABC-QU-OPE-0527081-Pompe-ICB Sampling Plan for 8 NYA | bailey.gleason@sanofi.com | 8NYA | 2024-06-19 | 2024-06-19 | N/A
3611 | QC Micro - 51NYA - No reportable Non-Viable result for Weekly EM site 152-3 | Craig.stevenson@sanofi.com | 51 NYA Room 152 | 2024-06-03 | 2024-06-19 | It is expected that non-viable results be collected for site 152-3 at 51NYA.
3610 | Cerezyme - RF639E003 - 8NYA - TCU SP not aligning with PV | michael.knapp@sanofi.com | 8NYA 1403 | 2024-06-19 | 2024-06-19 | Proper control of temp through TCU
3609 | 8NYA - EGHBU006 - 807194 Incorrect Batch Material Stock Quantity Created | wilfredo.lopez@sanofi.com | 8RM1350 | 2024-06-19 | 2024-06-19 | Expected that the MES provides the correct steps required for the operation including references to the correct EWI procedure. Expected that the correct output quantity is achieved for batch material.
3608 | 8NYA - RF1N8E006 807194 100L batch created stock for 20L | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-06-19 | 2024-06-19 | N/A
3607 | Particulate observed on the wall of down flow booth DFB10100001 | thomas.loik@sanofi.com | N/A | 2024-06-12 | 2024-06-12 | Wall is cleaned with IPA wipes post use.
3606 | 8NYA GEMBA Request- Multiple samples found in cold room | richard.manasian@sanofi.com | 1420 | 2023-08-21 | 2024-03-12 | Samples should have been discarded upon completion of investigation. NRSR was not generated
3605 | 8NYA - 705661 - Compromised Seal on Fill Line of 0.1M NaOH Tote | wilfredo.lopez@sanofi.com | 8RM1212 | 2024-06-17 | 2024-06-17 | Expected that all seals remain integral
3604 | HOF Cleaning Procedure | zachary.rickard@sanofi.com | 8NYA CNC | 2024-06-14 | 2024-06-14 | Cleaning procedure will be located in the correct EWI or be correctly referenced when needed.
3603 | 8NYA-Pompe RFiNAE007 X2831 Potential for Column overload/ Automation Intervention during live GEMBA troubleshooting event | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-06-17 | 2024-06-17 | It is expected that the maximum load volume for each column is maintained against the requirements of
3602 | 51NYA Cell culture correction to iLims lot number | adriana.alban@sanofi.com | 51NYA/140 | 2024-03-23 | 2024-06-07 | It is expected that the lot number is accurate on iLIMS labels/test requests.
3601 | Gemba needed for failed shipping barrel inspection. | kenny.leland@sanofi.com | 8 NYA room 1703 | 2024-06-07 | 2024-06-07 | No burrs should be found.
3600 | 8NYA- Pompe-RF1NAE008 DO Probe 1 malfunction | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-06-17 | 2024-06-17 | It is expected that both probes remain functional and the values are averaged to determine the pF of the product
3599 | 8NYA- Pompe- RF1NAE007 Pompe eluate SUMS negative weight values preventing recipe from filling eluate SUMs | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-06-15 | 2024-06-15 | It is expected that the load cells are tared and reading a positive value or Zero
3598 | 8NYA-Cz-X2230 DO Probe Divergence placing recipe in hold | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-06-17 | 2024-06-17 | The average of both probes is used to control SUB DO, and should remain within acceptable range of each other
3597 | 807194 - RF1N8E005 - Low Condo result | william.lee@sanofi.com | 1350 | 2024-06-16 | 2024-06-17 | no specific instructions about how to proceed with record
3596 | 8NYA - RF1NIE007 transfer line is stuck in welder #2009967 part of welder looks broken off | bailey.gleason@sanofi.com | CNC | 2024-06-16 | 2024-06-16 | N/A
3595 | 8NYA - Sample testing results not showing in PI historian from 10JUN24 - 16JUN24 for SP operations | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-06-16 | 2024-06-16 | N/A
3594 | Rzecipe went into hold during qual skid cip due to hi hi level sensor on the bubble trap | pavel.dobrovolskiy@sanofi.com | 8NYA 1303 | 2024-06-15 | 2024-06-15 | Its expected that the recipe maintain the appropriate flow rate and not cause the level in the bubble trap to exceed hi hi
3593 | 705691 - RF573E012 - MES record - Tote transfer set up prior to all 3 tote SUT | bailey.gleason@sanofi.com | 1350 - Solution prep | 2024-06-15 | 2024-06-15 | Record should execute without mixing instalation steps
3592 | 8NYA-Gemba-Downstream Cz- HOF does not have cleaning procedure | aidan.mula@sanofi.com | 8RM1210 | 2024-06-14 | 2024-06-14 | Determination at Gemba, uncertain if cleaning procedure should be required
3591 | 8NYA - RF551E015 Incorrect Condo Sample value pulled from PI and was not corrected by operators | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-06-15 | 2024-06-15 | hold gemba to see path forward | No | No Deviation
3590 | 8NYA - Media batches performed night of 14-15JUN24 connections performed without QA present | bailey.gleason@sanofi.com | 8NYA | 2024-06-15 | 2024-06-15 | Hold gemba to see path forward | Yes | No Deviation
3589 | AHU 7 at 51 shut off | Alex.Biskaduros@Sanofi.com | 51 NYA room 149 to 84 and additional rooms and rooms 152 ton 149 | 2024-06-14 | 2024-06-14 | AHU 7 should not have shutdown unexpectedly | Yes | Deviation
3588 | To perform base prime pump cap remove | bhargav.patel@sanofi.com | 8 NYA / 1403 | 2024-04-04 | 2024-04-04 | Automation adjustment should not require to adjust pump speed
3587 | 8NYA - Cerezyme EGHBN002 - 760741 Active Media Tote Leaking at Fill Transfer Line Seal | wilfredo.lopez@sanofi.com | 8RM1212 | 2024-06-14 | 2024-06-14 | Expected that all seals remain integral |  | Deviation
3586 | 76NYA HA - Multiple Instances (6ish) of non viable units not aligned to wall clock | Craig.stevenson@sanofi.com | 76NYA, Room 355 | 2024-05-08 | 2024-06-13 | It is expected that all climets are set +/- 1 minute from the wall clock.
3585 | SUB X-2250 Media Charge Recipe Was Not Progressing | brian.o'rourke@sanofi.com | 8NYA Room1403 | 2024-06-13 | 2024-06-13 | QU-WIN-0046472 Media Setup 500L SUB - Task 2 |  | No Deviation
3584 | 8NYA-Cz RF639E002 Unable to restart SUB recipe following MODBAD on Bleed Line Pendotech sensor | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-06-13 | 2024-06-13 | It is expected that operators respond to alarms and are able to restart recipe based on criteria set forth in EWI-000004 Alarm Response
3583 | No Instructions for C-Flex to AdvantaFlex welding | elon.elzam@sanofi.com | Upstream suite 2 room 1403 | 2024-06-08 | 2024-06-07 | FBL-EWI-000031 Child Sartorius Biowelder Operations, Task 2 | No | Deviation
3582 | 8NYA Pompe Unaccounted POU Pompe Media 805012 EGHDS011 inadvertantly sampled under 50L wave EOR sample | judith.keller@sanofi.com | 8NYA/Upstream | 2024-06-11 | 2024-06-12 | It is expected that samples are sampled and labeled correctly when submitting to QC for testing.
3581 | Pompe, RF1NAE008, Previous Base setup record consuming SU when still needed for shared install | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-06-13 | 2024-06-13 | Previous record to only consume partial amount of SU so that the next record can acquire same SU for install
3580 | 45NYA Support APU-Room 1131- chart recorder FB-1242-002 found out of tolerance during an unscheduled final calibration before being removed from service. | robert.bauch@sanofi.com | 45-1131 | 2024-06-13 | 2024-06-13 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3579 | VI Skid X-2865 Failed to reach Acid Set Point | elon.elzam@sanofi.com | Upstream suite 2, room 1403 | 2024-06-13 | 835 | FBL-CST-000031, initial pH needs to be between 3.50-3.70 | Yes | No Deviation
3578 | Lot EBYSEP017B 768281(Unlabeled p/n) Qty received by L&P does not match TO or inventory | alan.rousseau@sanofi.com | room 1017 | 2024-06-12 | 2024-06-12 | Expected result is that SAP inventory would be correct and match qty received
3577 | 8NYA - Cerezyme EGHSL006, 832327 - Insufficient EWI Instructions for Filter Testing Operation | wilfredo.lopez@sanofi.com | 8RM1350 | 2024-06-13 | 2024-06-13 | It is expected that batch records and EWIs contain the necessary steps referencing proper filter testing procedures. References to WFI wetting requirements must be included in the EWIs as needed.
3576 | 8NYA - Cerezyme EGHSL006 - Missing MES Step for Hexylene Glycol Manifold Addition | wilfredo.lopez@sanofi.com | 8RM1210 | 2024-06-12 | 2024-06-12 | It is expected that batch records contain the necessary steps referencing EWI procedures to be performed. The record needs to be updated to include the missing step as was the case for batch records used in manufacturing of part numbers 768506 and 805006
3575 | Unused harvest bag defect | david.loik@sanofi.com | 51 NYA | 2024-06-12 | 2024-06-12 | Harvest bag would be assembled correctly by the manufacturer.
3574 | 8NYA - Gemba for seal leak from 805013 media batch | william.lee@sanofi.com | 8NYA/1350 | 2024-06-12 | 2024-06-12 | Seal should not have leaked
3573 | Bulk Intermediate Product Reconciliation Log Missing Entry | craig.valery@sanofi.com | 74 NYA | 2024-04-24 | 2024-06-12 | it is expected all BIRF entries are filed in
3572 | 8NYA - Incorrect cleaning on material in MAL/PAL | Craig.stevenson@sanofi.com | 8NYA MAL/PAL | 2024-06-12 | 2024-06-12 | It is expected that everyone entering a classified space, clean equipment correctly prior to entering.
3571 | 45 NYA Main Air Compressors- No PM established PM schedule | wayne.fecteau@sanofi.com | 45 NYA Penthouse | 2024-06-12 | 2024-06-12 | QU-SOP-0036445- ASSET MANAGEMENT, Step 6.1.3. Equipment was placed into service in the field without PM strategy assigned and equipment activation in EAM | Yes | Deviation
3570 | Disposal record for upstream media went red when doing a take-over. | christian.galvez-tello@sanofi.com | 8NYA RM1360 | 2024-06-12 | 2024-06-03 | N/A
3569 | Freezer F-15831 went outside of validated temperature of -70 to -90C | lauren.ruffing@sanofi.com | 68NYA 4740 QC Micro | 2024-06-12 | 2024-06-12 | Freezer F-15831 stays inside of validated temperature of -70 to -90C per QU-SOP-0021304
3568 | APS 2R Vial Media Fill, 1 vial not accounted for upon receipt into QC (10kvial fill) | kathryn.lentoni@sanofi.com | 68NYA, 4700, QCM | 2024-06-11 | 2024-06-11 | Expected results are all 9820 vials are accounted for during turbidity reads
3567 | CZ, RF639E002. Harvest Wye not in the pinch clamp | joseph.gonzalez@sanofi.com | 1403 | 2024-06-12 | 2024-06-12 | Automation to intervene and open pinch clamp.
3566 | CZ, RF639E002. Harvest Wye not in the pinch clamp | gon2934@gmail.com | 1403 | 2024-06-12 | 2024-06-12 | Automation to intervene and open pinch clamp.
3565 | CZ, RF639E002. Harvest Wye not in the pinch clamp | gon2934@gmail.com | 1403 | 2024-06-12 | 2024-06-12 | Automation to intervene and open pinch clamp.
3564 | CZ, RF639E002. Harvest Wye not in the pinch clamp | gon2934@gmail.com | 1403 | 2024-06-12 | 2024-06-12 | Automation to intervene and open pinch clamp.
3563 | Nano Equil/Wetting cannot be run concurrently. | Dang.tran@sanofi.com | 8NYA, RM1503 | 2024-06-11 | 2024-06-11 | The nano-form and nano-filter-wet recipe uses a common phase that has hard coded logic. The expectation is that the equipment used should only look at conditions specific to the skid that is running. However, the code is not setup that way and could potentially lead to problems in the future. |  | No Deviation
3562 | Vacuum fuse pump broke down for SMA Atrium during Fill operations | valeria.moreno@sanofi.com | 76 NYA- room 355 | 2024-06-11 | 2024-06-11 | Refer to QU-SOP-0074306, step 6.5.1 for HA fill sampling. Refer to QU-SOP-0006133, for SMA operations
3561 | Harvest Overflow pump tripping pressure sensor | ralston.najarro@sanofi.com | 8NYA1403 | 2024-06-10 | 2024-06-10 | Pump should be able to divert harvest to waste without pressure sensor tripping
3560 | Gas flow alarm on X2162 / Keep using media | ryan.devanna@sanofi.com | 1403 8 New York Ave | 2024-06-10 | 2024-06-10 | n/a
3559 | leaking seal on the 200mM Acetic Acid part # 805004 | dennis.bazinet@sanofi.com | 8NYA CNC area | 2024-06-10 | 2024-06-10 | Sealed lines are not supposed to leak or drip after the line is sealed
3558 | test | sabrina.waring@sanofi.com | test | 2024-06-10 | 2024-06-10 | test
3557 | Qualified Observer field missing signature on Aseptic Qual paperwork | ryan.devanna@sanofi.com | 8 NYA, innoc prep | 2024-04-20 | 2024-06-10 | Management sign in place of operator
3556 | Qualified Observer field missing signature on Aseptic Qual paperwork | rdevanna1115@gmail.com | 8 NYA, innoc prep | 2024-04-20 | 2024-06-10 | Management sign in place of operator
3555 | 8NYA - QE Number Needed to Reference in Batch Records for MES Steps Performed Out of Sequence | wilfredo.lopez@sanofi.com | 8RM1350, 8RM1210 | 2024-06-07 | 2024-06-07 | Expected that all steps performed per MES records are signed when work is complete. Any steps not performed in sequence must be reflected in the records to ensure ALCOA principles are observed. | Yes | No Deviation
3554 | 8NYA - QE Number Needed to Reference in Batch Records for EWI Steps Not Performed | wilfredo.lopez@sanofi.com | 8RM1350, 8RM1210 | 2024-06-07 | 2024-06-07 | Expected that all steps performed per the EWIs are signed for in MES. Any steps not performed must be reflected in the records to ensure ALCOA principles are observed. | Yes | No Deviation
3553 | 8NYA - EGHBB014, 705662 Refractive Index Testing Performed with Incorrect System Setting per EWI | wilfredo.lopez@sanofi.com | 8RM1350 | 2024-06-09 | 2024-06-09 | Expected that the system be in the correct setting. Also expected the ability to change the settings as stated in the instructions.
3552 | 8NYA - EGHBB014, 705662 Automation Intervention Required During Mixing Timer Due to SUM Weight Interlock Condition | wilfredo.lopez@sanofi.com | 8RM1350 | 2024-06-09 | 2024-06-09 | Expected that the agitator stays on during the procedure to ensure proper mixing of material. | Yes | Deviation
3551 | Pompe, RF1NAE007, No instructions for harvest vessel overlay step | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-06-09 | 2024-06-09 | instructions on setting up harvest vessel overlay required per MABC-QU-OPE-0522556 | Yes | Deviation
3550 | Media bag leaking in barrel | david.loik@sanofi.com | 51 NYA | 2024-06-09 | 2024-06-09 | Media bag would not have leaked.
3549 | 8NYA-Gemba Outlet tubing leak (Caustic Buffer) | tanner.raiburn@sanofi.com | 8NYA-CNC | 2024-06-08 | 2024-06-08 | Find a path forward
3548 | 8NYA- Cerezyme- RF639E002 DO LoLo alarm placed SUB X2230 in hold. Per Alarm Response, GEMBA is required | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-06-08 | 2024-06-08 | It is expected that SUB maintain DO parameters set forth by CST. It is expected that Aber pulses and probe recalibration happen at seperate times
3547 | 8NYA - Gemba for autoclave over heating during VI probe load | william.lee@sanofi.com | 8NYA/1350 | 2024-06-07 | 2024-06-07 | Autoclave should have completed recipe with no excursions
3546 | POMPE, RF1NAE007, Failed weld on bleed line | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-06-07 | 2024-06-07 | weld to pass visual inspection
3545 | GEMBA - 8 NYA - Shaker Timer Frozen in MES | adam.keister@sanofi.com | 8 NYA | 2024-06-07 | 2024-06-07 | For record to advance as intended
3544 | Work order not completed contemporaneously | kevin.southwick@sanofi.com | 51 NYA | 2024-05-13 | 2024-06-06 | It is expected that work orders are processed contemporaneously per QU-SOP-0072346. | Yes | No Deviation
3543 | 8NYA - Gemba for 3K SUM bag leak | william.lee@sanofi.com | 8NYA/1350 | 2024-06-07 | 2024-06-07 | SUM bag should not have leaked
3542 | 8NYA APU-Room 1333-Dept CC8-Temperature indicator FB-2771-034 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1333 | 2024-06-03 | 2024-06-07 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3541 | Operations not signed for at the time of execution | brendan.dickhaut@sanofi.com | 8NYA Room 1403, US Ballroom | 2024-06-06 | 2024-06-06 | Good Documentation Practices, MABC-QU-SOP-0074684, 6.1.2 - all entries are to be made sequentially at the time of execution.
3540 | 8NYA - RF1YIE028 Bioburden Sample Leaked prior to transport | bailey.gleason@sanofi.com | 8NYA | 2024-06-06 | 2024-06-06 | Hold gemba to see path forward
3539 | RQ NaOH filter tested as bubble point per QU-SOP-0046034 attachment 1, not QU-OPE-0309296 STEP 7.11.22 | dianne.moustafa@sanofi.com | 76 NYA/364 | 2024-06-04 | 2024-06-06 | When is process use OPE
3538 | Ant found in 76-205 | Craig.stevenson@sanofi.com | 76NYA, Room 205 | 2024-06-06 | 2024-06-06 | It is expected that no pests are found in controlled areas
3537 | Pompe Antifoam line not set up to dedicated Antifoam port | Bhargav.patel@sanofi.com | 8 NYA / 1403 | 2024-06-03 | 2024-06-04 | FBL-EWI-000011
3536 | 8NYA GEMBA Upstream Aseptiquik G partially Torn Within Adhesive Seal Mark Cerezyme SUB Lot EGHBN002 PN: 761217 | gary.dodakian3@sanofi.com | 8NYA 1210 | 2024-06-06 | 2024-06-06 | Aseptiquik is integral
3535 | CZ Thaw sampling ViCell ran for FZ instead of CZ | phoebe.karkos@sanofi.com | 8NYA Innoc prep | 2024-06-05 | 2024-06-06 | document event
3534 | Unable to perform verification step in MES | Jonathan.Ahiaba@sanofi.com | 8NYA, 1403 | 2024-06-06 | 2024-06-06 | verification step is able to be signed by a train operator with operator access level in MES
3533 | Cold room 68-4755 dropped below 2 degree C | justin.carpenter@sanofi.com | 68-4755 | 2024-06-05 | 2024-06-05 | Cold room stays between 2-8 degree C
3532 | CV swab sample cannot be documented contemporaneously in EBR. Post CIP operations must be performed and verified before the operator can sign for CV swab sampling. | brendan.dickhaut@sanofi.com | 8NYA Room 1403, US Ballroom | 2024-06-06 | 2024-06-06 | Good Documentation Practices, MABC-QU-SOP-0074684, 6.1.2 - all entries are to be made sequentially at the time of execution
3531 | SUT Failed on WFI Tote Fill | kurt.marsocci@sanofi.com | 8 NYA CNC Area | 2024-06-05 | 2024-06-05 | Replace line and continue filling | No
3530 | CEDEX LDH2B calibration did not pass (do not have another Cal B to re calibrate) | phoebe.karkos@sanofi.com | 8NYA 1403 Upstream P2 | 2024-06-05 | 2024-06-04 | 2 Options: skip only the test that is not working or skip cedex entirely for daily sampling
3529 | No Water Bath/Thermometer Information Recorded During Growth Promotion Testing of Lot 24FN01195 For 2X TSA Media | Rebecca.Schmidt@sanofi.com | 68NYA/4700/QCM | 2024-06-04 | 2024-06-05 | Water bath and thermometer information needs to be recorded when used. QU-SOP-0059766 step 6.4.4 requires the media to be melted and cooled in two separate water baths.
3528 | HA-CMC Batch EBYCMC098B (part 2139-32) has exceeded its maximum mixing time per QU-OPE-0520874 step 7.4.7 | jan.stepanik@sanofi.com | 76 NYA / room 229 | 2024-06-05 | 2024-06-05 | Total mix time 12 - 24 hrs.
3527 | Rupture Disc blew out of the Tank V-3521 during transfer of HA Slurry | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2024-06-04 | 2024-06-04 | The rupture disc should not blow out during this process because the disc is used for safety purposes.
3526 | 8NYA- Gemba- Solution Transfer Tote Setup (CNC) EWI discrepancy | alejandro.martinez2@sanofi.com | 8NYA/CNC-Solution Fill Area | 2024-06-04 | 2024-06-04 | ITSCR to correct the observation
3525 | Pompe Media tote BHT2042005 not reading properly | robert.morawski@sanofi.com | 8NYA CNC | 2024-06-04 | 2024-06-04 | Tote weigh screen showed correct tote volume.
3524 | 8NYA- Support APU-Facilities-TOC FB-4002-005 with no as found data during an unscheduled calibration | robert.bauch@sanofi.com | 8-1170 | 2024-05-29 | 2024-06-04 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | No Deviation
3523 | Pompe, RF1NAE007, Antifoam instructions not clear. Antifoam attached to wrong line on SUB | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-06-03 | 2024-06-04 | instructions should be clear to install antifoam line onto SUB line with dip tube
3522 | SP1 pH Out of Range for Multiple Steps, EGF4B05A, B, and C | alyson.belleville@sanofi.com | 51 NYA, 138 | 2024-05-21 | 2024-06-03 | It is expected that the pH would remain within range throughout the operation.
3521 | Cerezyme - RF639E002 - 8NYA - pH Control Recipe Adjustment Needed | michael.knapp@sanofi.com | 8NYA 1403 | 2024-06-04 | 2024-06-04 | Low side pH control at this stage of process
3520 | 76-AHU 2A fan failure | john.beaule@sanofi.com | 76-224, 227 and 229 | 2024-06-03 | 2024-06-03 | FBS-SOP-000747, area pressures
3519 | 8NYA GEMBA - Submit results section in EWI for BDT | jack.laplante@sanofi.com | 8NYA 1373 | 2024-06-04 | 2024-06-04 | Looking for path forward.
3518 | 8NYA GEMBA - 832327 Buffer Manifold/5ft extension SUT not in BOM but is on EWI | jack.laplante@sanofi.com | 8NYA CNC | 2024-06-03 | 2024-06-03 | Looking for path foward.
3517 | 8RM 1333 Cedex unable to | jacob.roman@sanofi.com | 8RM1333 | 2024-06-02 | 2024-06-02 | The cedex is expected to pass cal
3516 | Cerezyme - RF639E002 - 8NYA - pH Probe Calibration Formula Result Error | michael.knapp@sanofi.com | 8NYA 1403 | 2024-05-31 | 2024-06-03 | Correct decision for probe calibrations
3515 | Pompe, RF1NAE007, 8NYA, Guard filter bleed valve droplet discovered | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-06-03 | 2024-06-03 | No droplet should be present outside bleed valve
3514 | No Product - No Lots - 76 NYA - QC Micro - No Reportable TOC/Conductivity results for site 222 | craig.stevenson@sanofi.com | 76NYA Room 107 / 68 NYA Room 4700 | 2024-03-09 | 2024-06-03 | It is expected that site 222 to have TOC and Conductivity sampled on a weekly frequency.
3513 | 8NYA Gemba - seal on 832327 (caustic buffer) | jack.laplante@sanofi.com | 8NYA CNC | 2024-06-03 | 2024-06-03 | Looking for path foward.
3512 | Lost hairnet while de-gowning from solution prep | christian.galvez-tello@sanofi.com | 8NYA RM1360 | 2024-06-03 | 2024-06-03 | N/A | Yes | Deviation
3511 | 51NYA Thyrogen APU-Room 129- FB-0578 Pressure Gauge overdue for calibration. | robert.bauch@sanofi.com | 51-129 | 2024-06-01 | 2024-06-03 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
3510 | 76NYA-Biosurgery APU-Room 3103-HABP-Pressure Indicator Transmitter F-6938-036 overdue for calibration. | robert.bauch@sanofi.com | 76-3103 | 2024-06-01 | 2024-06-03 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
3509 | 76NYA-Support APU-Room 107-QCBP-Viscometers F-15970 and F-15971 are overdue for calibration | robert.bauch@sanofi.com | 76-107 | 2024-06-01 | 2024-06-03 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
3508 | ATF failed inspection | Ralston.najarro@sanofi.com | 8NYA1403 | 2024-06-02 | 2024-06-02 | No holes in ATF | Yes | Deviation
3507 | Incorrect EWI instructions for ATF connection | Ralston.najarro@sanofi.com | 8NYA1403 | 2024-06-02 | 2024-06-02 | EWI should have clear and concise instructions | Yes | No Deviation
3506 | 8NYA - GEMBA CZ MMC Column failed 2nd qualification | alex.tello@sanofi.com | 1303 | 2024-06-01 | 2024-05-31 | Qual results to be within range
3505 | X2230 pressure sensor not reading correctly | Ralston.najarro@sanofi.com | 8NYA1403 | 2024-05-30 | 2024-05-31 | Pressure sensor should be reading correct pressure
3504 | 8NYA- EBR needs update to correlate with FBL-EWI-000015 instructions for step 21.9.7 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-05-31 | 2024-06-01 | N/A
3503 | 8NYA - Instruction in FBL-EWI-000015 step 12.4.3 has a safety concern | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-06-01 | 2024-06-01 | Hold gemba to see path forward
3502 | 8NYA - No instructions to remove clamps from WFI and Glycol ports in FBL-EWI-000015 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-06-01 | 2024-06-01 | N/A
3501 | MSAT contractor entered the controlled area 186E through restricted emergency exit corridor 186F at 45 NYA | Jeremy.Wooden@Sanofi.com | 186E was accessed via 186F at 45 NYA | 2024-05-31 | 2024-05-31 | It is expected that contractors would be escorted by trained personnel and not access the controlled areas through restricted areas.
3500 | 8NYA GEMBA - CZ MMC Column failed Qualification after packing | alex.tello@sanofi.com | 1303 | 2024-05-31 | 2024-05-31 | Qual results to be within range
3499 | X2240 temperature not coming into range | Ralston.najarro@sanofi.com | 8NYA1403 | 2024-05-31 | 2024-05-31 | Temperature must be at 37.00 degrees
3498 | ATF10 failed visual inspection | Ralston.najarro@sanofi.com | 8NYA1403 | 2024-05-31 | 2024-05-31 | ATF10 should be free of defects | Yes | Deviation
3497 | CE SDS incubation temperature below acceptable range | aimee.park@sanofi.com | 68NYA, QC Chemistry 2nd floor | 2024-05-28 | 2024-05-31 | Both temperatures would be within the acceptable range during incubation
3496 | 8NYA-Gemba-Downstream-Pompe Nano- 2 issues with nanofiltration record for shipment | aidan.mula@sanofi.com | 8RM1503 | 2024-05-31 | 2024-05-31 | Record should have correct SU set, additionally, material that was taken out against the initial order should be used in that order.
3495 | 8NYA - FBL-EWI-000019 tote fill execution needs to be split into 2 seconds 1 for CNC and 1 for in Solution prep | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-05-31 | 2024-05-31 | N/A
3494 | 68NYA Support APU-Validation-F-14727 Temperature Monitor with probe A found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | used in Northborough | 2024-05-29 | 2024-05-31 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3493 | 8NYA - RF1NIE006 and RF1YIE027 steps not executed per FBL-EWI-000015 for inflation set up | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-05-31 | 2024-05-31 | receive deviation, connect lines, move forward.
3492 | 8NYA-Cz-RF639E002 SUB Inoc Timeline shift | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-05-30 | 2024-05-30 | SUB Inoculation after DO Control | Yes | No Deviation
3491 | No Product - No Lot - 68NYA - QC Micro - 2 instances of Controlled Print bot being used | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2024-05-16 | 2024-05-29 | It is expected that any analtycial worksheets containing original assay data be Controlled Printed.
3490 | EWI instruction for PH control clarification | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-05-30 | 2024-05-30 | instructions on seeing prompt before sampling so that PH sample is correct value
3489 | Media Charge SUB weight delay | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-05-24 | 2024-05-24 | SUB weight to increase to show an actuate value during media charge
3488 | Concerns Around SCN Process Robustness - | benjamin.Larson@sanofi.com | N/A | 2024-05-29 | 2024-05-29 | A determination as to if a deviation has occurred.
3487 | 8NYA-Multiple Products- Gibco PBS used for ViCell dilutions (KANBAN) defaults to Vendor Expiry. No governing procedure for assigning expiry | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-05-29 | 2024-05-29 | Its is expected that there is a governing procedure on use of reagents for 8NYA and their associated expiry
3486 | Controlled Copy used on Settle Plate paperwork | Andrew.Brundige@sanofi.com | 68 NYA/4700/QCM | 2024-05-22 | 2024-05-29 | per QU-SOP-0034503 step 6.2 the worksheet should have been printed as controlled print.
3485 | Wave F hold caused loss of temp and overlay control | David.Loik@sanofi.com | 51 NYA | 2024-05-29 | 2024-05-29 | The added weight associated with sampling would not cause a hold. The hold would not cause the process to discontinue control of critical parameters.
3484 | 51NYA - No reportable result for daily CU sample 51-DI2-5 (1219883) | craig.stevenson@sanofi.com | 51 NYA Room 146 | 2024-05-26 | 2024-05-29 | It is expected that sampling vials are integral to perform testing.
3483 | HVAC Ceiling Covers Partially Open/Hanging | brian.o'rourke@sanofi.com | 8NYA Room1403 | 2024-05-29 | 2024-05-29 | Per QU-WIN-00225554 General Area Clearance, Task 1: Verify room, floor, and ceiling integrity is intact | Yes | No Deviation
3482 | DP alarm from 51-116 to hallway 81 from 1905 to 2200 on 28Feb24 | alyson.belleville@sanofi.com | 51-116 | 2024-05-28 | 2024-05-29 | It is expected that there would be no differential pressure alarms from the clean space to the uncontrolled hallway.
3481 | 8NYA - Instructions in FBL-EWI-000015 for SUM install are referencing the previous orientation of the SUM line ports | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-05-28 | 2024-05-28 | N/A
3480 | 8NYA-Gemba-Downstream- Cz DS was moved to cold room prior to BIRFs being printed | aidan.mula@sanofi.com | 8RM1220 | 2024-05-18 | 2024-05-28 | Steps should be signed contemporaneously with performance.
3479 | 8NYA - Lack of instructions to cut seal after tube sealing tote transfer connection in FBL-EWI-000031 | bailey.gleason@sanofi.com | 8NYA | 2024-05-28 | 2024-05-28 | N/A
3478 | Incorrect version of MBR used. | Nicole.lufkin@sanofi.com | 51 New York Ave | 2024-03-10 | 2024-05-28 | Correct version of MBR should have been used
3477 | 51NYA-Thyrogen APU-Room 150-F-0218-243 was found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 51-150 | 2024-05-23 | 2024-05-28 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3476 | 8NYA - GEMBA CZ - Incorrect recipe launched | alex.tello@sanofi.com | 1303 | 2024-05-28 | 2024-05-28 | launch recipe with CV
3475 | Base prime of Pompe SUB X-2240 has too much air in line | dana.hopkins@sanofi.com | 8NYA room 1403 | 2024-05-27 | 2024-05-27 | After prime, line should be full of base
3474 | Failed weld on media line of X-2240 | dana.hopkins@sanofi.com | 8NYA room 1403 | 2024-05-27 | 2024-05-27 | Weld should be successful | Yes | Deviation
3473 | Cerezyme - RF639E002 - 8NYA - Failed SUB bag (visual inspection) | michael.knapp@sanofi.com | 8NYA 1403 | 2024-05-26 | 2024-05-26 | No hair | Yes | Deviation
3472 | Pompe - RF1NAE007 - 8NYA - Torn aseptiquik membrane across sterile boundary during probe installation | michael.knapp@sanofi.com | 8NYA 1403 | 2024-05-26 | 2024-05-26 | Successful aseptic connection of probe to SUB
3471 | Pressure decreasing during pressure stabilization test on SUB 2230 | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-05-24 | 2024-05-24 | SUB bag pressure held with minimum loss.
3470 | Pompe - RF1NAE007 - 8NYA - Unclear EWI instructions for Media Setup | michael.knapp@sanofi.com | 8NYA 1403 | 2024-05-25 | 2024-05-25 | Operation instructions are clear
3469 | Multiple 51 NYA PMs had no comments or extensions associated to PMs completed after due date some WOs have been closed. | ALEX.BISKADUROS@SANOFI.COM | 51 NYA HVAC systems | 2024-05-23 | 2024-05-24 | All extensions attached to associated PMs | Yes | No Deviation
3468 | 8NYA - RF1NME008 Conversation on 805015 transfer. | bailey.gleason@sanofi.com | 8NYA | 2024-05-24 | 2024-05-24 | N/A
3467 | Gemab:51NYA, minicapt 100M (FB-3469) charging port not working | adam.rowe@sanofi.com | 51NYA, 152 | 2024-05-23 | 2024-05-23 | Unit should charge when connected to charger
3466 | 8NYA GEMBA - bad aseptic connection | jack.laplante@sanofi.com | 8NYA RM1350 | 2024-05-24 | 2024-05-24 | Looking for path foward.
3465 | 51 NYA Q Fraction pooling table unable to complete calculations based on a communication issue between DeltaV and PI | craig.valery@sanofi.com | 51 NYA | 2024-05-23 | 2024-05-23 | it is expected that data from DeltaV is sent to PI
3464 | Pompe MES seed train record Day 0 | catherine.pongan@sanofi.com | 1403 | 2024-05-23 | 2024-05-23 | VCD and viability on Day 0 be properly evaluated against criteria per QU-OPE-0522556 |  | Deviation
3463 | 45 NYA - EM/CU Missed weekly DI Water Site 419 Week of 05May2024-11May2024. | craig.stevenson@sanofi.com | 45 NYA, Mechanical Penthouse | 2024-05-11 | 2024-05-23 | It is expected that DI site 419 be collected and tested for Bioburden and TOC on a weekly frequency per QU-SOP-0071840 V14.0 Attachment 1 Page 1
3462 | Cz record RF639E002 MBR for MO update require GEMBA for QU-SOP-0041372 | Bhargav.patel@sanofi.com | 8 NYA / 1403 | 2024-05-24 | 2024-05-23 | SOP instruction follow
3461 | 8NYA GEMBA - shipping bags failed pre visual inspection | marco.vera@sanofi.com | Room 1703 | 2024-05-16 | 2024-05-16 | To pass inspection
3460 | Shipping Conditions of EBYHAM005 Send Outs | kristina.erickson@sanofi.com | QCBP 76 NYA RM 107 | 2024-05-23 | 2024-05-23 | Dry Ice shipment with airtight bag
3459 | 8NYA -RFNME007 Recipe filled bag to outside of weight range. | bailey.gleason@sanofi.com | 8NYA | 2024-05-23 | 2024-05-23 | N/A
3458 | 8NYA, Small Scale WFI Addition, No instructions | scott.howland@sanofi.com | 8NYA | 2024-05-23 | 2024-05-23 | N/A
3457 | 8NYA APU-Dept CC8-4 EasyCal found out of tolerane or no as found data collected during a scheduled calibration | robert.bauch@sanofi.com | 8NYA | 2024-04-22 | 2024-05-23 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3456 | 8NYA-Gemba- Pompe Downstream UF-B Load Material expires before the nano record performs take out. | aidan.mula@sanofi.com | 8RM1703 | 2024-05-22 | 2024-05-22 | Record should automatically take out material prior to expiry.
3455 | Incorrect Test Program used to test filter during SP1 EFG3B04B process | sylvester.kegbeh@sanofi.com | 51 NYA ROOMS 138 | 2024-05-04 | 2024-05-21 | It is expected that the GZ1041-05 filter be tested per QU-SOP-0072849 section 6.3. post IPA wetting.
3454 | TSH- Buffer prep- 45 NYA - Small sample line detached from bag 6 during Buffer fill | Jeremy.Wooden@Sanofi.com | 45 NYA room 190 | 2024-05-22 | 2024-05-22 | The sample line would remain integral and stay attached to the bag.
3453 | No instruction on VI Skid breakdown | Ralston.najarro@sanofi.com | 8NYA1403 | 2024-05-22 | 2024-05-22 | instructions on how to break skid down
3452 | Observation sections not circled for two aseptic quals performed 20APR24 | megan.botti-marino@sanofi.com | 8 NYA | 2024-04-20 | 2024-05-21 | It is expected that no or yes for observations be recorded/circled during the manufacturing process.
3451 | Raw material sample 1216947 was missing the "sampled date" in biolims | lauren.ruffing@sanofi.com | 68NYA 4700 QC Micro | 2024-05-20 | 2024-05-20 | All required sample data is listed in biolims when sample is received/ QU-SOP-0070558
3450 | 76R&D-Multiple errors related to lims from single operator | craig.stevenson@sanofi.com | 76NYA | 2024-05-17 | 2024-05-21 | Samples be accounted for prior to submission to QCM
3449 | Incorrect Clime selected for sample 1056573 in batch 202401115-1 | megan.botti-marino@sanofi.com | 8 NYA | 2024-01-30 | 2024-05-16 | It is expected that the correct Climet be Selected. There is evidence to show which Climet was selected though (MES and zero filter prinouts) and the correct Climet was used.
3448 | 68NYA Support APU-NBOVAL-F-7913 IRTD found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | NBOVAL | 2024-04-25 | 2024-05-21 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3447 | 8NYA APU-Dept CC8-3 UV EasyCal's failed calibration for no as found data taken. | robert.bauch@sanofi.com | 8-1333 | 2024-04-04 | 2024-05-21 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | No | Deviation
3446 | Cold Room 1220 at 8NYA Out of Temperature Range | wilfredo.lopez@sanofi.com | 08RM1220 | 2024-05-20 | 2024-05-20 | Expected temperature range of 2-8C to be maintained. Manufacturing team is in need of guidance on how to handle samples needing to be transported to QC for testing. Sample delivery fridge also out of service. Need to know proper sampling steps in LIMS.
3445 | 8NYA CV CZ skid X3201 Bioburden sample expired | cleiton.lopes@sanofi.com | 8NYA MFG/ 68NYA QCM | 2024-05-18 | 2024-05-20 | Samples to be delivered and tested within 24hrs from collection time
3444 | Freezer Bag Stock Not Created | adam.keister@sanofi.com | 8 NYA, freezer farm | 2024-05-16 | 2024-05-16 | To generate stock for freezer bag and use for thaw when needed | Yes | Deviation
3443 | Clean Steam Generator MECO 5000 #/HR located in 76 NYA emergency shutdown | brizaida.medina-vega@sanofi.com | 76 NYA Clean Steam System | 2024-05-20 | 2024-05-20 | No unplanned Clean Steam Shutdown | Yes | Deviation
3442 | Purification thermometers in use outside of calibrated range | craig.valery@sanofi.com | 51NYA | 2024-05-20 | 2024-05-20 | Equipment calibration should reflect process ranges
3441 | A gemba is needed for: "Shipping bags failed visual inspection". | kenny.leland@sanofi.com | 8NYA\1703 | 2024-05-17 | 2024-05-17 | Pass visual inspection
3440 | 76NYA-Biosurgery APU-Room 3103-HABP-Pressure Indicator Transmitter F-6938-057 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 76-3103 | 2024-05-16 | 2024-05-20 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3439 | WAVE E not rocking and temp control disabled | nicole.lufkin@sanofi.com | 51 NYA | 2024-05-19 | 2024-05-19 | rocking and temp need ot be controlled at setpoints
3438 | High Pre-freezer Temperature alarm after FB DGF4B19 was moved into the freezer | sylvester.kegbeh@sanofi.com | 80NYA | 2024-05-18 | 2024-05-18 | It is expected that the temperature be less than or equal to -70C per QU-OPE-0529697 section 4.0 step 4.5.2 | Yes | Deviation
3437 | 45 Pressure differential weekly log sheet not completed for 16MAY24 | jesse.demers@sanofi.com | 45 NYA | 2024-05-16 | 2024-05-18 | QU-FOR-0091683 | Yes | Deviation
3436 | TSH - 45 NYA - Differential Pressure issues between cold room 174/186E and 189B/188 | Jeremy.Wooden@Sanofi.com | 45 NYA rooms 189B, 174 and 186E | 2024-05-18 | 2024-05-18 | It is expected that all DP readings would be above action level and that the doors between 189B and the MSAT area would stay fully closed. | Yes | Deviation
3435 | SUB bag leaked during decontamination of SUB X2250 (Pompe) | Brandon.steff@sanofi.com | 8RM1403 | 2024-05-17 | 2024-05-17 | ~60kg of base is added to fully decontaminate SUB | Yes | Deviation
3434 | TSH Buffer Prep - Action notice cleanings performed on approved rather than effective version of QU-SOP-0073883. | Jeremy.Wooden@Sanofi.com | 80 NYA rooms 101 and 102 | 2024-05-17 | 2024-05-17 | It is expected that the effective version would be used to perform the cleanings.
3433 | TSH - Buffer Prep- 45 NYA Room 190 DP issue related to RAHU-5 | Jeremy.Wooden@Sanofi.com | 45 NYA Room 190 | 2024-05-16 | 2024-05-16 | It is expected that the air handler would maintain Differential Pressure in room 190.
3432 | 76 R&D LOOP AND DISTRIBUTION SHUTDOWN DUE LOW LEVEL TANK ALARM. | MACJONES.OSEIKUFFOUR@SANOFI.COM | BUILDING 76NYA 347 MECHANICAL ROOM | 2024-05-17 | 2024-05-17 | QU-SOP-0012101
3431 | Connected Pompe Media Tote not readying properly | robert.morawski@sanofi.com | 8NYA CNC | 2024-05-16 | 2024-05-16 | Media tote gave a proper reading in delta v which would allow for the valve to open for the tote swap over.
3430 | Pompe - RF1N9E008 - 8NYA - 50L Wave Equilibration Timer Automation intervention | michael.knapp@sanofi.com | 8NYA 1403 | 2024-05-16 | 2024-05-16 | The 50L wave equilibration timer is completed and the wave is ready for inoculation. | Yes | Deviation
3429 | Nitrates results does not conform to PQ criteria | MA-BioCampus_F&E-Compliance@sanofi.com | 8NYA/Room 1170 WFI System 2 return drops at loops 3A, 3B and 4 | 2024-05-10 | 2024-05-16 | Per PQ-WFI-8NYA-CQ-0005 expected result for Nitrates ≤ 0.06 PPM (Ph. Ru, ChP 2015)
3428 | 8NYA-Pompe- PC02 Prrobes 1 and 2 not re-calibrated correctly per EWI | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-05-16 | 2024-05-16 | It is expected that EWI instructions are followed to determine the path forward based on analytical results, and in a timely manner | Yes | Deviation
3427 | 76NYA-Biosurgery APU-Room 231-Packaging-FB-6580-001 and FB-6580-003 no as found data taken for the last 2 test points | robert.bauch@sanofi.com | 76-231 | 2024-05-16 | 2024-05-16 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3426 | PM Code <PM-014-11 QT> (nonGMP variable PM) Not Executed Since 09August2022 under WO 3040167 (Missed Last 6 PMs) | MA-BioCampus_F&E-Compliance@sanofi.com | N/A | 2022-08-09 | 2024-04-24 | PMs expected to be generated upon closure of last PM code | Yes | Deviation
3425 | WO 2832505 executed and closed by same person | dlspecialistgembarequest@sanofi.com | 08-1503 | 2023-05-10 | 2024-05-15 | different person (typically supervisor) shall review/close WO (MABC-QU-SOP-0074684 GDP SOP)
3424 | Shipping bag Protocol and EWI inconsistent on Pass/Fail determination. | Juan.Bonilla@sanofi.com | 8NYA/1703 | 2024-05-15 | 2024-05-15 | To effectively determine weld failures during inspection.
3423 | 8NYA-Pompe- DO Probe 1 found not fully installed on X2250 during routine process monitoring | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-05-15 | 2024-05-15 | It is expected that the probes remain attached to the SUB after installation and that the area is kept free and clear of clutter when the area is not in use
3422 | Pompe Media buffer FIT has failed twice. | nicholas.grote@sanofi.com | 8RM1350 | 2024-05-15 | 2024-05-15 | that either the primary or secondary wetting will result in a passing FIT test.
3421 | Pompe Pre-Nanofiltration SUT Incorrect | stephanie.aceto@sanofi.com | 8NYA Room 1503 | 2024-05-14 | 2024-05-14 | Tubing material of construction would align to the EWI and the factory calibrated settings within the flow transmitter.
3420 | 8 NYA GEMBA - Pompe Nano filtration pressure spike on pressure control valve | alex.tello@sanofi.com | 1503 | 2024-05-15 | 2024-05-15 | No HIHI pressure alarm
3419 | Pompe Nano- HI Pressure Alarm during Equil | nathaniel.barnes@sanofi.com | 8 NYA 1503 | 2024-05-15 | 2024-05-15 | Correct the ramp rate to allow pressure to stabilize before reaching the HIHI Alarm
3418 | Pompe Nano Filter Leaking At Brass Clamped Outlet Connection | patrick.mitchell@sanofi.com | 8NYA/1503 | 2024-05-14 | 2024-05-14 | It is expected that the filter shouldn't leak.
3417 | Wrong tubing material being utilized for Pompe Nano process | edgar.vale@sanofi.com | 8NYA/1503 | 2024-05-14 | 2024-05-14 | Correct tubing would be used for the process
3416 | 8NYA-Gemba- Pompe Downstream Nano Leak on Second Filter at Tri-Clamp Inlet | aidan.mula@sanofi.com | 8RM1503 | 2024-05-14 | 2024-05-14 | Filter should not leak during process
3415 | 500L SUB Agitator HUB & Agitator Shaft Investigation - Thermo Fisher | sheila.pandolfo@sanofi.com | 8 NYA Penthouse | 2024-04-12 | 2024-05-10 | It is expected the sterile boundary of the SUB is maintained.
3414 | Leak observed during Pompe Nano Equil process | edgar.vale@sanofi.com | 8NYA /1503 | 2024-05-14 | 2024-05-14 | there would be no leaks during the process
3413 | 8NYA - RF1NSE003 UFDF B Load filter Failed second FIT | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-05-13 | 2024-05-13 | N/A | Yes | Deviation
3412 | 76 NYA - HA - Calonge identified in 3106 | craig.stevenson@sanofi.com | 76NYA, Room 3106 | 2024-05-13 | 2024-05-13 | It is expected that employees follow appropriate gowning SOPs
3411 | 8NYA - Improper cleaning of Fume Hood in 1350 | craig.stevenson@sanofi.com | 8 NYA Room 1350 | 2024-05-13 | 2024-05-13 | All cleaning be performed for the appropriate contact time. | No | No Deviation
3410 | 8NYA-Gemba-Fz Butyl Blank Run record does not have mix recipe in EBR | aidan.mula@sanofi.com | 8RM1323 | 2024-05-12 | 2024-05-12 | Recipe should exist in the record for mixing in order to properly sample the blank run. | No | Deviation
3409 | Stock creation issue for Pompe Shipping Bag | zachary.rickard@sanofi.com | 8NYA/1703 | 2024-05-10 | 2024-05-10 | Stock will properly create.
3408 | 51NYA: pH standardization failed | adam.rowe@sanofi.com | 51NYA 122 | 2024-05-12 | 2024-05-12 | pH probe should standardize.
3407 | Shipping bags failed visual inspection | Juan.Bonilla@Sanofi.com | 8NYA\1703 | 2024-05-11 | 2024-05-11 | Pass visual inspection
3406 | Gemba-8NYA- Pompe-EGHDS001 - Shipping bag transfer recipe ended prematurely. | stanley.durand@sanofi.com | 8 NYA Room 1703 | 2024-05-10 | 2024-05-10 | It's expected for recipe to end after filling the final bag. |  | Deviation
3405 | 8NYA - GEMBA - Recipe Does Not End for FZ Shipping Barrel, Automation Intervention Needed | adam.keister@sanofi.com | 8 NYA 131 | 2024-05-10 | 2024-05-10 | For recipe to end when it should
3404 | 8NYA-Pompe- Automated Media Line Valve found in incorrect state (AUTO), causing top off operations to be delayed | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-05-10 | 2024-05-10 | Automagte dvalves in correct state in DeltaV | Yes | Deviation
3403 | 8NYA- Pompe- Base line not primed at start conclusion of Top Off operations | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-05-10 | 2024-05-10 | Base Line prime per EWI
3402 | GEMBA 8NYA UPSTREAM ATF Drain Tee Ready DAC Aseptic Connection Failed Pre-Use Visual Inspection | gary.dodakian3@sanofi.com | 8 nya Suite 2 Upstream 1403 | 2024-05-10 | 2024-05-10 | All Ready DAC seals should be integral and not breached when opened.
3401 | Pompe Post Viral SUM Temp Setting Incorrect | zachary.rickard@sanofi.com | 8NYA/1703 | 2024-05-09 | 2024-05-09 | Product temp setpoints will align with control strategies and be set accordingly. | Yes | Deviation
3400 | Missed Stability Initial Timepoint Window for DGF4B17 Bioassay Test | larissa.fairweather@sanofi.com | 68 NYA / 4700 / QCM | 2024-04-17 | 2024-05-10 | Per stability test request and stability protocol QU-OPE-0224343, the expected start of testing should have been within the testing window of 03Apr2024-13Apr2024.
3399 | Missing witness signature on 2 bulk intermediate reconciliation log (TSK fractions and TSK fraction 8) | craig.valery@sanofi.com | 51NYA RM132 | 2024-05-05 | 2024-05-10 | It is expected that witness steps are executed real time
3398 | Incorrect storage condition listed in PompeICB Sample Plan for Drug Substance samples | rondik.ahmed@sanofi.com | 8NYA/1503 | 2024-05-09 | 2024-05-09 | Per QU-OPE-0527081 the storage condition for all the neoGAA Drug Substance samples should be 2-8 | Yes | Deviation
3397 | AHU outage at 76 NYA | nicholas.macdonald@sanofi.com | Building 76 NYA | 2024-05-09 | 2024-05-10 | It is expected that the air handlers at 76 NYA remain operational. | Yes | No Deviation
3396 | Power Interruption (Potentially Line 1) | tanni.arif@sanofi.com | 76, 51, 80 | 2024-05-09 | 2024-05-09 | ensure AHUs and CSG is up running immediately | Yes | Deviation
3395 | Pompe - RF1NAE005 - 8NYA - Aber probe A and B discrepancy and trend review for SUB top-off | michael.knapp@sanofi.com | 8NYA 1403 | 2024-05-10 | 2024-05-10 | Probes are accurate.
3394 | 51NYA Unplanned Shutdown of AHU-7 | Ryan.Geraghty@sanofi.com | 51NYA: Rooms 149, 152, 152A, 152B, 153 | 2024-05-09 | 2024-05-09 | It is expected that the AHU would not power down unexpectedly | Yes | Deviation
3393 | 8NYA - Gemba Load filter failed FIT a second time RF1NSE002 | william.lee@sanofi.com | 8NYA/1350 | 2024-05-09 | 2024-05-09 | Filter should have passed FIT test
3392 | 68NYA-Metrology-no product impact-Metrology Standard S-1174 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2024-03-25 | 2024-05-09 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3391 | CST/DeltaV Discrepancies - Pompe Upstream | drew.hopkins@sanofi.com | 8NYA | 2024-05-09 | 2024-05-09 | Alarm limits align to or are within process AR per control strategy document.
3390 | Missing a BIFR for SP2 DGF4B19B | sylvester.kegbeh@sanofi.com | 51 NYA ROOMS | 2024-05-08 | 2024-05-08 | It is expected that the SP2 BIRF be returned or filed with the Q record per QU-OPE-0433145 section 5.0 step 5.1
3389 | Gemba – 8NYA – 805003 tote spill in cnc | william.lee@sanofi.com | 8NYA/CNC | 2024-05-09 | 2024-05-09 | 805003 tote should not be leaking
3388 | Buffer leak in CNC Corridor 1210 at 8NYA | Edgar.vale@sanofi.com | 8NYA, Corridor 1210 | 2024-05-09 | 2024-05-09 | Pinch clamp would be secured tightly to prevent leaks | Yes | No Deviation
3387 | Pompe Nano Equil recipe went into hold due to high pressure | Edgar.vale@sanofi.com | 8NYA, room 1503 | 2024-05-09 | 2024-05-09 | No holds during process
3386 | 8NYA - RF525E013 Pre use nano filter failed FIT a second time | bailey.gleason@sanofi.com | 8NYA RM1373 | 2024-05-08 | 2024-05-08 | N/A
3385 | LAL sample associated with Pompe buffer lot that does not exist | katherine.milliken@sanofi.com | 8 NYA | 2023-09-26 | 2024-05-08 | It is expected that the Samples would not be available in OneLIMS for sampling once a lot has been cancelled | Yes | Deviation
3384 | 8NYA-Gemba- Pompe Downstream UF-B Leak at Buffer Supply Pressure Transmitter | aidan.mula@sanofi.com | 8RM1503 | 2024-05-08 | 2024-05-08 | SUT should not leak if installed correctly. | Yes | Deviation
3383 | GEMBA-8NYA- EGHUA003 - Pompe UF-A stock creation expiry discrepencies. | stanley.durand@sanofi.com | 8 NYA room 1503 | 2024-05-01 | 2024-05-07 | It's expected for expiries to align with stock creation | Yes | No Deviation
3382 | 8NYA APU-Room 1403-UV EasyCal FB-4823 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1403 | 2024-02-21 | 2024-05-07 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
3381 | 51NYA-Thyrogen APU-Room 138 Purification-UV EasyCal F-4235-009 had no as found data taken during an unscheduled calibration. | robert.bauch@sanofi.com | 51-138 | 2024-04-25 | 2024-05-07 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
3380 | ANX Eluate Filter Failed 2nd FIT Attempt - RF1Y6E002 | nicholas.grote@sanofi.com | 08-1290 | 2024-05-07 | 2024-05-07 | Passing FIT result
3379 | 8NYA - EGHAE002 - Pompe ANX CX Values Failed to Propagate for Material Storage Unit | wilfredo.lopez@sanofi.com | 8RM1503 | 2024-05-06 | 2024-05-06 | That the step executes automatically and the CX values are written to the storage unit
3378 | Incorrect iLIMS collection date selected | craig.valery@sanofi.com | 51NYA RM132 | 2024-05-06 | 2024-05-06 | it is expected that the correct dates are selected |  | No Deviation
3377 | Pompe - RF1NAE005 - 8NYA - Bleed and Base tote RFID reading follow-up | michael.knapp@sanofi.com | 8NYA 1403 | 2024-05-06 | 2024-05-06 | All totes are able to be properly read through deltaV. | Yes | No Deviation
3376 | 805004 - BOM Discrepancies within EBR RF1N1E003 | nicholas.grote@sanofi.com | 08-1350 | 2024-05-06 | 2024-05-06 | All SUT that is needed for weight out should be included in the BOM.
3375 | HA Lots - EBYHAM002, EBYHAH005; Steps executed outside of Batch Record instruction | charlie.houle@sanofi.com | 76 NYA; Room 3102 (HA Purificaiton) | 2024-03-20 | 2024-05-02 | QU-OPE-0434886, v.15.0, step 7.19.23 states if the pH is between 6.00 and 6.60 to N/A section 7.20. The expected result would be that section 7.20 would be N/A'd, not executed.
3374 | Pompe- ANX Load Dilution- HIHI Weight Alarm | nathaniel.barnes@sanofi.com | 8 NYA 1503 | 2024-05-06 | 2024-05-06 | Automation intervention to allow recipe to complete and forward process.
3373 | 8NYA - RF1NQE003 UFDF filter failure | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-05-05 | 2024-05-05 | N/A
3372 | UV Meter failed standardization at the Equilibration step during TSK DGF4B22A Process at 51NYA. | sylvester.kegbeh@sanofi.com | 51NYA | 2024-05-05 | 2024-05-05 | Per QU-OPE-0295249 and QU-SOP-0014246, UV mete should be standardized prior to loading the product to the column
3371 | 8 NYA - LAL Sample 1179869 (POM_COP_808928_EFF) P/N: 808928 LOT: EGHCE002. LAL Pompe sample submitted in polypropylene tube. | roodline.bertrand@sanofi.com | 68NYA/4700, QC MICRO | 2024-05-04 | 2024-05-05 | QU-SOP-0070395 STEP # 6.3.3 samples must be in a polystyrene tube for LAL testing.
3370 | 8NYA - RF1NFE005 Condo out of range 805002 | bailey.gleason@sanofi.com | 1350 | 2024-05-05 | 2024-05-05 | hold gemba to see path forward
3369 | GEMBA for Pompe SUB X2250 Base tote switchover | Brandon.Steff@sanofi.com | 8RM1403 | 2024-05-05 | 2024-05-05 | RFID correctly reads tote in parking spot.
3368 | Gemba – 8NYA – 805009 Sodium Meta-Periodate discarded incorrectly into drain | william.lee@sanofi.com | 8NYA/1350 | 2024-05-03 | 2024-05-04 | 805009 Sodium Meta-Periodate should have been discarded by hazardous waste pickup.
3367 | Gemba-8NYA Downstream- Low Low DP Alarm in Room 1323, DP is not stabilizing | aidan.mula@sanofi.com | 8RM1323 | 2024-05-04 | 2024-05-04 | Determination of impact at Gemba
3366 | Gemba – 8NYA – ANX filter failed FIT three times | william.lee@sanofi.com | 8NYA/1350 | 2024-05-03 | 2024-05-03 | Filter should have passed FIT test | Yes | Deviation
3365 | 8NYA-805009 new MBR version has doubled the weight of the acetic acid dispense | christopher.arnold@sanofi.com | 8RM1350 | 2024-05-03 | 2024-05-03 | That a new MBR version would not change the formulation unless it was designed to do so. | Yes | Deviation
3364 | 8NYA - RF1NDE010 Stock cannot be created within the record | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-05-03 | 2024-05-03 | hold gemba to see path forward. create stock manually and have MLCs create stock in SAP
3363 | 51CV-0008 VP-107-04 240004932 failed COND (for RQ) | larissa.fairweather@sanofi.com | 68 NYA / 4700 / QCM | 2024-03-04 | 2024-03-29 | Per QU-FOR-0005429 attachment 22, sample should pass per USP <645>
3362 | 8NYA-GEMBA-Downstream Copper Follow Up on Path Forward for next lot from QE-030515 | aidan.mula@sanofi.com | 8RM1503 | 2024-05-03 | 2024-05-03 | Discuss path forward in Gemba
3361 | Stock creation could not be performed . | jermaine.kellam@sanofi.com | 8RM1350 | 2024-05-02 | 2024-05-02 | Stock creation should have occurred normally
3360 | 8NYA - RF1NME006 BOM issue. 200L bag was added to BOM and cannot be used. Resulting in stock creation error and change of BOM. | bailey.gleason@sanofi.com | 8NYA | 2024-05-02 | 2024-05-03 | hold gemba to see path forward
3359 | 76 NYA - QC Micro - No reportable Non-Viable EM result for site 425-5 in April 2024. | craig.stevenson@sanofi.com | 76 NYA Room 425 | 2024-04-30 | 2024-05-02 | It is expected that site 425-5 have Non-Viable testing performed on a monthly frequency.
3358 | ANX Eluate filter failed FIT twice | jermaine.kellam@sanofi.com | 8RM1373 | 2024-05-02 | 2024-05-02 | The filter should pass on initial testing.
3357 | 8NYA - EGHDT - Fabrazyme Post Viral Mix Recipe Required Automation Intervention | wilfredo.lopez@sanofi.com | 8RM1313 | 2024-05-02 | 2024-05-02 | That the agitation and mixing procedure takes place and that the equipment operates normally
3356 | 8NYA - EGHDS007 - Fabrazyme Post Viral Recipe Needing Automation Intervention to End Transfer Recipe | wilfredo.lopez@sanofi.com | 8RM1313 | 2024-05-02 | 2024-05-02 | That the recipe ends automatically at the end of the transfer process
3355 | Concentration factor of A280 values for TSH Q fractions not reported on final result | alyson.belleville@sanofi.com | 51-131 | 2024-05-02 | 2024-05-02 | It is expected that the actual A280 values would be reported in iLIMS
3354 | 51NYA: Leak from connection from media to colder app during refeed on E4 | adam.rowe@sanofi.com | 51NYA, 140 | 2024-05-02 | 2024-05-02 | There should be no leaks during the operation
3353 | UFDF-A, Second FIT Attempt Failed - RF1NQE003 | nicholas.grote@sanofi.com | 08-1373 | 2024-05-02 | 2024-05-02 | FIT test passes and there is not a Signiant release of WFI from the outlet side of the filter.
3352 | Missing POU sample from Cz Media (Part# 760741) | robert.morawski@sanofi.com | 8NYA CNC | 2023-12-17 | 2024-05-02 | Samples would have been submitted in BioLIMS at the time of sampling. | Yes | Deviation
3351 | Gemba - 8 NYA - EGHDS007 - Shipping bag failed visual inspection. | stanley.durand@sanofi.com | 8NYA 1313 | 2024-05-02 | 2024-05-02 | Bags are expected to not have questionable welds.
3350 | Liquid Spill in Room 1422 (CNC area) | robert.morawski@sanofi.com | 8NYA CNC 1422 | 2024-05-02 | 2024-05-02 | Process waste piping was aligned properly with the drain hub so liquid did not splash onto the floor.
3349 | UFDF Filter Tested in ANX MO | nicholas.grote@sanofi.com | 08-1350 | 2024-04-30 | 2024-05-01 | DS filters should be tested in the correct records.
3348 | Pompe - RF1NLE004 - 8NYA - Seed Train Media Assigned 14 day Expiry | michael.knapp@sanofi.com | 8NYA 1403 | 2024-05-01 | 2024-05-01 | Correct expiry is assigned to the media batch
3347 | EBR Pointing to Non-Existent PS80 Setup Child in EWI | nicholas.grote@sanofi.com | 08-1350 | 2024-05-02 | 2024-05-02 | EBR should point to the correct instruction. The same task is performed for DTT preparation and operators are well versed in performing this task. The EWI Child can be duplicated and edited so that it pertains to PS80.
3346 | 76-AHU 20 Fan failure (Fault overload) | john.beaule@sanofi.com | 76-3102 | 2024-05-01 | 2024-05-01 | FBS-SOP-000747, area pressures
3345 | GEMBA- 8 NYA- 1501 MAL CNC Side Door Stuck Open | katherine.reardon@sanofi.com | 8 NYA 1501 MAL | 2024-05-01 | 2024-05-01 | Expected result is that PAL leading into 1503 ballroom is established as MAL in order to carry out necessary POMPE processing while the MAL door is being fixed.
3344 | Fraction table for Q DGF4B19B says to not pool fraction 3, historically pooled | alyson.belleville@sanofi.com | 51-131 | 2024-04-30 | 2024-04-30 | The expectation is that MFG pools fractions based on information from the fraction table. There's been no deviation to that. Fraction 3 is normally included in pooling.
3343 | April 2024 Late PMs | Elizabeth.cook@sanofi.com | 51 NYA - Gas Shed | 2024-05-01 | 2024-05-01 | All extended PMs are completed before due date | Yes | Deviation
3342 | Bug Found in Solution Prep Room 1350 | nicholas.grote@sanofi.com | 08-1350 | 2024-05-01 | 2024-05-01 | No bugs inside the lab.
3341 | 76-AHU 20 Fan failure (Fault overload) | john.beaule@sanofi.com | 76-3102 | 2024-05-01 | 2024-05-01 | FBS-SOP-000747, area pressures | Yes | Deviation
3340 | Problems with Elution on ANX for Pompe | patrick.mitchell@sanofi.com | 8NYA/1503 | 2024-04-30 | 2024-04-30 | It is expected that the bag not have any leaks or exposure to the environment.
3339 | 8NYA-DOWNSTREAM-POMPE-EGHUA003-COLLECTION SUM CONNECTION | MARC.LATTA@SANOFI.COM | 1503 | 2024-04-30 | 2024-04-30 | SUT referred by EWI should be on the BOM
3338 | To mitigate a PPQ IPC failure we would like to increase the load density/resin concentration to 15g/L. There is an offset in the UV sensor on column 1 that has potentially resulted in a lower loading which is potentially impacting the eluate conc. Need to make the change as soon as possible before the current column is done loading/eluting | dana.hopkins@sanofi.com | 8NYA room 1403 | 2024-04-30 | 2024-04-30 | Consistent column loading
3337 | HA Lot - EBYFRM021; Undefined line/spike on Fermentation chart | charlie.houle@sanofi.com | 76 NYA; Room 364 (HA Fermentation) | 2024-04-30 | 2024-04-30 | QU-OPE-0309296, v19.0, step 7.11.20 gives instruction for review of the Fermentation chart. The instructions list the ranges for each process parameter. It is expected that each parameter stayed within range during Fermentation
3336 | 3 Failed FITs on Millex filters at 51NYA | nicholas.shoenfelt@sanofi.com | 51-143 | 2024-04-25 | 2024-04-25 | It is expected that the filters pass pre-use FIT.
3335 | WIP label reprinted and illegible label was discarded and not reconciled in the record. | adam.hildebrand@sanofi.com | 51RM140 | 2024-01-28 | 2024-04-30 | It is expected that labels will not be discarded and will be reconciled appropriately.
3334 | Residue found on agitator HUB with evidence of leak by-Pompe PPQ4 bag | nicole.stlawrence@sanofi.com | 8 NYA | 2024-04-29 | 2024-04-29 | The epected results are there is no evidence of leak by on the agitator HUB.
3333 | Agitator failed to start when prompted. | patrick.mitchell@sanofi.com | 8NYA/1503 | 2024-04-29 | 2024-04-29 | It is expected that the agitator would be turned on and start mixing. | Yes | Deviation
3332 | 8NYA - Recipe for Pompe Copper pooling filter is not available in Delta V | bailey.gleason@sanofi.com | 8NYA | 2024-04-29 | 2024-04-29 | Hold gemba to see path forward | Yes | Deviation
3331 | 8NYA - RF1Y5E001/4428001 Pompe Copper pooling filter Needs to be wetted and does not have connection. | bailey.gleason@sanofi.com | 8NYA | 2024-04-29 | 2024-04-29 | Hold gemba to see path forward
3330 | Humidor at 74NYA Laydown area in use, not calibrated | david.zhong@sanofi.com | 74 NYA Laydown area UT1-02 | 2024-02-01 | 2024-04-29 | BQ-GOP-000248: Equipment is overdue for calibration. | Yes | Deviation
3329 | Wave F not rocking, Temp control not working, Gassing not working | adam.hildebrand@sanofi.com | 51RM152B | 2024-04-29 | 2024-04-29 | It is expected that the Wave would be rocking and in a state of control within process limits.
3328 | Innoc prep room 8RM1343 DP alarm during zero filter for Pompe thaw | Brandon.Steff@sanofi.com | 8RM1343 | 2024-04-29 | 2024-04-29 | Air handlers maintain differential pressure and no EMS alarms occur during operations
3327 | Media leak during POU sample | Ralston.najarro@sanofi.com | CNC AREA | 2024-04-29 | 2024-04-29 | No leaks | Yes
3326 | 8 NYA -POMPE -COPPER -High high pressure during flowthrough-EGHCE001 | MARC.LATTA@SANOFI.COM | 1503 | 2024-04-29 | 2024-04-29 | there should not be a high hig hpressure alarm during flowthrough
3325 | Incorrect MES BOM for Liquid Weigh Outs - RF1NBE001 (804993) and RF1N5E001 (805010) | nicholas.grote@sanofi.com | 1350 | 2024-04-27 | 2024-04-27 | BOM should have the correct parts and raw materials listed for operation.
3324 | 51NYA: Wave F had a communication error and went into hold | adam.rowe@sanofi.com | 51NYA, 152 | 2024-04-26 | 2024-04-27 | Wave F should not be in hold, the recipe should be running
3323 | 8NYA - Gemba needed for automation intervention for the sanitization of WFI loop 2. | kenny.leland@sanofi.com | 8 NYA/ Room 1323 | 2024-04-24 | 2024-04-24 | UFDF recipes should not be running during the sanitization of the WFI loop.
3322 | GEMBA - 8 NYA - Baseline NWP Low on Pompe UF/DF A skid, X3431 | adam.keister@sanofi.com | 8 NYA 1503 | 2024-04-25 | 2024-04-25 | Path forward on how to re install membranes to re-establish baseline NWP.
3321 | Bulk HA - Media Fill; RQ-F076-MFR-76NYA-CQ-0013 - Particulate found in two bottles | charlie.houle@sanofi.com | Room 354 (Aseptic Fill Core); QC Micro lab | 2024-03-29 | 2024-03-29 | No particulate should be observed
3320 | Sample fridge fb-2956 in sample drop off room alarming for high temp | Brandon.Steff@sanofi.com | 8RM1410 | 2024-04-25 | 2024-04-25 | Samples fridge temp stays within 6-8C
3319 | 8 NYA GEMBA- FZ UFDF Pump stopping during hot WFI | alex.tello@sanofi.com | 1323 | 2024-04-25 | 2024-04-25 | for pump to work properly
3318 | MCC X2831 column 2 required in process pH calibration | Brandon.Steff@sanofi.com | 8RM1403 | 2024-04-20 | 2024-04-20 | pH for equilibration buffer stays within range of recipe
3317 | Cz MMC Column C3112010 damage during movement | Christian.Venegas@Sanofi.com | 8NYA Cell Culture Phase 1 Room 1303 | 2024-04-21 | 2024-04-21 | To get a Phath forward
3316 | MCC 2811 Return pipe positioned incorrectly leading to NaOH spill | Brandon.Steff@sanofi.com | CNC space 8NYA | 2024-04-25 | 2024-04-25 | Waste pipe drains waste down drain without spilling or spashing
3315 | Wave E recipe hold | adam.hildebrand@sanofi.com | 51RM152B | 2024-04-25 | 2024-04-25 | Recipe is expected to run steady state with no issues unless otherwise detected through built in markers of the process.
3314 | UFDF WFI drop needs to be dropped from recipe | nathaniel.barnes@sanofi.com | 8NYA 1323 | 2024-04-24 | 2024-04-24 | Drop released by automation to run WFI Sani
3313 | Thyrogen Discard GZ5157-00 BGF1S13 procedural gap | joshua.brooks@sanofi.com | 80NYA (freezer location) | 2024-04-24 | 2024-04-24 | No Gap
3312 | Phase 1 UF\DF Skid Rupture Disk Leaking During Equil | Christian.Venegas@Sanofi.com | 8NYA Cell Culture Phase 1 Room 1323 | 2024-04-24 | 2024-04-24 | To get a Phath forward | Yes | Deviation
3311 | Turbo and standard suitability for TOC machines has not been performed | kathryn.lentoni@sanofi.com | 68NYA, 4700, QCM | 2022 | 2024-04-24 | Expected results are Turbo and Standard mode suitability should have been performed annually
3310 | Clamps on ATF bottomed out and able to be spun | gregory.klotz@sanofi.com | 8NYA | NA | NA | expect that SUT received will be integral and clamps will not be bottomed out and freely spun
3309 | 8NYA - FZ GEMBA - PUMP NOT WORKING DURING HOT WFI FLUSH - UFDF | alex.tello@sanofi.com | 1323 | 2024-04-23 | 2024-04-23 | for pump to work properly
3308 | TSH Harvest Bag Missing Pinch Clamps on Sample Bags | james.letourneux@sanofi.com | 51NYA140 | 2024-04-24 | 2024-04-24 | All harvest bags should have pinch clamps installed along each segment of tubing. I am not aware of an explicit instruction in a validated document that dictates this, but without these clamps, it is more likely (or certain) that sterility is compromised.
3307 | Wave in hold and unable to be restarted. | adam.hildebrand@sanofi.com | 51RM152B | 2024-04-24 | 2024-04-24 | It is expected that the Wave process is uninterrupted.
3306 | WFI Loop 2 Failed to Initialize Daily Sanitization | wilfredo.lopez@sanofi.com | 8RM1323 | 2024-04-24 | 2024-04-24 | It is exepected that the sanitization operation takes place at 0100 each day. All WFI drops must be released and ready for use at that time. No batching or recipes that required use of WFI must be taking place at this time for the operation to be successful without automation intervention.
3305 | Room 1333 Refrigerator Temperature Alarms missing comments | michael.knapp@sanofi.com | 8NYA 1333 | 2024-03-31 | 2024-04-23 | Alarms are commented on and QA reviewed within 72 hrs
3304 | HA Media Fill P/N: GZ3107-01 LOT: EBYMFC001 contained particles in 2 bottles. | roodline.bertrand@sanofi.com | 76 NYA | 2024-03-21 | 2024-03-21 | Per QU-SOP-0014243 samples should be negative for turbidity.
3303 | 8NYA - Out of KCL | stanley.durand@sanofi.com | 8NYA | 2024-04-23 | 2024-04-23 | Per FBL-EWI-000010 and FBL-EWI-000022 KCL is required for storage of pH probes
3302 | E4 refeed performed with Q status media | David.Loik@sanofi.com | 51 NYA | 2024-04-23 | 2024-04-23 | Media is available for use as needed.
3301 | Gemba needed for 8 NYA UFDF A cleaning with CV. Automation intervention needed of recipe went into hold. | kenny.leland@sanofi.com | 8 NYA/ Room 1503 | 2024-04-23 | 2024-04-23 | LOTO should have been performed at prompt for sampling. | Yes | Deviation
3300 | 68 NYA - QC Micro - EMS alarms not responded to in adequate time | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2024-02-05 | 2024-04-23 | It is expected that the Alarm Responder, Impacted Department and QA complete acknowledgement of EMS Alarms within 72 hours of the original alarm. | Yes | No Deviation
3299 | 45-51NYA Thyrogen Blue Eluate Pool LOT: DGF3B21A- Bioburden sample 2400004317 was placed in the incorrect incubator | roodline.bertrand@sanofi.com | 68NYA/4700, QC MICRO | 2024-04-20 | 2024-04-23 | Per QU-SOP-0073109, rhTSH Blue Eluate Pool samples must be incubated in a 30-35 C temperature range for 3 to 5 Days. | Yes | Deviation
3298 | 8 NYA GEMBA Request- Broken Window room 1503 | richard.manasian@sanofi.com | 1503 | 2024-03-21 | 2024-03-21 | Building would not have damage to window
3297 | Material 713989 was used for slurry bag setup but was not part of the BOM | MARC.LATTA@SANOFI.COM | 1303 | 2024-04-07 | 2024-04-18 | There is a discrepancy between the BOM and the EWI | No | Deviation
3296 | Pompe- RF1NAE002- 8NYA- VI waste overfilling with air | jeffrey.boldizar@sanofi.com | 8NYA 1403 | 2024-04-22 | 2024-04-22 | Able to manage waste without additional change outs due to pressure. | Yes | No Deviation
3295 | VI Skid mini eluate went to waste | kurt.marsocci@sanofi.com | 8 NYA/ Room #1403 | 2024-04-21 | 2024-04-21 | Wait for the next mini eluate push to complete eluate 7 | Yes | No Deviation
3294 | GEMBA for Counts on TSA Negative Control from Aseptic Operations at 45 NYA | jessica.kychler@sanofi.com | QC Micro 68 New York Avenue, floor 4, room 4700 | 2024-04-11 | 2024-04-20 | It is expected that negative control plates be free of microbial contamination
3293 | Wrong standardization for pH | daniela.medina2@sanofi.com | 8NYA/ 1350 Support Services | 2024-04-19 | 2024-04-19 | To sample with the correct standard,
3292 | Seed train for pompe validated prior to restart of the ppq restart, but QC microbiology report not yet validated | katherine.milliken@sanofi.com | 8 NYA | 2024-04-18 | 2024-04-18 | It is expected that the microbial validation would occur concurrently with the manufacturing validation.
3291 | 8NYA - 1000L Bag failed pre inspection HU1001286187 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-04-19 | 2024-04-19 | hold gemba to see path forward | Yes | Deviation
3290 | 8NYA - Filter tested under incorrect parameters RF1CBE007_003 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-03-26 | 2024-04-19 | N/A
3289 | Thyrogen - SP1 - Bubble Trap PRV leaking during rinse phase | cindy.guerra@sanofi.com | 51NYA - 138 | 2024-04-19 | 2024-04-19 | PRV would not leak during rinse phase |  | Deviation
3288 | R2A Mislabeled as M-ENDO Submitted for Growth Promotion Testing | jessica.kychler@sanofi.com | QC Micro 68 New York Avenue, floor 4, room 4700 | 2024-04-16 | 2024-04-18 | It is expected that M-ENDO is purple in color and 1 sleeve is submitted in light-proof packaging to QC Micro
3287 | 8NYA-250L multiple SUM' P/N 692168 Pre-Use Inspection failures | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-04-12 | 2024-04-18 | It is expected that there are no apparent manufacturing defects associated with material for use in cGMP production activities | No | No Deviation
3286 | Leak from bad seal on 100mM Sodium Acetate pH 5.6 Buffer - 805003 | nicholas.grote@sanofi.com | 1202 | 2024-04-18 | 2024-04-18 | Good seal, no leaking. | Yes | Deviation
3285 | 8NYA-Pompe-RF1AE002-Eluate 6 not collected due to skid troubleshooting activities | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-04-17 | 2024-04-14 | It is expected that VI mini eluates are collected
3284 | Incorrect incubator selected for bioburden sample incubation | megan.botti-marino@sanofi.com | 68 NYA, QC Microbiology Lab | 2024-04-06 | 2024-04-17 | Per attachment 1 of MABC-QU-SOP-0073109, coliform bioburden samples are to be incubated at 34-36 degrees C.
3283 | Pompe - RF1NAE002 - 8NYA - MCC Bubble Trap Low Alarm | michael.knapp@sanofi.com | 8NYA 1403 | 2024-04-17 | 2024-04-17 | No bubble trap alarm for air. Line is able to be manually primed per EWI.
3282 | GEMBA - 8 NYA - Slurry percentage average out of range for column packing | adam.keister@sanofi.com | 8 NYA RM 1303 | 2024-04-17 | 2024-04-17 | Path forward | Yes | Deviation
3281 | Critical Utility site 68-RO-2 not aligned between A/N SOP and loop location | craig.stevenson@sanofi.com | 68NYA Room 1310 | 2024-04-17 | 2024-04-17 | It is expected that QU-SOP-0074751 V6.0 Attachment 1 and 2, QU-SOP-0064149 V4.0 Attachment 3, and QU-SOP-0022403 V18.0 Table 5 align with sampling requirements and specifications | Yes | Deviation
3280 | EMS alarms for Cryo-storage Freezers at 45NYA and 74NYA | Adriana.Alban@sanofi.com | 74NYA/166 and 45NYA/101 | 2024-04-15 | 2024-04-15 | Is it expected that QCA's send out notifications to area owner. It is expected that EMS connection with equipment is maintained.
3279 | Gemba – 8NYA – Insect captured in room 1350 | william.lee@sanofi.com | 8NYA/1350 | 2024-04-16 | 2024-04-16 | There should be no insects in solution prep room 1350 |  | Deviation
3278 | NBO Cleaners not trained on cleaning procedure for QCRM lab room 1204 | Marc.Guerin@Sanofi.com | 11 Forbes Road NBO, QCRM Lab Room 1204 | 2024-04-15 | 2024-04-12 | expected results is that all cleaning personnel are to be trained on MABC-QU-SOP-0014502-Disinfection Program for the NBO QC Raw Materials Sampling Suite, Associated Area prior to performing any cleaning activities
3277 | Weight not taring on VI-Skid mixer. | Jonathan.Ahiaba@sanofi.com | 8NYA, 1403 | 2024-04-10 | 2024-04-10 | The weight indicate was tared on deltaV |  | No Deviation
3276 | 68NYA Support APU-Room 1830-Metrology Pressure Standard FBS-0073 and Flow standard S-1357 found out of tolerance during scheduled calibrations | robert.bauch@sanofi.com | 68-1830 | 2024-02-27 | 2024-04-16 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3275 | Missed Sample for Eluate 4 - 50mL Retain for MSAT | phoebe.karkos@sanofi.com | 8NYA 1403 | 2024-04-15 | 2024-04-15 | MABC-QU-OPE-0527081-Pompe-ICB Sampling Plan for 8 NYA table 9 prefiltration eluate retain for PPQ for MSAT 50ml in sampling bag
3274 | Unlabeled 50 mL sample bag in the QC bin/shelf inside 8NYA laydown fridge | rondik.ahmed@sanofi.com | 8NYA/1410/Fridge FB-2956 | 2024-04-14 | 2024-04-14 | All samples that are dropped off in the fridge to be picked up must be labeled
3273 | 8NYA – 705682 batch discarded due to agitator issue | william.lee@sanofi.com | 8NYA/1350 | 2024-04-13 | 2024-04-13 | Agitator should not have moved out of place and bent. | Yes | No Deviation
3272 | Pompe - RF1NAE002 - 8NYA - VI and MCC Communication Error | michael.knapp@sanofi.com | 8NYA 1403 | 2024-04-14 | 2024-04-14 | VI mixer and MCC communicate to ensure proper elution transfer.
3271 | Pompe - RF1N9E005 - 8NYA - Minicapt unable to achieve required flowrate | michael.knapp@sanofi.com | 8NYA 1343 | 2024-04-14 | 2024-04-14 | Minicapt operates without flow alarms
3270 | 8 NYA and 45-51 NYA CV Conductivity samples tested after failed standard verification (Lots: RF626E001-004, 51CV-0005) | Andrew.Brundige@sanofi.com | 68 NYA/4700/QCM | 2024-04-13 | 2024-04-13 | Per QU-OPE-0115243 Step 6.3.3 the standard verification must be between 81.48 and 86.52
3269 | 8NYA-Pompe-EGHAF002- VI skid hold due to high pressure between skid and eluate SUM | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-04-14 | 2024-04-14 | It is expected that SUM bag is in suitable condition to recieve all eluate pushes
3268 | 8NYA-Pompe-Flask Media P/N 805012 found with open clamp and wetted sterile boundry prior to use in Production activities | john.lytle@sanofi.com | Cold Rm 1420/ 8NYA | 2024-04-14 | 2024-04-14 | it is expected tyhat pich clamps be closed prior to dispensing aliquots
3267 | 68 NYA - RO Tank Low Level Because Air Compressor Lost Air | yunbo.lou@sanofi.com | 68 NYA / 5600 | 2024-04-14 | 2024-04-14 | The air compressor is expected to be working. | Yes | Deviation
3266 | GEMBA - 8 NYA - Wrong Column Clean Recipe Launched | adam.keister@sanofi.com | 8 NYA, 1303 | 2024-04-13 | 2024-04-13 | To get recipe to change over to "with CV," with automation's help
3265 | 51NYA: 7.5 hour time limit exceeded for inoc procedures/microcarriers found in clean steam piping | adam.rowe@sanofi.com | 51NYA, 140 | 2024-04-12 | 2024-04-12 | There should not be any microcarriers in the clean steam piping.
3264 | 8NYA Pompe retain EGHAF002 ELU-3 leak | cleiton.lopes@sanofi.com | 68NYA | 2024-04-13 | 2024-04-13 | Retains to pass visual inspection at receipt | Yes | Deviation
3263 | Missing zip tie on 2000L Sartorius Bag 813913 | zachary.rickard@sanofi.com | 68NYA | 2024-04-12 | 2024-04-12 | Bag should be properly assembled with all zip ties in place | Yes | Deviation
3262 | Supplier Event required for AirGas, unable to recieve CO2 delivery due to new tank seals being applied post DOT stop. | sarah.adams@sanofi.com | 68 NYA, QCC | 2024-04-11 | 2024-04-11 | Tucks seals match provided paperwork.
3261 | Approved version of Document Used for Aseptic Requalification EM | jessica.kychler@sanofi.com | QC Micro 68 New York Avenue, floor 4, room 4700 | 2024-04-02 | 2024-04-12 | It is expected that the effective version of a document will be printed for use | Yes | Deviation
3260 | Pompe - RF1NAE004 - 8NYA - SUB Atypical DO Sparge and ATF Exchange Trends | michael.knapp@sanofi.com | 8NYA 1403 | 2024-04-12 | 2024-04-12 | No atypical trends during processing
3259 | Unused SUT issued to Manufacturing Orders in Room 1605 | sheila.pandolfo@sanofi.com | 8 NYA Room 1605 | 2023-06-03 | 2024-04-11 | It is expected that parts not required for a manufacturing order are not on the Bill of Materials. It is expected that parts issued to a manufacturing order are used. If a part is scanned in and not used, it is expected the part is reversed from the Manufacturing order and the parts are returned to inventory.
3258 | Unable to QC Release Lots with Cancelled Samples in OneLims | amy.roy@sanofi.com | 68NYA/4700/QC Micro | 2024-03-28 | 2024-03-28 | QC should be able to perform lot dispostion of Onelims lots with cancelled samples. "All tests done" should only be a block for non-cancelled samples
3257 | Weight indicator FB-4201-006 on Eluate sum PBM285001 did not tare properly before mini-eluate collection | Jonathan.Ahiaba@sanofi.com | 8NYA, 1403 | 2024-04-10 | 2024-04-10 | The weight indicate was manually tared
3256 | EWI instructions for FBL-EWI-000032 Blank run Pompe Task 2 steps 2 and 6 are swapped | Brandon.Steff@sanofi.com | 8RM1403 | 2024-04-10 | 2024-04-10 | Column outlet line is welded to right side of the column connect assembly containing sample bag prior to welding column inlet line to left side of column without sample bag T connection
3255 | EWI 32 Step not executed prior to Buffer Prime start | Yves.Antoine@sanofi.com | Room 1403 | 2024-04-10 | 2024-04-10 | We expect that EWI 32 should be followed throughly prior to buffer prime start
3254 | Gemba – 8NYA – Propylene glycol below target in 705682 batch | william.lee@sanofi.com | 8NYA/1350 | 2024-04-10 | 2024-04-10 | Target should be reached for Propylene Glycol addition
3253 | pCO2 transmitter PI pulls are reversed in ATF sampling loop of SUB record RF1NAE004 | Brandon.Steff@sanofi.com | 8RM1403 | 2024-04-10 | 2024-04-10 | Transmitter 1 result is pulled into pCO2 transmitter 1 section of record.
3252 | Anti-foam stir plates and scales not onboarded in Gage system | dom.arone@sanofi.com | 8 NYA, RM 1333 and 1403 | 2024-04-10 | 2024-04-10 | QU-OPE-0521196, Attachment 1 and 2 decision tree
3251 | Gemba – 8NYA – Leak noticed on SUT during buffer preparation | william.lee@sanofi.com | 8NYA/1350 | 2024-04-10 | 2024-04-10 | Bag should not leak from bottom outlet
3250 | Gemba 51NYA: pressure relief valve on E4 breached during SIP cooldown | adam.rowe@sanofi.com | 51NYA, 140 | 2024-04-09 | 2024-04-09 | If the jacket pressure is out of range, it should have been adjusted before starting SIP.
3249 | PM Parts Attachment not Signed for in SP1 Pack Record EGFCP001 | craig.valery@sanofi.com | 51 NYA | 2024-02-23 | 2024-04-08 | it is expected that applicable information is recorded in mbr
3248 | Pompe - RF1NAE002 - 8NYA - MCC port allocation error | michael.knapp@sanofi.com | 8NYA 1403 | 2024-03-28 | 2024-04-09 | The valves are appropriately identified and allocated at the time of installation.
3247 | 51 AHU, Loss of Corridor Differential Pressure | craig.valery@sanofi.com | 51 NYA | 2024-04-09 | 2024-04-09 | it is expected that environmental conditions remain with in process specifications
3246 | 8NYA - Gemba - BRR Kickback for RF563E002_001 -multiple FIT attempts | alejandro.martinez2@sanofi.com | 8RM1350 | 2024-03-13 | 2024-04-27 | CAPA needed to implement clarity to EWI-21 ; equipment issues shouldn't count as failures in MES
3245 | Unverifiable source for bioburden EM settle plate sampling end time | Rebecca.Schmidt@sanofi.com | 68NYA/4700/QCM | 2024-03-23 | 2024-04-01 | QU-SOP-0057388 Step 6.2.4, sample time period must be recorded after the sample lid was replaced once testing is complete.
3244 | 8NYA APU-CC8-Room1403-FB-2938-009 Temperature Indicator on skid PBM-2570001 Product Mixer bag found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1403 | 2024-04-04 | 2024-04-05 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
3243 | HA Pre-filtration LOT: EBYHAM004 P/N: GZ5615-10 Sample# 24MS00479. LAL sample container submitted with multiple cracks. | roodline.bertrand@sanofi.com | 68NYA/4700, QC MICRO | 2024-04-03 | 2024-04-08 | Per QU-SOP-0014245 LAL samples must be uncompromised and submitted in a polystyrene container.
3242 | 8NYA - Multiple lots missing BIRF entries | stanley.durand@sanofi.com | 8NYA room 1204 | 2024-03-29 | 2024-04-08 | Per QU-SOP-0074092 all entries for removal and storage for drug substance during transport should filled in BIRFs.
3241 | R&D - Bioburden Sample 24MS00497- TfR-GAA EOH Pool A LOT: EFGC0065. Sample tested on an LAL assay instead of bioburden | roodline.bertrand@sanofi.com | 68NYA/4700, QC MICRO | 2024-04-06 | 2024-04-07 | Per the test request and QU-OPE-0174277 the sample should've been tested for bioburden.
3240 | Incorrect equipment selected in DBS prep record | David.Loik@sanofi.com | 51 NYA | 2024-04-07 | 2024-04-07 | The correct bottle type would be selected
3239 | Thyrogen Q Fraction 8 A280 Sample for L: DGF3B18A ID: 2400003963 unable to be validated without additional material. Results needed to forward process. | sarah.adams@sanofi.com | 45 NYA | 2024-04-06 | 2024-04-06 | Readings obtain be within acceptance criteria
3238 | WORK order # 3174661 Replace Column High Level read switch LSH722.0. | andrew.cao@sanofi.com | building 76 NYA, room 347 | 2024-04-07 | 2024-04-07 | FBL-SOP-000365 | Yes | No Deviation
3237 | Perfusion failure, pH, and weight alarms causing hold on FZ bioreactor X-2210 | brendan.dickhaut@sanofi.com | 8NYA Room 1333, US Ballroom | 2024-04-07 | 2024-04-07 | Per FBL-EWI-000010, CNC operations, Media Changeout, part 3 - complete media change out, step 4; operator is to remove all pinch clamps and ensure a clear flow path for automatic changeover. Expected result is that all clamps are removed.
3236 | 8 NYA - LAL Sample 1115179 - BUFFER_705675 submitted to QCM in polypropylene tube and tested by QCM | Rebecca.Schmidt@sanofi.com | 68NYA/4700/QCM | 2024-04-03 | 2024-04-07 | Per QU-SOP-0070395 Step 6.3.2 Store samples at 2-8C in polystyrene containers. Step 6.3.3 Ensure all samples are received and diluted in polystyrene or dehydrogenated glass tubes prior to sample analysis.
3235 | 8NYA - No second verify provided for manual entries on several records. | bailey.gleason@sanofi.com | 8NYA | 2024-04-07 | 2024-04-07 | Hold gemba to see path forward
3234 | Vi skid did not add acid as intended | Ralston.najarro@sanofi.com | 8NYA1403 | 2024-04-05 | 2024-04-05 | VI shouldve added acid after the sampling step was complete | No | Deviation
3233 | GEMBA 8 NYA Upstream Pompe Allowable Recipe Acid Doses Increased for VI-2865 for Remainder of Train 3 Run Lot EGHAF002 PN 807198 | gary.dodakian3@sanofi.com | 8 nya Suite 2 Upstream 1403 | 2024-04-04 | 2024-04-04 | The allowable number of acid doses in the VI skid recipe should get the mini eluate to the correct range.
3232 | 51NYA Plates arrived after 24 hours | abel.agnes@sanofi.com | 68NYA 4th Floor | 2024-04-04 | 2024-04-04 | EM plates are incubated within 24 hours per MABC-QU-SOP-0070558 attachment 5.
3231 | The clean steam to the header went down at 10:43 PM last night. The reactor(s) lost steam to the agitator last night due to the event. The unit that was in service was the X-9120 CSG2. | stephen.walker@sanofi.com | 76NYA FZ 1st area | 2024-04-04 | 2024-04-05 | steam supply valve not supposed to be closed | Yes | Deviation
3230 | Temperature Gauge F-18482 display not working 76 NYA Fermentation | dianne.moustafa@sanofi.com | 76 NYA/364 | 2024-04-04 | 2024-04-04 | It is expected that the pressure gauge read out on the display
3229 | MCC/VI Column 3 Eluate Issue | bhargav.patel@sanofi.com | 8 NYA / 1403 | 2024-04-04 | 2024-04-04 | Eluate from column 3 collected to VI SUM | Yes | Deviation
3228 | Incorrect Filter Test Selected | craig.valery@sanofi.com | 51 NYA RM 139 | 2024-03-17 | 2024-04-03 | It is expected that filter part number GZ2381-01 is tested under corresponding GZ2381-01 test program per QU-SOP-0072849
3227 | Buffer 705665 1.0M NAOH - Clamp not fully closed, slight leak by compromising 1 of 3 aseptic connections | nicholas.grote@sanofi.com | 1202 | 2024-04-01 | 2024-04-04 | Clamps are aligned on the tubing and fully closed to prevent leaks.
3226 | GEMBA: 51NYA, chipped sight glass | adam.rowe@sanofi.com | 51NYA, 140 | 2024-04-03 | 2024-04-03 | Sightglass should not be damaged
3225 | MCC Skid # 3 X-2831001 UV Transmitter Local Scaling Incorrect | kelley.huang@sanofi.com | 8NYA 1403 | 2024-04-03 | 2024-04-03 | Post-Column UV Transmitters should be scaled 0 - 10 OD for Pompe processing per MABC-QU-SOP-0016987
3224 | Gemba: Settle plate paperwork not control printed for execution | judith.keller@sanofi.com | 68NYA/4700/QCM | 2024-03-25 | 2024-04-01 | It is expected that analytical worksheets where data is recorded are control-printed.
3223 | 8NYA-Pompe- MES rounding result inaccurately during calculation for probe differences on steps that require comparison of local transmitters and offline readings | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-04-03 | 2024-04-03 | it is expected that MES round correctly to ensure ranges are met/ allow forward processing | Yes | Deviation
3222 | SUT not used for buffer changeout when required | zachary.rickard@sanofi.com | 8NYA/1210 | 2024-03-28 | 2024-04-03 | All steps and materials called out in the EWI will be used in the process.
3221 | Pompe - RF1NAE002 - 8NYA - MCC missing VI Skid ID Tag ID | michael.knapp@sanofi.com | 8NYA 1403 | 2024-04-03 | 2024-04-03 | MCC and VI are set to communicate correctly with each other.
3220 | Fz Shipping Bag Load Take out | zachary.rickard@sanofi.com | 8NYA/1313 | 2024-03-28 | 2024-04-03 | Product to be used in a process will be properly released and issues within a record will be addressed at time of occurrence.
3219 | Pompe - RF1NAE002 - 8NYA - Harvest Overflow Pump ModBad Error | michael.knapp@sanofi.com | 8NYA 1403 | 2024-04-03 | 2024-04-03 | Overflow pump operates without alarms
3218 | During the processing of Lot EBYSEP013, resuspension pump suffered a mechanical failure. | Jan.Stepanik@sanofi.com | 76-237 | 2024-04-03 | 2024-04-03 | Pump should not have failed.
3217 | 76 NYA: 2 approved HA Ferm operation worksheets used for manufacturing operations, 12MAR24 | megan.botti-marino@sanofi.com | 76 NYA | 2024-03-12 | 2024-04-03 | It is expected that effective documents be used to execute processes.
3216 | Particles found during routine processing in the Gemco Dryer. | Jan.Stepanik@sanofi.com | 76 NYA / room 227 | 2024-04-03 | 2024-04-03 | No stain should have been present
3215 | Multiple errors, HA Fill, Dynamic EM, L/N: EBYMFL001 | craig.stevenson@sanofi.com | 76NYA, Room 355 | 2024-03-21 | 2024-04-03 | All data to be collected per QU-SOP-0074306
3214 | incorrect time put in for storage of media | nicole.lufkin@sanofi.com | 51 NYA 140 | 2023-12-07 | 2024-04-02 | QU-SOP-0003344
3213 | Record conflicts with EWI instructions | Ralston.najarro@sanofi.com | 8NYA1403 | 2024-04-02 | 2024-04-02 | Status of MCC step should be open so record can move forward
3212 | 8NYA Cell Culture 250L mixer PBM-2547001 Product load adjust Calibration stickers expired | Christian.Venegas@Sanofi.com | 8NYA Cell Culture Phase 1 Room 1333 | 2024-04-02 | 2024-04-02 | To get a Phath forward
3211 | pO2 was out of spec for T5 on E3 | adam.hildebrand@sanofi.com | 51-RM140 | 2024-04-02 | 2024-04-02 | Per MBR the pO2 range should be between 30-131
3210 | Peristaltic Pump (P-2205001) PM Went Overdue. Pump P-2205001 could not be located although 2 x P-2215001 pumps were found | dom.arone@sanofi.com | 8 NYA, RM 1333 | 2024-04-01 | 2024-04-02 | PM's must be completed by Infor assigned due date.
3209 | 8NYA - Date not set for Facilities cleaning, Preventing record closure. | bailey.gleason@sanofi.com | 8NYA | 2024-04-02 | 2024-04-02 | Hold gemba to see path forward | Yes | Deviation
3208 | 8NYA - No second verify provided, Preventing Record closure. | bailey.gleason@sanofi.com | 8NYA | 2024-04-02 | 2024-04-02 | Hold gemba to see path forward
3207 | Two extended due PM's for the month of March went overdue | Barry.Haines@Sanofi.com | 8NYA | 2024-04-01 | 2024-04-01 | That PM's are completed on time and before the due date
3206 | 8NYA Downstream Column Qual. CNC Buffer setup (regarding a small leak on hose of 705661 buffer ) | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1303 | 2024-04-02 | 2024-04-02 | To get a Phath forward to continue with record and CV |  | No Deviation
3205 | pH probe signal dropping out | brendan.dickhaut@sanofi.com | 8NYA room 1403 | 2024-04-01 | 2024-04-01 | pH probe signal is not interrupted. | Yes | No Deviation
3204 | Incorrect Bioburden Incubation for 8 New York Ave Downstream CV Sample of Skid X-3200001 | Rebecca.Schmidt@sanofi.com | QC Micro Bioburden Lab | 2024-03-31 | 2024-04-01 | QU-SOP-0073109 Attachment 1, the incubation temperature must be 30-35C for CV Rinses | Yes | No Deviation
3203 | MCC X2831 status incorrectly changed in MES | Ralston.najarro@sanofi.com | 8NYA1403 | 2024-04-01 | 2024-04-01 | Status of MCC step should be green so record can move forward
3202 | 8NYA Upstream: VI skid Set up GEMBA for EWI update / MES take out and VI manifold set up | amanda.sorel@sanofi.com | 8 NYA 1403 | 2024-03-27 | 2024-03-27 | It is expected that the EWI would give instructions in order of how they should be performed. | Yes | Deviation
3201 | 45-51NYA Thyrogen DGF2B16- Missed Plate Count Verification for rhTSH Formulated Product (sample #: 2400001848, part #: 5157-00, lot: DGF2B16) | larissa.fairweather@sanofi.com | 68 NYA / 4700 / QCM | 2024-03-04 | 2024-03-29 | Per QU-SOP-0019490 Step 6.3.18, Thyrogen Formulated Bulk/Formulated Product (5157-00) plate counts require verification by a second trained analyst with a notation stating the count has been verified.
3200 | Thyrogen- Equipment Prep - incorrect autoclave cycle number recorded on released Carboy Cap. | Jeremy.Wooden@Sanofi.com | Caps were autoclaved and released in room 169 at 45 NYA. | 2024-02-18 | 2024-03-28 | It is expected that the Autoclave tag affixed to the carboy cap would have the correct run number prior to autoclave and release for use in production. | Yes | No Deviation
3199 | 8NYA APU-Room 1323-CC8-Flow Indicators FB-2943-015, FB-2943-016, FB-2943-025, FB-2943-026 did not generate passing calibration data due to a failed pump on UF/DF Skid X-3070001 | robert.bauch@sanofi.com | 8-1323 | 2024-03-29 | 2024-04-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3198 | Harvest/Overflow line for SUB X2240 will not reach drain if set up per FBL-EWI-000011 | Brandon.Steff@sanofi.com | 8RM1403 | 2024-04-01 | 2024-04-01 | Harvest overflow filter will reach HBOD drain
3197 | Pompe - RF1NAE002 - 8NYA - Probe Assembly Takeout Error | michael.knapp@sanofi.com | 8NYA 1403 | 2024-03-31 | 2024-03-31 | Takeouts work.
3196 | 8NYA-Pompe-EGHSD004 Equilibration Media P/N 805012 10L bag found with open clamp and liquid contacting the sterile boundary. | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-03-30 | 2024-03-30 | It is expected that sterile boundaries remain integral by closing the associated pinch clamps
3195 | 8NYA - Fabrazyme - Butyl Skid X3001 Failed CV | stanley.durand@sanofi.com | 8NYA Room 1323 | 2024-03-28 | 2024-03-30 | CV results for swab #3 to be ≤2.4 ppm per QU-FOR-0090360
3194 | 8RM1220 Cold Room Temperature Out of Range | wilfredo.lopez@sanofi.com | 8RM1220 | 2024-03-29 | 2024-03-29 | It is expected that the cold rooms remain within temperature range specified
3193 | Pompe - RF1NAE002 - 8NYA - MCC Buffer Prime Recipe Unable to Advance | michael.knapp@sanofi.com | 8NYA 1403 | 2024-03-30 | 2024-03-30 | Recipe proceeds without interruption.
3192 | 692175 3D Tote Bag W/ Pump - Around Loop, 3000L - Incorrect tubing configuration | nicholas.grote@sanofi.com | 1202 | 2024-03-30 | 2024-03-30 | Bag configuration should match specification from QU-OPE-233697
3191 | 8NYA-Pompe-EGHSD003 CEDEX Sample analysis run with an incorrect Test selection | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-03-30 | 2024-03-30 | It is expected that "NH3B" be selected as a test method per EWI instructions.
3190 | Gemba needed for 8 NYA slurry concentration above target. | kenny.leland@sanofi.com | 8NYA room 1303 | 2024-03-29 | 2024-03-29 | Concentration should be in the range of 40 - 60%.
3189 | MMC clean recipe trying to pull buffer from wrong port. | Juan.bonilla@sanofi.com | 8NYA1503 | 2024-03-29 | 2024-03-29 | Pull buffer from correct port
3188 | GEMBA: 51NYA, missing entry on BIRF form | adam.rowe@sanofi.com | 51NYA, 148 | 2024-03-27 | 2024-03-29 | BIRF form should be filled out for all movements in and out of cold/warm rooms. |  | Deviation
3187 | Buffer 760720 40MM SODIUM PHOSPHATE - Manual Weigh Out Values Don't Align | nicholas.grote@sanofi.com | 1350 | 2024-03-04 | 2024-03-26 | Manual weigh out will be properly documented with correct weights.
3186 | Broken bottle-HA-CMC for Sepramesh lot EBYMES008 | Karen.Hempsted@sanofi.com | 80 NYA | 2024-03-28 | 2024-03-28 | Per 7.6.3.1-If any bottles or shipper are damaged contact QA to determine if an Event needs to be initiated. | Yes | Deviation
3185 | 8NYA-Pompe RF1NAE002/ EGHAF002 EWI-000010 Typographical error in REF1 under Child "SUB Daily Probe Activities/ M800/ New M400 In Process pH Probe Standardization" | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-03-29 | 2024-03-29 | it is expected that the table reference for "No Adjustment required" have a action value of "delta </= 0.02" pH units
3184 | Damaged Eluate bag SUT/Damaged SUT | marcelo.dossantos2@sanofi.com | 8NYA/1333 | 01/26//24 | 2024-01-26 | SUT should pass visual inspection prior to install
3183 | steam leak on MCC Skid X-2831 during SIP | Yves.Antoine@sanofi.com | Room 1403 | 2024-03-28 | 1201 | We expect the recipe to work properly and pressure hold test to detect all leaks prior to steam start.
3182 | 8NYA-GEMBA-Downstream Fz Butyl Column went dirty expired | aidan.mula@sanofi.com | 8RM1323 | 2024-03-28 | 2024-03-28 | Discuss impact at Gemba
3181 | Fz Sub X-2210 pc02 deviation alarm | kurt.marsocci@sanofi.com | phase 1 room 1333 | 2024-03-27 | 2024-03-27 | pcO2 Probe 1 and pcO2 probe 2 should be within 10mmhg +/- | Yes | No Deviation
3180 | 8NYA - RF647E005_001 Needed for shipment Record kicked back by QA | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-03-23 | 2024-03-27 | Hold gemba to see path forward
3179 | 51 NYA CSG X-9120 was reactivated prior to Action Notice release | stephen.walker@sanofi.com | 51 NYA | 2024-03-27 | 2024-03-27 | The CSG would not have been reconnected to the header until the Action Notice was closed and the Change Control was in the proper state. | Yes | Deviation
3178 | Missing Logbook | Robin.Loguidice@sanofi.com | 76 NYA QCBP Lab | 2024-03-11 | 2024-03-27 | logbook could be located
3177 | GEMBA for base leaking from VI skid base barrel elbow | Brandon.Steff@sanofi.com | 8RM 1403 | 2024-03-27 | 2024-03-27 | Base barrel can be primed manually without leaks | No | Deviation
3176 | 8NYA - RF525E011 Nano filter did not pass inspection, HU 1001279679 has a broken air vent | bailey.gleason@sanofi.com | 1350 | 2024-03-27 | 2024-03-27 | order a new part, hold Gemba for defective part
3175 | General Use valves cannot be located | Adriana.Alban@sanofi.com | 51-143 | 2024-03-27 | 2024-03-27 | It is expected that equipment is available for PM |  | Deviation
3174 | Leak observed during nano filtration | pavel.Dobrovolskiy@sanofi.com | 8nya 1323 | 2024-03-27 | 2024-03-27 | Its expected that during nano filtration the filter will not leak.
3173 | During resupension of lot EBYSEP010, an abnormality was noticed on one of the mixer blades. | Jan.stepanik@sanofi.com | 76-237 | 2024-03-26 | 2024-03-26 | No nymber should have been present on the mixer blade.
3172 | GZ95001858 Out of Spec | hanh.do@sanofi.com | 11 Forbes Rd | 2024-03-27 | 2024-03-27 | Passing the result
3171 | 8NYA - RF1Y5D003_001 FIT tested under incorrect parameters | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-09-03 | 2024-03-27 | hold gemba to see path forward
3170 | 8NYA APU-Room 1503-PF8-UV Indicator FB-4023-006 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 8-1503 | 2024-03-26 | 2024-03-27 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
3169 | Nanofilter PRE USE Fit failed integrity test | nicholas.grote@sanofi.com | 1373 | 2024-03-26 | 2024-03-26 | Passing FIT
3168 | Thyrogen - GZ5158-02 - Supplier Notification | Jarad.Kukla@sanofi.com | 11 Forbes | 2024-03-26 | 2024-03-26 | There would be no issue with the manufacturing of this product.
3167 | 8NYA - Fabrazyme - Butyl Column over validated number of cycles | stanley.durand@sanofi.com | 8 NYA room 1303 | 2023-08-08 | 2024-03-26 | Validated number of column cycles to match CST.
3166 | ZUMERROR - Discrepancy with HU Takeout | nicholas.grote@sanofi.com | 1503 and 1350 | 2024-03-20 | 2024-03-26 | SUT should only be consumed into a single record.
3165 | 8NYA-Cerezyme -EGHDT011 - During freeze bag operations, bag #7 on ROSS shelf #20 underfilled. | stanley.durand@sanofi.com | 8 NYA Room 1703 | 2024-03-26 | 2024-03-26 | Per FBL-EWI-000043 it's expected that all freeze bags are filled to spec (7.3-7.7kg)
3164 | Failed Bubble trap visual inspection during swabbing of the MCC skid 2831 | bhargav.patel@sanofi.com | 1403 | 2024-03-26 | 2024-03-26 | System
3163 | UFDF Retentate FIT Launched from incorrect BF with incorrect formula | Nicholas.grote@sanofi.com | 8 NYA Room 1290 | 2024-03-25 | 2024-03-25 | FITs are launched from correct locations within the batch record with the correct formula. | Yes | Deviation
3162 | Pompe - RF1NAE002 - 8NYA - SUB Perfusion Start Prompt arrived early | michael.knapp@sanofi.com | 8NYA 1403 | 2024-03-25 | 2024-03-17 | The Perfusion Start prompt does not appear until the required growth time is completed.
3161 | 51NYA: wrong sample ID in MBR and entered in BGA | adam.rowe@sanofi.com | 51NYA, 152 | 2023-12-31 | 2024-03-25 | Should be written as W125D1
3160 | 8 NYA Locker Room Soap Dispensers Empty | thomas.loik@sanofi.com | N/A | 2024-03-25 | 2024-03-25 | Soap dispensers are filled.
3159 | Supplier Alert - GZ5158-01: FILTER,OMEGA MAXI, 10KD,2SQFT, PES MEMBR | jarad.kukla@sanofi.com | 11 Forbes | 2024-03-25 | 2024-03-25 | There would be no issue with the manufacturing of this product.
3158 | GEMBA - 8 NYA - Leak from top inlet bleed valve of nano filter during equil | adam.keister@sanofi.com | 8 NYA, 1503 | 2024-03-25 | 2024-03-25 | Path Forward
3157 | 8NYA Downstream UF/DF- EGHUP007-Missing sample-Bioburden and endotoxin sample not taken at the end of concentration 2 | MARC.LATTA@SANOFI.COM | 1503 | 2024-03-24 | 2024-03-24 | bioburden and endotoxin sample must be taken at the end of concentration. 2
3156 | MCC X2831 CIP CV rinsate samples collected into incorrect container (Polypropylene vs polystyrene) | Brandon.Steff@sanofi.com | 8RM 1403 | 2024-03-22 | 2024-03-23 | MCC endotoxin CV rinsate samples are collected in Polystyrene LAL tube. | Yes | No Deviation
3155 | 8NYA - RF1C3E007 Nano filter First FIT failed | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-03-22 | 2024-03-22 | hold gemba to see path forward
3154 | Connection point on ATF drain tee seal not properly adhered | Ralston.najarro@sanofi.com | 8NYA1403 | 2024-03-21 | 2024-03-21 | Seal should be adhered so that connection isn't exposed
3153 | No Products - No Lots - 76NYA - 102 Assets Not Deactivated in INFOR (CCR 2021FRACC0018) | tanni.arif@sanofi.com | 76NYA FZ 1st area | 2021-02-21 | 2024-03-21 | FBL-SOP-000852 - decomissioning
3152 | 8NYA - RF555E004 - Record needed prior to shipment on 27MAR24. P/N 714020 BOM status change | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-03-22 | 2024-03-22 | Hold gemba to see path forward |  | Deviation
3151 | 8NYA -GEMBA Request -failed visual inspection on SUT-Vendor deviation needed | richard.manasian@sanofi.com | 1703 at 8NYA | 2024-03-21 | 2024-03-21 | SUT would be free from defects
3150 | Errors in approved MES record for Pompe SUB | nicole.stlawrence@sanofi.com | 8 NYA | 2024-03-15 | 2024-03-22 | It is expected that the approved MES record does not contain these errors. | Yes | Deviation
3149 | 8NYA - RF1NNE002 805015 P/N 714020 cannot be used in this operation. | Bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-03-22 | 2024-03-22 | Hold gemba see path forward | Yes | Deviation
3148 | 8NYA-GEMBA-Downstream Cz Nano-Particulate noted inside pre-nano assembly tubing | aidan.mula@sanofi.com | 8RM1503 | 2024-03-22 | 2024-03-22 | Material should be sterile and free of particulates
3147 | 761221, EGHBV008, Cerezyme Q Membrane Load Filter Tested with Incorrect Parameters | wilfredo.lopez@sanofi.com | 8RM1350 | 2024-03-22 | 2024-03-22 | That the filter recipes are checked against the critical deviation interim control document for correct testing parameters
3146 | 8NYA-GEMBA-Downstream-Cz Freeze HOFF Loading not signed in record | aidan.mula@sanofi.com | 8RM1210 | 2024-03-21 | 2024-03-22 | Determination of impact at Gemba
3145 | 3000L tote bag had a leak during transfer | jermaine.kellam@sanofi.com | 8RM1202 | 2024-03-21 | 2024-03-21 | The bag should not have leaks | Yes | No Deviation
3144 | 8NYA-Pompe- RF1NAE002/ EGHAF002 Drain Tee found with non-integral sterile boundry prior to installation during ATF build | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-03-21 | 2024-03-21 | It is expected that components with sterile boundries remain integral prior to installation
3143 | Media 805013 (EGHMD003) Endotoxin Sample Received Empty - Sample 1080389 | Nicholas.grote@sanofi.com | 1350 | 2024-03-21 | 2024-03-21 | Sample vials are filled with correct sample volume.
3142 | GEMBA - 8 NYA - HOF Failed to Launch | adam.keister@sanofi.com | 8 NYA, CNC | 2024-03-21 | 2024-03-21 | Document and assess delay in HOF launching
3141 | Nano used for RF1C3E006 failed first FIT - related event QE-018859 | nicholas.grote@sanofi.com | 1290 | 2024-03-21 | 2024-03-21 | Max allowable flow is 70 ml/min, test result failed at 75.2 ml/min.
3140 | Balance used prior to LEQ Relocation Qualification report signed; unit within calibration. | sarah.adams@sanofi.com | 45 NYA | 2024-03-19 | 2024-03-20 | Instrument to not be used prior to Qualification release
3139 | GEMBA - 8 NYA - Calculation Issue with CZ Freeze Record | adam.keister@sanofi.com | 8 NYA, 1703 | 2024-03-21 | 2024-03-21 | Path forward with ROSS load parameters, need to have HOF loaded by around 2230 tonight
3138 | 68NYA-Support APU-Room 1740-Validation-F-9240 Humidity Calibrator found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1740 | 2024-03-19 | 2024-03-21 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3137 | 45NYA- Support APU-Facilities-Room 181-F-6155-045 and F-6155-040 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 45-181 | 2024-03-20 | 2024-03-21 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3136 | 8NYA - Fabrazyme - TCU parameters do not align with UR (User requirements specification) - UR F008-ED-UR-0087 | Thomas.Asmar@sanofi.com | 8 NYA - Room 1333 | 2022-10-17 | 2024-03-20 | It is expected that the TCU parameters align with the UR.
3135 | 8NYA-Fz-Thermo 250L SUM bag found to have potential Mfg defect during pre-installation inspection | john.lytle@sanofi.com | Rm 1333, 8NYA | 2024-03-20 | 2024-03-20 | 250L SUM bag needs agitator oriented in correct position to ensure correct deployment in the associated mixer |  | Deviation
3134 | QU-SOP-0068354-LIMS Workflow Operations (v5.0) is approved without document control approval | helma.kuruvilla@sanofi.com | 68NYA | 2024-03-15 | 2024-03-20 | It is expected that Document Control would be included as the final approver for all document revisions per QU-SOP-0074143.
3133 | 76 NYA, During routine EM of the HA fill areas non-viable site 217 selected/sampled twice and no reportable result for site 220 | megan.botti-marino@sanofi.com | 76 NYA, HA fill areas | 2024-02-07 | 2024-03-20 | It is expected that all associated non-viable sites be sampled weekly.
3132 | Media mistakenly put in warm room instead of cold room. | stephen.roberts2@sanofi.com | 51NYA Room 148 | 2024-03-20 | 2024-03-20 | Per batch record media 609616 needs to be warmed for a minimum 48 hours and maximum 5 days. Warming occurred prematurely. 686470 does not need to be warmed at all. |  | Deviation
3131 | Slow Leak From Inlet Bleed Valve on Nano Filter | adam.keister@sanofi.com | 8 NYA, 1503 | 2024-03-20 | 2024-03-20 | Path Forward |  | Deviation
3130 | 51 NYA, approved paperwork for two aseptic challenges printed 08MAR24 and signed as confirmed effective when assay performed on 10MAR24 | megan.botti-marino@sanofi.com | 51 NYA | 2024-03-10 | 2024-03-20 | It is expected that the effective version of documents be used and that the working copy block be signed to verify the paperwork is effective prior to work execution.
3129 | 8NYA Suite 2 Upstream Vortexer for ViCell not Compliantly Onboarded | stephanie.aceto@sanofi.com | 8 NYA, Room 1403 | 2024-03-19 | 2024-03-19 | All equipment being introduced for use in processing of GMP products would follow the onboarding process per QU-SOP-0044843. |  | Deviation
3128 | Unable to generate goods receipt | kevin.ornell@sanofi.com | 80-102 | 2024-03-19 | 2024-03-19 | 5 goods receipt labels would be generated
3127 | pH overshot setpoint step 7.9.13.6 EBYFRM013 | dianne.moustafa@sanofi.com | 76 NYA/364 | 2024-03-19 | 2024-03-19 | Step 7.9.13.6 pH not to overshoot pH setpoint
3126 | Fabrazyme - EGHCE002 - 8NYA - Vicell Viability below 80% | dana.hopkins@sanofi.com | 8RM1333. | 2024-03-15 | 2024-03-15 | That the viability as tested will result in a result > 80%. |  | Deviation
3125 | Activities performed in 8RM1503 while Action Notice work was being performed and before | aidan.mula@sanofi.com | 8RM1503, 8RM1323 | 2024-03-15 | 2024-03-15 | That per the action notice instructions, no open processing will occur. Performing set-up and sampling activities likely violated this restriction. | Yes | No Deviation
3124 | 8NYA-downstream-CZ Q-membrane-EGHBV007-SUM bag pre-inspection failed 2x | lukus.palzkill@gmail.com | 8NYA downstream suite 2 room 1503 | 2024-03-18 | 2024-03-18 | The bags should not have physical defects that could cause a leak or affect the product.
3123 | Loss of 68NYA RODI "storage and distribution" | MA-BioCampus_F&E-Compliance@sanofi.com | 68NYA | 2024-03-18 | 2024-03-18 | It is expected the the Air Compressor remain in service and the RO Storage Tank not dump. | Yes | Deviation
3122 | Thyrogen - Buffer Prep - 45 NYA Rounding error on pH adjustment worksheet for Buffer 1636 | Jeremy.Wooden@Sanofi.com | 45 NYA Room 190 | 2024-03-05 | 2024-03-14 | That the pH titration amount would be rounded correctly |  | Deviation
3121 | 76 NYA PUR EBYHAM002: Tank V-3521 Scale went off after transfer from FERM to PUR. | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2024-03-15 | 2024-03-15 | Per QU-OPE-0434886 (step# 7.7.4.6, 7.7.9 , 7.7.10.4 and 7.7.11), Tank V-3521 should be used for the first precipitation processes such as measuring the amount of alcohol to be added as well as to measure the amount of supernatant to be decanted. | Yes | Deviation
3120 | 68NYA-Support APU-Room 4700-QCM68-F-15179 Thermometer found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-4700 | 2023-12-21 | 2024-03-18 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3119 | 8NYA APU-Room 1363-PF8-EMS Temperature Indicator FB-2969-108 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 8-1363 | 2024-03-04 | 2024-03-18 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3118 | Pompe - RF1NAE002 - 8NYA - Discoloration on probe aseptiquik sterile boundary. | michael.knapp@sanofi.com | 8NYA 1403 | 2024-03-17 | 2024-03-17 | No discoloration on membrane.
3117 | 8NYA - Fabrazyme - EGHUP006 Final yield calculation below normal range. | stanley.durand@sanofi.com | 8 NYA rm 1323 | 2024-03-17 | 2024-03-17 | Following the input of the A280 results from QC the yield calculation should be within normal range.
3116 | Fabrazyme - EGHCE002 - 8NYA - SUB Harvest Overflow Pump Failure | michael.knapp@sanofi.com | 8NYA 1333 | 2024-03-16 | 2024-03-16 | overflow pump operates as intended |  | No Deviation
3115 | Fabrazyme - EGHCE002 - 8NYA - Vicell Viability below 80% | michael.knapp@sanofi.com | 8NYA 1333 | 2024-03-17 | 2024-03-17 | Daily viability >80% |  | Deviation
3114 | 8NYA-No Associated Product- Rm1343 Differential Pressure alarms/ Room status check | john.lytle@sanofi.com | Rm 1343 8NYA | 2024-03-13 | 2024-03-15 | It is expected that differemntial pressure be maintained and corrective action be taken to remediate root cause |  | Deviation
3113 | A Gemba is needed for: RF1BRE018 Butyl Load filter tested with incorrect recipe. | jim.ramjattan@sanofi.com | 8NYA, Room 1290 | 2024-03-16 | 2024-03-16 | That filter recipes are checked against the critical deviation interim control document for correct testing parameters. |  | Deviation
3112 | 8NYA - Cerezyme Lot EGHDT010 last ROSS bag underfilled during freeze bag filling operations. | stanley.durand@sanofi.com | 8 NYA Room 1703 | 2024-03-16 | 2024-03-16 | It is expected the CZ ross fill is weighed to 7.3-7.7kg per the batch record. | Yes | Deviation
3111 | Fabrazyme - EGHCE002 - 8NYA - Vicell Viability below 80% for second day | michael.knapp@sanofi.com | 8NYA 1333 | 2024-03-16 | 2024-03-16 | Daily viability >80% |  | Deviation
3110 | Tempterature excursion at 51NYA RM88 (UF/DF) | cindy.guerra@sanofi.com | 51NYA, RM 88 | 2024-03-15 | 2024-03-15 | That the cold room will be controlled within the expected parameters with no excursions.
3109 | 45 NYA - RO Pipe Burst on PreTreatment Skid | kevin.southwick@sanofi.com | 45 NYA | 2024-03-15 | 2023-12-15 | It is expected that the pipes on the 45 NYA RO/DI Pretreatment Skid not burst and continue operating.
3108 | SUB 2210 low viability | jeffrey.boldizar@sanofi.com | 8NYA 1333 | 2024-03-15 | 2024-03-15 | Vi-Cell viability to be above 80 Percent per MABC-QU-OPE-0524531-Control Strategy for Fabrazyme Integrated Continuous Biomanufacturing
3107 | Ross Shipper (F008-SHP-100) qualification protocols do not include RQ enrollment section per procedure. | aveer.shah@sanofi.com | 8 NYA | 2022-08-17 | 2024-03-12 | QU-OPE-0523549 page 6 - Equipment Description: Chamber
3106 | No sample in sample vial for Sample ID 1107439 - EGHUP006 | sarah.adams@sanofi.com | 68 NYA, QCC | 2024-03-14 | 2024-03-14 | Sample to be filled. |  | No Deviation
3105 | Daily sampling record not update as intended for Pompe | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2024-03-14 | 2024-03-14 | Record should update with EWI alignment for sampling for post ATF
3104 | Cz MMC Pooling Filter Clog | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2024-03-13 | 2024-03-13 | Path forward to continue Load Pooling process. Filter typically handles 3 eluates with no issue. Tonight's pooling was only 2 eluates
3103 | HA-CMC LOT: EBYCMC044 P/N: 2139-32. Bioburden and retain vials were cracked for pour plate testing. | roodline.bertrand@sanofi.com | 68NYA/4700, QC MICRO | 2024-03-05 | 2024-03-12 | QU-SOP-0052170 | Yes | No Deviation
3102 | GEMBA for MCC X2510 Held due to columns out of sync | Brandon.Steff@sanofi.com | 8RM 1333 | 2024-03-13 | 2024-03-13 | MCC SUM is installed correctly, and global hold timer is not exceeded. |  | Deviation
3101 | 761219, EGHBR005 Cerezyme Zinc Buffer Out of Temperature Range | wilfredo.lopez@sanofi.com | 8RM1503 | 2024-03-12 | 2024-03-12 | It is expected that the buffer is chilled after batching and transport, prior to the start of chromatography operations so that it is ready and within temperature range for the process.
3100 | 45NYA-Support APU, QCC45- Room 1130, Micro balance F-17180 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 45-1130 | 2024-03-11 | 2024-03-13 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3099 | 8NYA APU-Room 1403-CC8-Bioreactor #1 BR-2220001-Temperature Indicators FB-4095-029 and FB-4095-032 found out of tolerance during a calibration | robert.bauch@sanofi.com | 8-1403 | 2024-03-12 | 2024-03-13 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
3098 | Disposition/reconciliation not completed for FB BIRFs during pooling | craig.valery@sanofi.com | 51 NYA RM87 | 2024-03-10 | 2024-03-12 | it is expected that BIRFs are completed for disposition and reconciliation at the time the intermediate is dispositioned/reconciled.
3097 | 12 Biweekly EM sits not sampled | craig.stevenson@sanofi.com | 76 New York Ave, Multiple Rooms | 2024-03-09 | 2024-03-12 | It is expetced that the assocated rooms are sampled for EM on a biweekly frequency. | No | Deviation
3096 | Fz zinc column cycle discrepancy | zachary.rickard@sanofi.com | 8NYA/1323 | 2020-10-08 | 2024-03-11 | The correct amount of cycles will be calculated. |  | Deviation
3095 | 8NYA-FAC Work Orders 2992870 and 2985930 not completed for Kuhner Incubators (Due 11Mar24) | john.lytle@sanofi.com | Rm 1343 8NYA | 2024-03-11 | 2024-03-12 | It is expected that the Incubators get the associated preventative maintenance according to the calibration program/ schedule |  | No Deviation
3094 | HA bottles TOR greater than 48 hours | valeria.moreno@sanofi.com | 76 NYA, room 360 | 2024-03-08 | 2024-03-11 | QU-OPE-0039925
3093 | 8NYA Fabrazyme - EGHSE006 - During SP pooling operations all 3 butyls pooled at same time. | stanley.durand@sanofi.com | 8 NYA Room 1323 | 2024-03-07 | 2024-03-09 | QU-WIN-0002455-Chromatography Operations (Per FBL-EWI-000033) "Perform Eluate Pooling" one butyl to be pooled at a time.
3092 | X-3350 pH found out of specification during chromatogram review | craig.valery@sanofi.com | 51 NYA RM 139 | 2024-03-10 | 2024-03-12 | it is expected that the pH remains with in specified range |  | Deviation
3091 | Incorrect Recipe run for autoclave hose load | jesse.demers@sanofi.com | 45RM169 | 2024-03-11 | 2024-03-12 | Expected recipe is porous load |  | Deviation
3090 | 8NYA-Fz-RF503E002 X2210 pCO2 Probe Divergence | john.lytle@sanofi.com | Rm 1333, 8NYA | 2024-03-12 | 2024-03-12 | Per MABC-QU-WIN-0000088-Skid In-Line Analytical Instruments, "if a probe is diverging towards an alarm limit or has triggered an alarm, assess the alarm per FBL-EWI-000004" Alarm Response. EWI-000004 Instructs the operator to have MSAT trend the divergence for a GEMBA and based on the data, disable the diverging probe per applicable procedure |  | No Deviation
3089 | Multiple Products - Multiple Lots - 8NYA - MCC SIP timer configured incorrectly in MES | michael.knapp@sanofi.com | 8NYA 1403, 1333 | 2024-02-15 | 2024-03-11 | Steam expiry timer is configured per the PQ documents. |  | Deviation
3088 | 8NYA - Cerezyme - EFGDT007 - Leak observed from pump pressure sensor membrane during shipping bag transfer. | stanley.durand@sanofi.com | 8NYA Room 1703 | 2024-03-09 | 2024-03-09 | No leaks should be observed to be coming from pressure sensor membrane of pump tubing. | No | Deviation
3087 | 8NYA-Gemba-Pompe Seed Media-805012-Formulation SFO missing 714047 | alejandro.martinez2@sanofi.com | 8RM1350 | 2024-03-11 | 2024-03-11 | MES BOM updates are underway
3086 | Broken thermometer in QCM | kathryn.lentoni@sanofi.com | 68NYA 4700, QCM | 2024-03-10 | 2024-03-10 | Thermometer not broken and functions properly
3085 | 8NYA CZ EGHBY005- Sample rinse volume flagging in OneLIMS for incorrect rinse even though rinse volume is correct per QU-SOP-0073109. | larissa.fairweather@sanofi.com | 68 NYA / 4700 / QCM | 2024-03-11 | 2024-03-11 | Per QU-SOP-0073109 Attachment 2, sample type CZ2_CM_761222_LOAD_BIOB should be tested 10mL of sample volume followed by a 1x100mL Fluid A rinse.
3084 | Tote auto step failure in MES | jeffrey.boldizar@sanofi.com | 8NYA CNC | 2024-03-10 | 2024-03-11 | Auto step to move the record in the correct direction. |  | No Deviation
3083 | 8NYA-Pompe-Cedex memory storage full, with no defined procedure to free up space for additional sample analysis | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-03-11 | 2024-03-11 | It is expected that instructions are provided and proceduralized for the removal of sample results from Cedex memory to allow for further use in production activities
3082 | 8NYA-DOWNSTREAM-Cz-Zinc-EGHBR005-skid setup-UV high out of range | MARC.LATTA@SANOFI.COM | 1503 | 2024-03-11 | 2024-03-11 | UV HIGH VALUE SHOULD BE READING THE HIGH VALUE INDICATED IN GAGE INSITE AT PLUS OR MINOUS 2%
3081 | Unverfiable source for EM Settle plate incubations | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2024-03-09 | 2024-03-10 | Per QU-SOP-0073597 Step 6.6.6 the incubation date and time must be recorded properly.
3080 | GEMBA for Cz Eluate SUM too full to mix | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2024-03-10 | 2024-03-10 | Bag deploys correctly and has sufficient head space for mixing.
3079 | Barrel of Growth phase media PN# 609616 was not warmed for the minimum 48 hours prior to use. | adam.hildebrand@sanofi.com | 51NYA RM 140 | 2024-03-09 | 2024-03-09 | It is expected the minimum warming time for media would be greater than or equal to 48 hours.
3078 | 8 NYA - GEMBA FZ UFDF probes not installed during equlibration | alex.tello@sanofi.com | 8NYA RM1523 | 2024-03-09 | 2024-03-09 | Probes to be installed and read PH and Condo
3077 | 8NYA-ReGemba-Downstream Cz Zinc- Buffer leaked at sample bag due to loose cap and open pinch clamp | aidan.mula@sanofi.com | 8NYA/RM1212 | 2024-03-08 | 2024-03-08 | Re-gemba to capture event. | Yes | Deviation
3076 | 8NYA-Gemba-Downstream Cz- Weld failure on manifold | aidan.mula@sanofi.com | 8NYA/ RM1703 | 2024-03-09 | 2024-03-09 | Determination of impact at Gemba | Yes | Deviation
3075 | Leak observed on Cz Zinc load SUM bag post Dillution | Edgar.vale@sanofi.com | 8NYA, room 1503 | 2024-03-09 | 2024-03-09 | Bag weld would be intact and leak would not occur | Yes | Deviation
3074 | Pre-Filtration Osmo result OOS for Fz Media | william.lee@sanofi.com | 8RM1350 | 2024-03-09 | 2024-03-09 | That the osmo result would be within process control range.
3073 | 2 8 NYA Aseptic Challenge Samples not Meeting Acceptance Criteria for Turbidity | jessica.kychler@sanofi.com | QC Micro 68 New York Avenue, floor 4, room 4700 | 2024-03-07 | 2024-03-07 | It is expected that aseptic challenge samples will be free from turbidity/particles
3072 | Thyrogen - Buffer Prep - 45 NYA DP alarm between room 190 and room 188 on 06MAR24. | Jeremy.Wooden@Sanofi.com | 45 NYA rooms 190 (Buffer Prep suite) and 188 (MSAT suite) | 2024-03-06 | 2024-03-08 | There would not be any DP alarms between rooms 190 and 188. | Yes | Deviation
3071 | Reconciliation required between FBL-EWI-000005 (Facility Disinfection) and MABC-QU-SOP-00730029 (Preparation of Disinfection and Cleaning Solutions) | christopher.arnold@sanofi.com | 8NYA | 2024-03-06 | 2024-03-06 | That the guiding documents would be aligned for acceptable cleaning solutions. That exceptions made into EBRs would contain the details to demonstrate compliant operation.
3070 | Fabrazyme - EGHCE001 - 8NYA - Low Low Temperature Alarm on SUB X2200 | michael.knapp@sanofi.com | 8NYA 1403 | 2024-03-07 | 2024-03-07 | Temperature is maintained and glycol levels do not fluctuate within TCU during operation.
3069 | HA and HA-CMC – Multiple Lots – 76 NYA Multiple Rooms – Multiple AHUs shutdown due to momentary fluctuation in the supply of electrical power | juan.ramos-torres@sanofi.com | 76 NYA | 2024-03-07 | 2024-03-07 | AHU should run at all times
3068 | X-3350 pH found out of specification during chromatogram review | craig.valery@sanofi.com | 51 NYA RM 139 | 2024-03-06 | 2024-03-07 | it is expected that the pH remains with in specified range
3067 | N2 Bulk cannot be received by QC; unable to verify that the tanker was sealed. | sarah.adams@sanofi.com | 80 NYA | 2024-03-07 | 2024-03-07 | Tanker is sealed upon arrival, Analyst witness the seal removal by driver.
3066 | Operator observed water intrusion from ceiling in room 76-229, MFG Purification Room. | jan.stepanik@sanofi.com | 76-229 | 2024-03-07 | 2024-03-07 | No water should have been dripping from the ceiling. | Yes | Deviation
3065 | 8NYAWFI Loop Used Power Surge Used Post Power Surge Pre Sani | thomas.loik@sanofi.com | 8RM1350 | 2024-03-07 | 2024-03-07 | Utilize Endo testing to keep solution | Yes | Deviation
3064 | HiHi temp QCA on D2 bioreactor | David.Loik@sanofi.com | 51 NYA | 2024-03-07 | 2024-03-07 | Pump would reset without intervention following outage. |  | Deviation
3063 | Continuous Non-Viable Data not collected during HA Fill Setup for CP006 and CP007 | craig.stevenson@sanofi.com | 76NYA, Room 355 | 2024-02-27 | 2024-03-07 | It is expected that CP006 and CP007 have continuous non-viable results during the duration of the set up, besides the hopper transfer into the filling table.
3062 | Sitewide Power Interruption | tanni.arif@sanofi.com | sitewide | 2024-03-07 | 2024-03-07 | Campus expected to ride through power outage
3061 | No reportable result for Sample ID 915694 Cerezyme ICB Butyl A Eluate DGHBN021 (HL, BL) | sarah.adams@sanofi.com | 68 NYA, QCC | 2024-02-26 | 2024-03-06 | Reportable results
3060 | EGHSE006, 762731 Fabrazyme SP Leak on Pump Tubing During Eluate Pooling | Kelvin.glean@sanofi.com | 8 NYA, 1423 | 2024-03-07 | 2024-03-07 | No air bubbles in the lines after priming and during operation. No defects on the SUT.
3059 | Cz, Fz, Pompe Multiple Lots_8NYA_QU-WIN-0002454 V.13,( Seed Train) LFH diagram title is Media Equilibration | nicole.stlawrence@sanofi.com | 8 NYA/Small Scale | 2024-03-06 | 2024-03-06 | It is expected QU-WIN-0002454 V.13 step 12.6.3 diagram title should be specific to all operations or there should be a diagram for each operation.
3058 | 3K SUM Leak when making 705673 - 10 mM Phosphate, 45% Ethylene Glycol (only WFI had been added when leak occured) | nicholas.grote@sanofi.com | 1350 | 2024-03-06 | 2024-03-06 | Bag remains integral and no leak occurs. | Yes | Deviation
3057 | No Product - No Lots - 8NYA - Two GMP AHU PM WO Extensions Missing QA Approval | MA-BioCampus_F&E-Compliance@sanofi.com | AHUs feeding 8NYA (rooms 1111 and 2141) | 2024-01-25 | 2024-03-06 | FBL-SOP-000081 Section 6.5.5 (required signatories) | Yes | Deviation
3056 | Pompe PPQ1 EGHSD001 Low viability | Jacob.Roman@sanofi.com | 8RM1403 | 2024-03-06 | 2024-03-06 | Per QU-OPE-0522556, the viability criteria of greater than or equal to 85% must be meet in less than or equal to 7 days.
3055 | 8NYA-Cz-X2560 MCC recipe held due to low Harvest material weight, X-2220 SUB recipe found to be in hold due to ATF B Harvest pressure interlock | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-03-05 | 2024-03-06 | It ias expected that Harvest material is collected into a secondary vessel with no restrictioons in order to supply product for continuous feed through MCC
3054 | Refractometer testing returned an OOS result for 705675 (10% Polysorbate, 3% TBP) | nicholas.grote@sanofi.com | 8RM1350 (Solution Prep) | 2024-03-06 | 2024-03-06 | That the material will be within all process control parameters.
3053 | Membrane deallocation issue | pavel.Dobrovolskiy@sanofi.com | 8nya 1323 | 2024-03-05 | 2024-03-05 | Its expected that at the end of each run the record will deallocate any equipment associated with the run
3052 | Cz record RF639D024 CNC SUB shop floor order will not terminate | dana.hopkins@sanofi.com | 8NYA room 1403 | 2024-02-06 | 2024-03-04 | Record should be fully executed before attempting to terminate
3051 | Low Viability Pompe PPQ1 | phoebe.karkos@sanofi.com | 8NYA 1403 | 2024-03-05 | 2024-03-05 | Viablitity should be >80% on day 6 (today is day 6), Scale up must be done by EOD Day 7 with viability >85%
3050 | Possible ATF fouling | Ralston.najarro@sanofi.com | 8NYA1403 | 2024-03-05 | 2024-03-05 | There shouldn't be a ring of cells on the ATF |  | Deviation
3049 | pH Sampling of buffer returned out of specification result | christopher.arnold@sanofi.com | 8RM1350 | 2024-03-04 | 2024-03-04 | The expected result is that the pH will be in range when tested. | Yes | Deviation
3048 | WO Inaccurately Closed | tanni.arif@sanofi.com | INFOR | 2018-06-01 | 2024-02-21 | WO PM checklist steps checked off if truly complete | Yes | Deviation
3047 | SP1 DGF4B22C Eluate found outside of the cold room at 51NYA. | sylvester.kegbeh@sanofi.com | 51NYA | 2024-03-03 | 2024-03-04 | Per QU-OPE-0079501 and QU-SOP-0074092 product should be stored at 6-10C. TOR per QU-SOP-0074092 is 58 minutes
3046 | Buffer 705681 leaked during transfer | chris.brewster@genzyme.com | Room 1202 at 8 NYA | 2024-03-03 | 2024-03-03 | It is expected that bags will not leak during transfer. | Yes | Deviation
3045 | 8NYA-Cz- Bleed pump not controlling properly after pressure spike | john.lytle@sanofi.com | Rm 1343 8NYA | 2024-03-01 | 2024-03-03 | It is expected that the bleed pump activate in order to control capacatance according to its automation
3044 | Tolerance Violation exception during weigh and dispense | nicholas.grote@sanofi.com | 1364 | 2024-03-01 | 2024-03-02 | Bags should be weighed out within the calibration range of the scale.
3043 | Cold Room 1220 Monthly Clean Expired | nihcolas.grote@sanofi.com | 8 NYA 1220 | 2024-03-02 | 2024-03-02 | Room status is cleaning ok.
3042 | SUB 2210 CO2 Set Dev Alarm | jeffrey.boldizar@sanofi.com | 1333 | 2024-03-02 | 2024-03-02 | CO2 probes to be within 20mmHg of each other | Yes | No Deviation
3041 | Sample was submitted with insufficient volume for Postnano drug substance PART: 728285, LOT: EGHDT004 | anastacia.hachem@sanofi.com | 4700 QCM | 2024-03-01 | 2024-03-01 | Per QU-OPE-0329679 step 9.2.2.5, use FBL-SOP-000019 to determine sample volume and media required for each sample to be tested | Yes | Deviation
3040 | Late PMs for month of February | kevin.southwick@sanofi.com | 8 NYA | 2024-03-01 | 2024-03-01 | SOP MABC-QU-SOP-0021310-Maintenance Operations, section 6.4.8. | No | Deviation
3039 | 51 NYA Gemba-Incorrect filter serial number recorded in MES FIT Record | jacob.kallon@sanofi.com | 51 NYA Room 139 | 2024-02-24 | 2024-04-01 | QU-SOP-0072849 |  | Deviation
3038 | 8NYA FZ Capto Butyl Eluate P/N 762729 EGHBE012 - Only 1 retain submitted instead of 2 | abel.agnes@sanofi.com | 68NYA 4th Floor | 2024-02-29 | 2024-03-01 | Two bags of 20mL retains are submitted for this process step with each sample only having one label affixed.
3037 | 8NYA APU-rooms 1503 and 1333- 5 items overdue for calibration | robert.bauch@sanofi.com | 8-1503, 8-1333 | 2024-03-01 | 2024-03-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before Replacement
3036 | No nonviable data output from equipment (Climet) during FB DGF2B16 Process at 51NYA | sylvester.kegbeh@sanofi.com | 51NYA | 2024-02-28 | 2024-02-28 | Per QU-SOP-0074750 step 6.7.7, nonviable data should be parsed directly into the system. |  | Deviation
3035 | Record RF639D024 hung up on Tote Removal for Buffer 760722, Tote ID: BHT1720016. | Yves.Antoine@sanofi.com | Room 1403 | 2024-02-29 | 2024-02-29 | To complete tote Disposal/removal
3034 | 20" filters were used for FZ ufdf instead of 30" per BQ-FRM-000753 | pavel.Dobrovolskiy@sanofi.com | 8nya 1323 | 2024-02-28 | 2024-02-28 | Its expected that EWIs and MES record instruct operator to connect the correct filters. |  | Deviation
3033 | FZ MCC skid x-2510, SUM PBM2520001 Eluate bag seems to be deflated | kurt.marsocci@sanofi.com | 8 NYA Phase 1 | 2024-02-28 | 2024-02-28 | Bag should not be this deflated when filling |  | Deviation
3032 | HA Lot - EBYFRM010; leak observed after media transfer to V-3020 | Charlie.Houle@sanofi.com | 76 NYA; Room 364 (HA Fermentation) | 2024-02-27 | 2024-02-27 | No leak should occur
3031 | 8NYA - Expired Media Spill in CNC about 10L 760741 | bailey.gleason@sanofi.com | 8NYA | 2024-02-28 | 2024-02-28 | Hold Gemba See path forward for cleaning |  | No Deviation
3030 | 8NYA - Gemba - Pompe - 805012 - Seed Media Inventory/MBR Version# Clarification | alejandro.martinez2@sanofi.com | 8RM1350 | 2024-02-22 | 2024-02-22 | Confirm only two 500mL bottles are needed for Pompe USP ops with MSAT/USP team. If so, MES to update BOM only calling for 1 x 692303, when taken out of BOM, MES should auto-generate 2 HU's for each bottle needed.
3029 | 80 NYA Warehouse - Winter Temperature Mapping Failure | regan.aghdam@sanofi.com | 80 NYA Warehouse - 149 Packing & Shipping, Corridor 156, Quarantine Cage 138, Acid/Corrosive Storage 145 | 2024-02-20 | 2024-02-27 | DFCOP-QU-OPE-0858653; 20C to 25C with allowable excursions ≤ 24 hours of 15C to 30C provided the mean kinetic temperature does not exceed 25C. | Yes | Deviation
3028 | 8NYA CZ DGHBN024 Myco/Viral Not Received in OneLIMS | abel.agnes@sanofi.com | 68NYA 4th Floor | 2024-02-27 | 2024-02-27 | These sample types are stored in -60C within 24hrs.
3027 | Atypical TSH Harvest ELISA Results for H12 (Lot EGF3B01) & H2 (Lot EGF1B02) | beth.grafton@sanofi.com | 68 NYA QCM | 2024-02-26 | 2024-02-27 | Harvest titer generally increases in a relatively linear shape.
3026 | Pompe - RF1N9E001 - 8NYA - Onelims EM start time not captured at start of LFH setup | michael.knapp@sanofi.com | 8NYA 1343 | 2024-02-27 | 2024-02-27 | EM sampling is initiated in onelims to capture setup start time. instructions to perform exist within the EWI. | Yes | No Deviation
3025 | Post sanitization WFI flush during system cleaning/storage - flowrate outside low limit UFB Lot: DGF2B16 | craig.valery@sanofi.com | 51NYA, Rm 88 | 2024-02-24 | 2024-02-27 | It is expected flow rate is within specified limits |  | Deviation
3024 | Equipment - 8NYA - Room 1403 Refrigerator removed from service in 2021 and not returned to service | michael.knapp@sanofi.com | 8NYA 1403 | 2021-03-29 | 2024-02-27 | The refrigerator is in service in EAM. |  | Deviation
3023 | 8NYA - Fabrazyme SP - EGHSE005 - Leaked observed from bag pH probe during standardization. | stanley.durand@sanofi.com | 8NYA Room 1323 | 2024-02-26 | 2024-02-26 | Per QU-WIN-0019199 "Bag Inspection" section 4.6 no leak to observed during bag inspection |  | Deviation
3022 | Cerezyme -RF639D024 - 8NYA - MCC Recipe unable to end due to Load Pump ModBad failure | michael.knapp@sanofi.com | 8NYA 1403 | 2024-02-26 | 2024-02-26 | Recipe end of run completes successfully. | Yes | No Deviation
3021 | Cerezyme - RF639D024 and RF639E001 (shared buffer) - 8NYA - Leak identified on new buffer install (Not In Use) | michael.knapp@sanofi.com | 8NYA 1210 | 2024-02-26 | 2024-02-26 | There are no leaks on the tubing following tube welding. | Yes | Deviation
3020 | TSH-Buffer bag 3 of buffer fill GZ1410-53 was outside tolerance of +/- 1% of target weight of 50.050kg. | kevin.ornell@sanofi.com | 45 NYA room 190 | 2024-02-25 | 2024-02-25 | Bag would be filled within tolerance. | Yes | No Deviation
3019 | 8NYA - RF555D008 Buffer created stock for 3114.313L when it should have been 2800.00 | bailey.gleason@sanofi.com | 8NYA | 2023-09-26 | 2024-02-25 | hold gemba to see path forward
3018 | 8NYA GEMBA - CZ pump tubing leak | alex.tello@sanofi.com | 8NYA RM1703 | 2024-02-24 | 2024-02-24 | no leak | No | Deviation
3017 | Crack observed on inside of sight glass of D1 reactor at 51NYA. | nicholas.shoenfelt@sanofi.com | 51-140 | 2024-02-22 | 2024-02-22 | It is expected that the sight glass surface be integral. |  | Deviation
3016 | cold room 147 temp alarm 10.09 | Nicole.lufkin@sanofi.com | 51NYA 151 | 2024-01-31 | 2024-02-23 | temp not above 10 |  | No Deviation
3015 | 51NYA Room 129A- Thyrogen APU- final rinse conductivity Indicator F-17195-020 on the parts washer was found out of tolerance during an unscheduled calibration. | robert.bauch@sanofi.com | 51-129A | 2024-02-22 | 2024-02-23 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
3014 | 68NYA - No reportable results for 7 Weekly CU Conductivity results | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2023-10-17 | 2024-02-22 | It is expected that all Critical Utility samples be tested after they are collected.
3013 | Analyst performed one assay occasion without training being signed off by a manager | paul.consolmagno@sanofi.com | 76NYA | 2024-02-20 | 2024-02-22 | The training paperwork would be signed by a managerial representative prior to assay execution.
3012 | Doc Control not added as a Final Authorizer per MABC-QU-SOP-0074143 | karolina.markle@sanofi.com | 68NYA, no lab involved | 2024-02-19 | 2024-02-22 | N/A
3011 | Incorrect formulation BOM 805012 Pompe media. | jack.laplante@sanofi.com | 8NYA RM1350 | 2024-02-22 | 2024-02-22 | Looking for path foward.
3010 | Moisture pool discovered in eluate agitator shield | jeffrey.boldizar@sanofi.com | 1333 | 2024-02-22 | 2024-02-22 | No leak should be present in enclosed system
3009 | 8NYA-Fz- RF503E002 Daily Sampling results (pCO2) exceed expected values | john.lytle@sanofi.com | Rm 1333, 8NYA | 2024-02-22 | 2024-02-22 | EWI instructions are clear on >/=20mmHg discrepency between BGA pCO2 results and local pCO2 transmitters, stating that the BGA should be recalibrated, the SUB resampled, and the re-sample results being used to perform an in process calibration of the associated local transmitters. This was not performed as the trend of conflicting adjustments made operators elevate results |  | Deviation
3008 | 8NYA-Fz-X2210 TCU not controlling SUB temp | john.lytle@sanofi.com | Rm 1333, 8NYA | 2024-02-21 | 2024-02-21 | It is expected that the SUB temp remain 36-38deg |  | Deviation
3007 | Atypical pH Trend for Q Equilibration step. | jacob.kallon@sanofi.com | 51 NYA, Room 132 | 2024-02-13 | 2024-02-18 | It is expected that the pH trend is within expected range per QU-SOP-0061107 |  | Deviation
3006 | 8NYA-X2210-RF503E002 SUB X-2210 pH adjusted without additional sampling performed per EWI | john.lytle@sanofi.com | Rm 1333, 8NYA | 2024-02-21 | 2024-02-21 | It is expected that a second sample is pulled to verify accuracy of 1st sample, prior to performing a 1 point Standardization on the SUB probes |  | Deviation
3005 | Multiple Products - Multiple Lots - 8NYA - Cryostat "Control Failed" EMS alarm | michael.knapp@sanofi.com | 8NYA 1280 | 2024-02-20 | 2024-02-21 | No alarms generated during operation |  | No Deviation
3004 | 8NYA FZ Transfer Filter was found broken upon inspection. | daniela.medina2@sanofi.com | 8NYA / 1350. | 2024-02-20 | 2024-02-20 | filter expected to pass visual inspection |  | No Deviation
3003 | Media POU sample was sampled in biolims but never received by QC | Brandon.Steff@sanofi.com | 8RM1333 8NYA | 2024-02-15 | 2024-02-15 | Biolims samples are samples in biolims, dropped off, picked up by MLCs for Transport and received by QC for testing
3002 | RTU at 51 NYA exceeded tier 2 AN allowable 3 hour downtime. | Alex.Biskaduros@Sanofi.com | 51 NYA RTU 12 affecting rooms 129, 129A, 141, 142 | 2024-02-20 | 2024-02-20 | tier 2 HVAC work to be completed within a 3-hour time frame.
3001 | A Witnessed step was not executed during Blue DGF2B20 Fraction pooling | sylvester.kegbeh@sanofi.com | 51NYA | 2024-02-04 | 2024-02-20 | Per QU-OPE-0107276, it is expected that the Witnessed steps be executed at the at the time of performance |  | Deviation
3000 | Blue DGF3B21 fractions pooled was stored without a BIRF. | sylvester.kegbeh@sanofi.com | 51NYA | 2024-02-08 | 2024-02-20 | It is expected that product be stored per QU-SOP-0074092
2999 | 8NYA - no Pompe buffers referenced in FBL-EWI-000022 for PH / Condo testing | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-02-20 | 2024-02-20 | Hold gemba to see path forward |  | No Deviation
2998 | 8NYA APU-CC8-FB-4784 UV easy cal found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8NYA | 2024-02-05 | 2024-02-20 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
2997 | Media observed leaking from colder apparatus connection during bioreactor D1 Transition refeed. | james.letourneux@sanofi.com | 51NYA140 | 2024-02-19 | 2024-02-19 | Sanitary connections should be secure and not allow egress or ingress of material at the connection.
2996 | 8NYA - RF621E030 Bioburden sample leaked upon arrival to QC lab | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-02-19 | 2024-02-19 | Hold gemba to see path forward
2995 | 8NYA - RF573E004 Record will not progress to allow for transfer. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-02-19 | 2024-02-19 | Hold gemba to see path forward |  | No Deviation
2994 | Cerezyme - RF639E001 - 8NYA - Retain samples delivered in 1ea single use sample bag | michael.knapp@sanofi.com | 8NYA 1403 | 2024-02-18 | 2024-02-18 | Samples are delivered in their own individual sample bags.
2993 | 8NYA-Fz-X2200 Production Media P/N 705691 Top tote drain to bottom tote during production use | john.lytle@sanofi.com | 8NYA CNC Space Rm1240 | 2024-02-17 | 2024-02-17 | It is expected that tote setup is complete and executed per EWI, therefore supply to the SUB would be accomplished using one media tote at a time to meet daily perfusion rate criteria
2992 | Tote bag leak on 760725 - 20 mM Glycylglycine, 27% Ethylene Glycol, pH 8.0 | nicholas.grote@sanofi.com | 1202 | 2024-02-17 | 2024-02-17 | Tote bag deploys correctly, no leak of buffer from tote. |  | Deviation
2991 | Media transfer line bad seal after creating media in solution prep | nicholas.grote@sanofi.com | 1202 | 2024-02-16 | 2024-02-16 | good seal, no leaking |  | No Deviation
2990 | 8NYA-Fz-RF503E002- ATF A Harvest Line has a small leak between the harvest guard filter and the associated bleed bag | john.lytle@sanofi.com | Rm 1333 | 2024-02-16 | 2024-02-16 | It is expected that the system remain integral during processing
2989 | TempTale discrepancy for Vesi Virus shipment #168276 to Geel | Adriana.Alban@sanofi.com | Geel,BE | 2023-12-12 | 2023-12-12 | It is expected that the temptale maintain monitoring until manually stopped
2988 | MCC X2543 bubble trap bottom sensor reading incorrectly | Brandon.Steff@sanofi.com | 8RM1333 8NYA | 2024-02-15 | 2024-02-15 | Both bubble trap sensors read "Air" until bubble trap is issued fill command, after which both sensors should read "liquid". "Air" reading should not cause recipe to hold prior to issuing "Fill" command. | Yes | Deviation
2987 | RF600D030 had no buffer consumption | dana.hopkins@sanofi.com | 8NYA room 1403 | 2023-12-06 | 2024-02-14 | Record should consume correct quantity of buffer
2986 | AHU 1 shutdown | patrick.hosegood@sanofi.com | 51 NYA | 2024-02-15 | 2024-02-15 | Unit stayed running
2985 | RTU1 air handling unit at 51NYA shutdown unexpectedly | nicholas.shoenfelt@sanofi.com | 51-125, 51-127, 51-122A, 51-122B | 2024-02-15 | 2024-02-15 | It is expected that the RTU does not trip during Hepa testing.
2984 | Missing Cell Culture Falcon Batch Record (Lot BCP008, Document number FBL-MBR-000124 | Amy.McIntyre@sanofi.com | 68 New York Avenue, room 1825 | 2024-01-09 | 2024-01-09 | Expected results are that the batch record would be in the RM reconciliation bin for RM to pick up and sign for in the issuance log sheet, per steps 6.8.2 through 6.8.6 of QU-SOP-0044980. Quarterly |  | Deviation
2983 | Missing Cerezyme Buffer (DGHAD022) Endotoxin Sample | ana.salichs@sanofi.com | 8NYA | 2023-11-05 | 2024-02-15 | Samples are received and tested by QC on time.
2982 | 8NYA CZ Capto Butyl Eluate P/N 761217 DGHBN024 - Only 1 retain submitted instead of 2 | abel.agnes@sanofi.com | 68NYA 4th Floor | 2024-02-13 | 2024-02-13 | 2 x 20 mL retains are received with one label on each per table 10 and 11 in CZ Sampling Plan MABC-QU-OPE-0044689. |  | Deviation
2981 | Cerezyme - RF639D024 - 8NYA - Incorrect Sampling Times recorded in OneLIMs | michael.knapp@sanofi.com | 8NYA 1403 | 2024-02-13 | 2024-02-13 | OneLIMS sampling is documented accurately. |  | No Deviation
2980 | Cerezyme - RF639D024 - 8NYA - Disable Capacitance Probe 2 | michael.knapp@sanofi.com | 8NYA 1403 | 2024-02-13 | 2024-02-13 | Capacitance probe is consistent and accurate.
2979 | Cerezyme - 705663 - EGHBN006 - 8NYA - RF557E006 - Weigh out bag outside of calibration range for scale | alejandro.martinez2@sanofi.com | 8NYA/8RM1364 | 2024-02-06 | 2024-02-14 | MES should stop operator from weighing out outside of range.
2978 | 2 QCA Trend Review PM's not completed | dom.arone@sanofi.com | 8 NYA | 2023-10-31 | 2024-02-06 | PM's must be completed, and completed by Infor assigned due date | Yes | Deviation
2977 | Training assay unable to be located to complete review. | sarah.adams@sanofi.com | 68 NYA | 2023-08-21 | 2024-02-12 | Training assay able to be reviewed
2976 | Fabrazyme - EGHCE002 - 8NYA - ATF A failure | michael.knapp@sanofi.com | 8NYA 1333 | 2024-02-12 | 2024-02-12 | ATFs remain operational, there are no diaphragm ruptures. | Yes | Deviation
2975 | 8NYA: EGHZE003 - Fabrazyme Zinc Eluate Samples Taken Prematurely | wilfredo.lopez@sanofi.com | 8RM1323 | 2024-02-13 | 2024-02-13 | That the steps are performed as they become active in MES. The samples should have been taken when the record prompted operators to do so. It is expected that the steps in the record are in the correct order, the steps become active as the operation progresses, and that the steps are aligned with the DeltaV operations.
2974 | 76 NYA PUR EBYHAH001: Missing Performed by Signature on step 7.11.18 of QU-OPE-0202445 | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2024-02-08 | 2024-02-09 | Per QU-SOP-0074684, steps 6.1.1 and 6.1.2 and per the ALCOA+ attributes, all spaces must be filled, and all signatures and entries are to be made sequentially at the time of execution. The signer performs the signatures prior to completing the next step in the operation.
2973 | RO#1 tank lid was found half way opened. | jesus.acosta2@sanofi.com | 51 NYA. Room 146 mechanical room | 2023-12-19 | 2023-12-19 | RO#1 tank lid should've been closed all the way.
2972 | TSH - EP - Cycle count on 45 Autoclave reset and is repeating cycle numbers already used in the past. | Jeremy.Wooden@Sanofi.com | 45 NYA room 169 | 2024-02-10 | 2024-02-10 | Cycle counts should not reset and repeat cycles already used. This is a unique ID that should continue to advance by 1 each time a cycle is performed.
2971 | visual inspection of qual skid failed X3201 during swabbing | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 1 Room 1303 | 2024-02-10 | 2024-02-10 | To get a Path forward to continue with record | No | Deviation
2970 | Manifold not attached to tote for part 832327 | dennis.bazinet@sanofi.com | 8NYA CNC area | 2024-02-10 | 2024-01-22 | Manifold should of been welded to bag right after tote was filled by SP | Yes | Deviation
2969 | SUB X2210 DO stabilization failure | Brandon.Steff@sanofi.com | 8RM1333 8NYA | 2024-02-10 | 2024-02-10 | DO probes hold steading within +/- 0.2% for 2 hours | Yes | No Deviation
2968 | 692169 SUT Failure - RF552E002 | leo.gillespie@sanofi.com | 1350 - Solution prep | 2024-02-09 | 2024-02-09 | Expect to follow EWI instructions to successfully connect recirculation line--but unable to do so | Yes | No Deviation
2967 | TSH - 45 NYA cold room 174 is in alarm for high temperature | Jeremy.Wooden@Sanofi.com | 45 NYA room 174 | 2024-02-09 | 2024-02-09 | That the temperature of the cold room would be in range.
2966 | Media Storage movement not recorded for 1 barrel of growth media | nicholas.shoenfelt@sanofi.com | 51-109 | 2024-02-04 | 2024-02-09 | It is expected that the movement into the cold room is document on the portable media storage log. |  | Deviation
2965 | 76NYA R&D - CIM - Dynamic EM plate 938687 (Site:TRR-MAS) no reportable results | craig.stevenson@sanofi.com | 76NYA, Room 390 | 2023-11-18 | 2024-01-25 | It is expected that if plates are submitted to QC Micro, the be accounted for when being received, counted and generate results.
2964 | Version 4.0 of QU-SOP-0071120 approved with incorrect versioning in revision history table (missing 4.0) | Amy.McIntyre@sanofi.com | QualiPSO CMS | 2024-02-06 | 2024-02-08 | Expected results: Version 4.0 would a row in the revision history table for the changes made in version 4.0. | Yes | Deviation
2963 | PBM-4220001 N2 Overlay Not Functioning Properly | adam.keister@sanofi.com | 8 NYA, 1503 | 2024-02-07 | 2024-02-08 | For N2 Overlay to function as it should |  | Deviation
2962 | 8NYA-Cz-Eluate SUM setup complete but not signed for at time of completion | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-02-01 | 2024-02-07 | It is expected that batch record steps are signed for as complete contemporaneously with the work performed |  | Deviation
2961 | TOC outage from 76NYA the EMS signal 76NYA_F076-AI-9626-026 | MA-BioCampus_F&E-Compliance@sanofi.com | 76NYA Room 347 | 2024-01-30 | 2024-02-07 | EMS data is expected to be received and stored in PI
2960 | MABC-QU-OPE-0152178 does not include post-nano SUM bag for the extractable / Leachable calculation | gregory.klotz@sanofi.com | 8NYA | NA | NA | Expected that the bag would be included in the calculation |  | Deviation
2959 | Cerezyme - RF639D024 - 8NYA - Eluate SUM unable to be mixed for sampling | michael.knapp@sanofi.com | 8NYA 1403 | 2024-02-07 | 2024-02-07 | SUM is installed such that it can be properly mixed for eluate sampling.
2958 | 68NYA Support APU- Room 4735-QC-F-16123-001 no as found data during a scheduled calibration | robert.bauch@sanofi.com | 68-4735 | 2024-02-06 | 2024-02-07 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2957 | 8NYA-Gemba-FZ- Seed Media Transfer BOM (CNC) needs to be updated | alejandro.martinez2@sanofi.com | 8NYA/CNC-Solution Fill Area | 2024-02-06 | 2024-02-06 | BOM should only have parts required, no extra parts, which can lead to confusion |  | No Deviation
2956 | High pressure gas mix out alarm on 12L Wave Bioreactor | Adriana.Alban@sanofi.com | 152B 51NYA | 2024-02-06 | 2024-02-06 | The filter heater to evaporate the condensation so as not to build up.
2955 | Gemba-8NYA-Downstream Fz- 2 records without steps signed at time of performance | aidan.mula@sanofi.com | 8RM1323 | 2024-02-06 | 2024-02-06 | Steps should be performed at time of operation. |  | Deviation
2954 | Broken Sartorius Pro Mixers X-4201007, X-4201008 to be Sent Off Site for Repair | michael.pagonis@sanofi.com | 8NYA: CNC and DS2 Room 1403 | 2023-10-20 | 2024-01-22 | Mixers to function as expected. To be able to connect to SUM bags and DeltaV without issues |  | Deviation
2953 | GEMBA for Fz missing Athena transactions and SUT take outs | Brandon.Steff@sanofi.com | 8RM1333 8NYA | 2024-01-21 | 2024-02-05 | Media tote take outs work correctly in MES | Yes | Deviation
2952 | Kickback Record: RF603E003 - Alternate Part Number 669716 was not issued to Process Order | alejandro.martinez2@sanofi.com | 8NYA/8RM1364 | 2024-02-01 | 2024-02-05 | Path-forward on charging non-BOM material in athena for QA to approve
2951 | 8NYA GEMBA - Damaged Nano filtration filters. | alex.tello@sanofi.com | 8NYA RM1503 | 2024-01-26 | 2024-01-26 | Pass visual inspection | Yes | Deviation
2950 | 3 Consecutive failures on exhaust filter F-15854 from E4, DGF4B22. See attached BP test and Pressure drop test which was prompted during teh 3rd consecutive test. | adam.hildebrand@sanofi.com | 51-140 | 2024-02-05 | 2024-02-05 | It is expected Post FIT's pass after the run has completed.
2949 | 3 Consecutive failures on exhaust filter F-15854 from E4, DGF4B22 | adam.hildebrand@sanofi.com | 51-140 | 2024-02-05 | 2024-02-05 | It is expected Post FIT's pass after the run has completed.
2948 | 8NYA GEMBA- CZ Shipping bags failed visual inpection | alex.tello@sanofi.com | 8NYA RM1503 | 2024-02-04 | 2024-02-04 | Pass visual inspection
2947 | X-2210 went into hold for a HI HI alarm during the inflation recipe | kurt.marsocci@sanofi.com | 8 NYA Phase 1 Room 1333 | 2024-02-05 | 2024-02-05 | Fix the HI HI alarm parameter
2946 | Process Temperature for TSK DGF2B16A and Blue DGF3B21pH is identified to be out of range during chromatogram review from PI | sylvester.kegbeh@sanofi.com | 51NYA | 2023-07-11 | 2024-02-04 | It is expected that the Temperature and pH should be in range per QU-SOP-0065572 and QU-SOP-0025632
2945 | Gemba – 8NYA – Filter assembly found to be damaged upon inspection-FZ | william.lee@sanofi.com | 8NYA/1290 | 2024-02-03 | 2024-02-03 | Filter bleed should not be damaged with crack
2944 | Seprafilm - 76NYA RM233 - EBYSEP007 - Seprafilm lot with atypical attribute | dawn.lam@sanofi.com | 76 nya Room 233 | 2024-02-05 | 2024-02-01 | According to QU-SOP-0072890 Die Cut and Inspection of Seprafilm® at 76 NYA Version 18.0. films should not contain foreign particles (>500 µm), holes (>500 µm), cracks/tears/nicks (>2 mm), fibers (>100 µm), cosmetic scratch and tape impressions.
2943 | Gas mix to Wave Bioreactor E reading 0.00 and cannot clear alarm & hold condition. | james.letourneux@sanofi.com | 51 NYA Room 152B | 2024-02-04 | 2024-02-05 | Gas flow in should be maintained at setpoint.
2942 | 760727 Buffer tote not recirculating. | robert.morawski@sanofi.com | 8NYA CNC | 2024-02-04 | 2024-02-04 | The tote to proper recirculate once the pump and glycol valve are opened.
2941 | 8NYA - Fabrazyme UF/DF - DGHUP022 - End of Hold BIOB Sample not received by QCM | stanley.durand@sanofi.com | 8NYA Room 1323 | 2023-11-14 | TBD | Per QU-WIN-0046060 |  | No Deviation
2940 | EBR Pointing to non-existent child in EWI-21 for FIT | nicholas.grote@sanofi.com | 1290 | 2024-02-02 | 2024-02-02 | EBR points to correct instruction. | Yes | No Deviation
2939 | Seprafilm - 76NYA RM233 - EBYSEP006 - Seprafilm lot with atypical attribute | dawn.lam@sanofi.com | 76 nya Room 233 | 2024-02-01 | 2024-02-01 | According to QU-SOP-0072890 Die Cut and Inspection of Seprafilm® at 76 NYA Version 18.0. films should not contain foreign particles (>500 µm), holes (>500 µm), cracks/tears/nicks (>2 mm), fibers (>100 µm), cosmetic scratch and tape impressions.
2938 | Unable to locate 3 lots of FZ ICB UF/DF for annual inspection | abel.agnes@sanofi.com | 68NYA | 2024-02-01 | 2024-02-01 | Annual inspection is completed on all reserves by 31JAN.
2937 | A 250L Eluate SUM bag, Part Number 692168 failed visual inspection. | kurt.marsocci@sanofi.com | 8 NYA Upstream Phase 2 | 2024-02-01 | 2024-02-01 | Supplier deviation and get a new bag to install
2936 | Missing Clarification DGF3B11B Bioburden Sample at 51NYA | sylvester.kegbeh@sanofi.com | 68NYA | 2023-08-08 | 2024-01-26 | It is expected that the sample be submitted by MFG and tested by QC.
2935 | No Products - No Lots - 76 NYA, QC Biopolymers - Two Controlled Worksheet Stamp Numbers Unaccounted for in the QC Biopolymers Laboratory | paul.consolmagno@sanofi.com | 76NYA | 2024-02-01 | 2024-01-01 | Per QU-SOP-0071640, stamps are controlled, tracked, and reconciled in the QC Biopolymers lab
2934 | 8 NYA: No zero filter data for operations collection batch 202305969-1 | megan.botti-marino@sanofi.com | 8 NYA | 2023-12-30 | 2024-02-01 | Per step 6.4 of MABC-QU-SOP-0030277, "Perform zero filter test per procedure only once prior to the start of the first operation per shift." | Yes | Deviation
2933 | Blue Pool Eluate DGF4B19 without BIRF was transfer from cold room 138 to 131 at 51NYA | sylvester.kegbeh@sanofi.com | 51NYA 138 and 131 | 2024-01-31 | 2024-01-31 | FBL-SOP-000371, QU-OPE-0107276 and FBL-SOP-000828 | Yes | Deviation
2932 | 8NYA APU-Room 1403 - Dept CC8- 8 items overdue for calibration | robert.bauch@sanofi.com | 8-1403 | 2024-02-01 | 2024-02-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date |  | Deviation
2931 | Gemba: Brucella Agar GZ95000506 24GN00168 stored at incorrect temperature in NBO/80NYA before delivery to QC | judith.keller@sanofi.com | 80NYA/ NBO / QCM 68NYA 4700 | 2024-01-29 | 2024-01-31 | Per SPC, Brucella agar should be stored at CRT.
2930 | One (1) Late Preventative Maintenance (PM) Work Order Not Completed by Extended Due Date of 31Jan2024 (WO 2493859 + WO 2493859) | MA-BioCampus_F&E-Compliance@sanofi.com | 68NYA room 1310 | 2024-02-01 | 2024-02-01 | Complete all PM tasks prior to PM closure (otherwise approved by QA) or a deviation is required per site specific SOP.
2929 | warm room temp alarm 148 at 32.96 | Nicole.lufkin@sanofi.com | 51NYA 148 | 2024-01-31 | 2024-02-01 | temp not below 36 |  | Deviation
2928 | 51NYA RO1 Low Level Tank Alarm Not Sending Signal to BMS | tanni.arif@sanofi.com | n/a | 2024-01-30 | 2024-01-16 | low level alarm is triggered BEFORE system goes into interlock to allow facilities an opportunity to respond
2927 | 8NYA - 3000L SUM bag with incorrect pressure sensor connection. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-01-31 | 2024-01-31 | N/A |  | Deviation
2926 | X-2220 Bleed Pump throwing pressure alarms | dana.hopkins@sanofi.com | 8NYA room 1403 | 2024-01-31 | 2024-01-31 | Bleed pressure should not throw nuisance alarms
2925 | MCC X2510 going into hold due EM not being able to be acquired | Ralston.najarro@sanofi.com | 8NYA1333 | 2024-01-31 | 2024-01-31 | Recipe should acquire EM |  | No Deviation
2924 | Cz MMC column slurry concentration above target range | zachary.rickard@sanofi.com | 8NYA/1303 | 2024-01-30 | 2024-01-30 | Slurry concentration will fall within range of 40-60% |  | Deviation
2923 | 8NYA APU-Room 1333 Dept CC8-FB-4773-047 and FB-4773-048 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1333 | 2024-01-30 | 2024-01-30 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
2922 | 8NYA Fabrazyme - LAL sample 1015765-FZ2_ZN_762732_ELUATE_ENDO submitted to QCM in polypropylene tube. | larissa.fairweather@sanofi.com | 68NYA/4700/QCM | 2024-01-29 | 2024-01-30 | Per QU-SOP-0070395 Step 6.3.2 Store samples at 2-8C in polystyrene containers. Step 6.3.3 Ensure all samples are received and diluted in polystyrene or dehydrogenated glass tubes prior to sample analysis.
2921 | Gemba-8NYA- CZ-Buffer 832327; Part number missing in Cz sampling plan | alejandro.martinez2@sanofi.com | 8NYA/8RM1350 | 2024-01-29 | 2024-01-29 | Add buffer part number (832327) into Cz sampling plan FBL-RPT-000538. | Yes | No Deviation
2920 | 51 NYA Thyrogen Purification-Blue process pH found out of range during Chromatogram review per QU-SOP-0061107 | jacob.kallon@sanofi.com | 51 NYA Room 138 | 2023-12-31 | 2024-01-29 | It is expected that pH trend remains in specification during all process per QU-SOP-0061107. | Yes | No Deviation
2919 | Thymus Collection Box | Tom.Stomski@sanofi.com | TSA Area | 2024-01-05 | 2024-01-08 | Can not have dulpicate lots numbers
2918 | GEMBA 51NYA: Wave F LowLow temp | adam.rowe@sanofi.com | 51NYA/152 | 2024-01-29 | 2024-01-29 | The wave should not gone into hold, but even if it did, it should not stop controlling critical parameters |  | Deviation
2917 | Gemba needed for 8 NYA MMC column cleaning. Wrong recipe loaded. | kenny.leland@sanofi.com | 8 NYA/ Room 1303 | 2024-01-28 | 2024-01-29 | Correct recipe should have been loaded.
2916 | Hose failure during transfer | maria.palmar@sanofi.com | 8 New York Ave | 2024-01-28 | 2024-01-28 | the hose to not get broken
2915 | 2 Endotoxin samples Cerezyme CM Column Eluate and CM Column Load 761222 EGHBY003 submitted in polypropylene tube rather than polystyrene tube | Stefanie.Erickson@sanofi.com | QC Micro LAL bench | 2024-01-24 | 2024-01-27 | for the samples to be submitted in polystyrene tubes
2914 | WFI Tote not connected to the correct port in CNC area | Edgar.vale@sanofi.com | 8NYA, room 1503 | 2024-01-27 | 2024-01-27 | WFI would be connected to the correct port
2913 | Aseptic connection was not taken at inflation moment | maria.palmar@sanofi.com | 8 NYA | 2024-01-27 | 2024-01-27 | To push back one step on DeltaV and be able to re do inflation
2912 | EWI for Sampling plan is not correct | maria.palmar@sanofi.com | 8 New York Ave | 2024-01-27 | 2024-01-27 | Per FBL-EWI-000009 we will need this syringe
2911 | 8NYA - RF613E002 - SUM set up SFO has P/N 713972 as a takeout step. This is not compatible. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-01-27 | 2024-01-27 | Hold Gemba to see path forward. |  | No Deviation
2910 | Gemba-8NYA-Downstream Cz Phenyl CIP Rinse Prompt was left unanswered, WFI tote emptied | aidan.mula@sanofi.com | 8RM1503 | 2024-01-27 | 2024-01-27 | Path forward on cleaning recipe |  | No Deviation
2909 | SIP on X2510( follow up to QE-006340 | marcelo.dossantos2@sanofi.com | 8NYA/1333 | 01/26//24 | 2024-01-26 | Option to re-SIP without going into clean expiry. |  | No Deviation
2908 | 8NYA - RF1AVE001 - EWI does not specify how to take sample in Collect sample Step in formulation. | bailey.gleason@sanofi.com | 8NYA | 2024-01-27 | 2024-01-27 | Hold Gemba to see path forward.
2907 | 8NYA-SUT Failure - 692174 3000 L tote bag- HU 1001257743 - Batch23GN01773 | leo.gillespie@sanofi.com | 1350 - Solution prep | 2024-01-26 | 2024-01-26 | Expect to find zero defects during inspection
2906 | Two Nano filters broken during Setup | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1703 | 2023-12-21 | 2023-12-26 | To take accountability for Nano filter due to a shortage
2905 | Around 1345 an operator noticed MCC skid 2560 was leaking from the HIBOD filter on the inlet side. There was around 3 litters that leaked onto the floor. The recipe was placed into hold and the HIBOD waste was cleaned up with IPA and Sporklenz. | kurt.marsocci@sanofi.com | 8 NYA Upstream Phase 2 | 2024-01-26 | 2024-01-26 | FBl-EWI-000032
2904 | SIP of X-2510 had DeltaV recipe stuck overnight. Meet do determine Sterile status | dana.hopkins@sanofi.com | 8NYA room 1333 | 2024-01-25 | 2024-01-26 | SIP cycle should run without pauses
2903 | Thyrogen - 51NYA - Lots TBD - EN-285 Original Results Incorrect INFOR Documentation | tanni.arif@sanofi.com | n/a | 2024-01-26 | 2024-01-16 | WO closes would reference WO for CM followups, upload documents, not not passing, etc
2902 | 8NYA GEMBA - CZ Nanofiltration pump alarm during Equilibration | alex.tello@sanofi.com | 8NYA RM1503 | 2024-01-14 | 2024-01-14 | Finish Equilibration without alarms
2901 | Wrong sterilization number recorded for carboy use | james.letourneux@sanofi.com | 51NYA Room 140; 51NYA Autoclave room 141 | 2023-09-26 | 2024-01-26 | Correct SFO should be recorded for autoclaved equipment.
2900 | 76NYA R&D - CIM - EM not performed in room 390 during October 2023 | craig.stevenson@sanofi.com | 76NYA, Room 390 | 2023-10-31 | 2024-01-26 | It is expected that Grade D rooms are sampled at a Monthly frequency for EM
2899 | Gemba - 705691 Fz Media RF573D033 showing two different expiries for SU's in MES | Nicholas.grote@sanofi.com | 1350/1443 | 2023-11-21 | 2024-01-25 | Expiries align
2898 | Seprafilm - 76NYA RM233 - EBYSEP004B - Seprafilm lot with atypical attribute | juan.ramos-torres@sanofi.com | 76 NYA Room 233 | 2024-01-25 | 2024-01-24 | According to QU-SOP-0072890 Die Cut and Inspection of Seprafilm® at 76 NYA Version 18.0. films should not contain foreign particles (>500 µm), holes (>500 µm), cracks/tears/nicks (>2 mm), fibers (>100 µm), cosmetic scratch and tape impressions.
2897 | GEMBA - 8 NYA - MCC butyl load bag foaming through top | adam.keister@sanofi.com | 8 NYA, 1333 | 2024-01-25 | 2024-01-25 | Path forward, next step is mix time then filtered transfer | Yes | Deviation
2896 | Leaking sample bag on E4 H4 | David.Loik@sanofi.com | 51 NYA | 2024-01-25 | 2024-01-25 | Sample bag would not be leaking. |  | No Deviation
2895 | 8NYA-Fz-X2200 hold for Hi Hi Weight Controller output prior to switching from top to bottom Media supply totes | john.lytle@sanofi.com | Rm 1333, 8NYA | 2024-01-25 | 2024-01-25 | it is expected that combined pump speeds do not exceed 54rpm per output alarm limit | Yes | No Deviation
2894 | Incorrect Buffer listed on BOM | kyle.tesorero@sanofi.com | 8NYA | 2024-01-25 | 2024-01-25 | Correct buffer 832327 should have been revised on CIP record when Production records were revised. |  | Deviation
2893 | 8NYA - Cerezyme Nano - EGHDS001 - Additional Mannitol buffer used during equil not captured in MES or SAP | stanley.durand@sanofi.com | 8NYA Rm 1503 | 2024-01-17 | 2024-01-25 | All materials should be captured in MES and SAP per QU-SOP-0019707
2892 | Clean Hold Time Documented Outside of Clean Hold Range | craig.valery@sanofi.com | 51 NYA | 2024-01-13 | 2024-01-25 | it is expected that clean hold time is within range
2891 | Phantom Inventory write-off due to BOM and Athena transaction discrepancies | elissa.karol@sanofi.com | 8NYA | 2023-08-31 | 2023-08-31 | Materials consumed in production and non-production activities are transacted in Athena in real time and costs charged to the activity.
2890 | RF545E004 - Amber clamp on tote bag not tightened, small volume of buffer leaked by out of tote | Nicholas.grote@sanofi.com | 1202 | 2024-01-24 | 2024-01-25 | Clamps fully tightened, no leak by occurs. | Yes | Deviation
2889 | Cz reserve/retain samples incorrectly labeled | zachary.rickard@sanofi.com | 68NYA | 2024-01-23 | 2024-01-24 | Samples will be labeled correctly
2888 | Several shippers p/n GZ8017-01 batch 23MN00588 were found sealed with 3" tape. The SPC calls for 2" tape. | alan.rousseau@sanofi.com | room 1017 74 NYA | 2024-01-24 | 2024-01-24 | Per SPC QU-OPE-0421321 shipper 8017-01 bottom seams should be sealed with 2" tape not 3 " tape.
2887 | 45 NYA Room 174 Magnehelic Gauge DP Reading Zero | jeff.hill@sanofi.com | 45 NYA Room in 174 to 186 (Cold Room) | 2024-01-24 | 2024-01-24 | QU-FOR-0091683, Attachment 1, DP must be above 0.02.
2886 | 8NYA APU-Room 1503 Pressure Indicator FB-4778-002 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 8-1503 | 2024-01-24 | 2024-01-24 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2885 | GEMBA 51NYA: pH reading on deltaV does not match sample on Wave E | adam.rowe@sanofi.com | 51NYA/152 | 2024-01-24 | 2024-01-24 | pH should be in range and should reflect sample results.
2884 | 51 nya, RO1 and RO2 loop was breach | andrew.cao@sanofi.com | BUILDING 51 NYA UTILITIES ROOM | 2024-01-23 | 2024-01-23 | FBL-SOP-000845 | Yes | Deviation
2883 | SUB 2220 max out pH loop out for air sparge | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2023-11-30 | 2023-12-01 | Air sparge should not be max out
2882 | TSH 45NYA: Nosebleed in RM 190 | cindy.guerra@sanofi.com | 45NYA Room 190 | 2024-01-23 | 2024-01-23 | Operator would not have nosebleed in MFG area |  | Deviation
2881 | 8NYA - RF542E001 - P/N 692169 3000L SUT leak during initial WFI. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-01-23 | 2024-01-23 | Do loss of SUT, abort recipe in delta V, change out Sum set up SUT, re launch recipe. Hold gemba about leaked bag
2880 | Gemba – 8NYA – Media totes left in ambient room temperature | william.lee@sanofi.com | 8NYA CNC | 2024-01-22 | 2024-01-23 | 3 x 1K totes should have been transferred to cold room
2879 | 68NYA Support APU-Room 1310- F-15586-041, F-15586-043, F-15586-045, F-15586-047 had no as found data collect during a scheduled calibration | robert.bauch@sanofi.com | 68-1310 | 2024-01-22 | 2024-01-23 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2878 | 68NYA Support APU-Room 1830-Metrology Pressure Standard S-1660 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2024-01-04 | 2024-01-23 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2877 | Media disposal was started as non shared but it was a shared media | dennis.bazinet@sanofi.com | 8NYA CNC area | 2024-01-22 | 2024-01-22 | Batch record would automatically advance but we had to do manual transactions to advance record
2876 | Seprafilm - 76NYA RM233 - EBYSEP003B - particles found in lot | fatima.fragaalvarez@sanofi.com | 76 NYA RM233 | 2024-01-20 | 2024-01-22 | QU-SOP--0011577 STEP 6.4
2875 | TSK bubble trap pressure relief valve found leaking during rinse step | jacob.kallon@sanofi.com | 51NYA, Room 132 | 2024-01-22 | 2024-01-22 | it is expected that the PRV does not leak within normal operating pressure range (</=25 psig) |  | Deviation
2874 | GEMBA 51NYA: Leaking media | adam.rowe@sanofi.com | 51NYA/152 | 2024-01-21 | 2024-01-21 | Bag should not be leaking
2873 | 51NYA: scale missing from list of selectable units for drop checking | adam.rowe@sanofi.com | 51NYA/152B | 2024-01-20 | 2024-01-20 | The scale should be selectable
2872 | 8NYA-Cz-MCC X2731 Column 1 remediation | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-01-20 | 2024-01-20 | It is expected that air is bled out of the MCC system prior to loading over a column, as air has a negative impact on the column functionality
2871 | 45 NYA Room 174 Magnehelic Gauge DP Reading Zero | jeff.hill@sanofi.com | 45 NYA Room in 174 to 186 (Cold Room) | 2024-01-20 | 2024-01-20 | QU-FOR-0091683, Attachment 1, DP must be above 0.02.
2870 | 8NYA-X2200 Base Pump leak during Production Activities | john.lytle@sanofi.com | Rm 1333, 8NYA | 2024-01-20 | 2024-01-20 | it is expected that the base line is unrestricted and has a free flow path to SUB for pH control
2869 | 8NYA-Cz-Perfusion Failure Train 3 | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-01-20 | 2024-01-20 | It is expected that ATFs maintain their operational integrity during perfusion activities |  | Deviation
2868 | Low pCO2 BGA reading on T4, E4 | David.Loik@sanofi.com | 51 NYA | 2024-01-20 | 2024-01-20 | pCO2 value would be within set range. |  | No Deviation
2867 | 45 NYA Air Handling Issue DP's Reading Zero | jeff.hill@sanofi.com | 45 NYA Room in 186C, 186E, 186F (Buffer Prep Hallway) | 2024-01-20 | 2024-01-20 | QU-FOR-0091683, Attachment 1, DP must be above 0.04.
2866 | Harvest SUM PBM2716001 on Train 2 SUB X-2230 has a leak | caleb.coppola@sanofi.com | 8RM1403 | 2024-01-20 | 2024-01-20 | bag doest not leak
2865 | 8NYA-Fz-Tear in Aseptikwik X connection during ATF B installation | john.lytle@sanofi.com | Rm 1333, 8NYA | 2024-01-19 | 2024-01-19 | When removing the sterile connection tab of an Aseptikwik X connector, it is expected that any damage to the membrane be assessed according to criteria listed in MABC-QU-WIN-0046672-Single Use Tubing |  | No Deviation
2864 | GEMBA for UPLC sample mislabeled/missing | Brandon.Steff@sanofi.com | 8RM1333 8NYA | 2023-12-03 | 2024-01-18 | UPLC samples are taken for SUB, Harvest and Eluate SUM each eluate day
2863 | Observed higher than usual pressure during SP2 load C-3370 | craig.valery@sanofi.com | 51NYA RM132 | 2023-12-10 | 2023-12-27 | normal operational pressure durning load is lower at a higher flow rate |  | No Deviation
2862 | 8NYA-Fz- defect found in ATF B permeate line tubing set prior to use in production | john.lytle@sanofi.com | Rm 1333, 8NYA | 2024-01-19 | 2024-01-19 | It is expected that the "plastic" protective lining over the intended sterile boundary line remain integral during prior to installation
2861 | Personnel observed entering room 129A without secondary gowning at 51 NYA> | nicholas.shoenfelt@sanofi.com | 51-129A | 2024-01-19 | 2024-01-19 | It is expected the secondary gowning will be doned and then removed and bagged prior to leaving room 51-129A per QU-SOP-0006952.
2860 | 8NYA Upstream FZ X-2200 EGHCE001 Low Side MFC-019 not Functioning as Expected Post Inoculation | gary.dodakian3@sanofi.com | Suite 1 Upstream 1333 | 2023-12-18 | 2024-01-18 | DO levels are maintained by the MFC post inoculation
2859 | Thyrogen Formulated bulk for EGFPB001 LAL sample unable to be located | judith.keller@sanofi.com | 68 NYA 4700 QCM | 2024-01-15 | 2024-01-18 | It is expected that samples are accounted for and tested.
2858 | 51NYA GEMBA: PVLL in warm room 148 | adam.rowe@sanofi.com | 51NYA/148 | 2024-01-14 | 2024-01-15 | Temp should be stable at setpoint
2857 | Multiple Products- Mutiple Lots- 8 NYA-Merck Potential Recall (CCNS-1005645) GMID 690579 - PMIX0200L1710547 200L Bag Assembly, Mixing | Brenda.Robillard@sanofi.com | 8NYA - potential other manufacturing areas | 2021-09-01 | 2024-01-08 | TC adapter fitting would be in specification.
2856 | Freezer FB-1244 found with no PM schedule while in use | jacob.kallon@sanofi.com | 80 NYA room 150 | UNKOWN | 2024-01-18 | It it expected that PM schedule be generated for new assets per MABC-QU-SOP-0036445
2855 | Gemba-8NYA-Downstream Cz Phenyl Blank Run Paperwork missing signatures | aidan.mula@sanofi.com | 8RM1503 | 2023-12-26 | 2024-01-18 | Steps should be performed at time of operation. | Yes | No Deviation
2854 | warm room temp alarm 148 at 34.47 | Nicole.lufkin@sanofi.com | 51NYA 151 | 2024-01-17 | 2024-01-18 | temp not below 36 |  | Deviation
2853 | ATF wetting leak during flush part of ATF wetting | Ralston.najarro@sanofi.com | 8NYA1403 | 2024-01-17 | 2024-01-17 | Media should flush through the ATF and into the waste tote
2852 | 8NYA Upstream D0 Sample Missed on Lot: DGHBN022 | gary.dodakian3@sanofi.com | Suite 2 Upstream 1403 | 2023-12-07 | 2024-01-17 | D0 samples are taken
2851 | 51 NYA Clean Steam EN-285 testing failed the Non-Condensable Gases (NCG). | brizaida.medina-vega@sanofi.com | 51 NYA Clean Steam Generators | 2024-01-16 | 2024-01-17 | Passing result for % NCG is less than or equal to 3.5% and the three samples from the Clean Steam Generator X9130 returned % NCG were 4.0%, 4.2% and 3.6% The three samples from the Clean Steam Generator X9120 returned were % NCG of 8.5%, 8.1% and 8.5 %. | Yes | Deviation
2850 | GEMBA for X2731 leak at buffer inlet | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2024-01-17 | 2024-01-17 | Colder connections do not leak | Yes | Deviation
2849 | GEMBA for X2200 TCU not controlling, Temperature too high | Brandon.Steff@sanofi.com | 8RM1333 8NYA | 2024-01-17 | 2024-01-17 | TCU controls to SP of 37C
2848 | Cerezyme - RF639D023 - 8NYA - ATF A Isolation Valve Failure During ATF B Changeout | michael.knapp@sanofi.com | 8NYA 1403 | 2024-01-16 | 2024-01-16 | ATF A remains operation during ATF B changeout.
2847 | Unicorn Calibration Certifications used for WAVE DeltaV upgrade IQOQ | Adriana.Alban@sanofi.com | 152B 51NYA | 2023-10-09 | 2024-01-16 | QU-SOP-0069490 Calibration Program
2846 | Bioburden result for Phenyl column CV came back >200 CFU | zachary.rickard@sanofi.com | 8NYA 1503 | 2024-01-16 | 2024-01-16 | CV samples will all come back with acceptable results
2845 | Shipper p/n GZ8017-01 batch 23MN0083 is missing the bottom foam insert | alan.rousseau@sanofi.com | room 1017 NYA | 2024-01-19 | 2024-01-12 | Per SPC QU-OPE-0421321 shipper 8017-01 should have both top, bottom and side foam inserts
2844 | Cerezyme - RF639D023 - 8NYA - ATF Wetting Media Drain Step not performed in EBR | michael.knapp@sanofi.com | 8NYA 1403 | 2023-12-16 | 2024-01-15 | Tote is drained per EBR.
2843 | GEMBA for SUB X2220 Low DO | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2024-01-14 | 2024-01-14 | DO control automatically starts after acknowledging growth prompt following SUB inoculation
2842 | Fabrazyme Capto Butyl Point Of Use 705672 DGHBE071 Bioburden Volume and Sample Time | Andrew.Brundige@sanofi.com | 68NYA, QC Micro | 2023-12-30 | 2024-01-13 | Per QU-SOP-0073109, a 10mL test volume must be filtered for sample 1007318(FZ2_CB_762729_POU705672_BIOB)
2841 | 8NYA GEMBA - 100L shipping bag inspection failure | alex.tello@sanofi.com | 8NYA RM1703 | 2024-01-14 | 2024-01-14 | Bag to pass inspection
2840 | Gemba-8NYA- Downstream Phenyl Column Packing- Automatic bed detection failed on Axichrom | aidan.mula@sanofi.com | 8RM1313 | 2024-01-13 | 2024-01-13 | Path forward on column operations
2839 | Unlabeled Sample 04Jan2024 | joshua.brooks@sanofi.com | 8NYA Sample laydown room 1410 | 2024-01-04 | 2024-01-05 | Samples are labeled.
2838 | 8NYA - No second verify provided for manual entry's, Records needed for next shipment. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2024-01-13 | 2024-01-13 | Hold Gemba to see path forward.
2837 | 8NYA - Tote chill step not executed in record. Record needed for next shipment. | bailey.gleason@sanofi.com | CNC | 2024-01-13 | 2024-01-13 | Hold Gemba to see path forward.
2836 | 8NYA GEMBA - EGHBY001 CM 200L load barrel leak | alex.tello@sanofi.com | 8NYA RM1503 | 2024-01-13 | 2024-01-13 | Path forward |  | Deviation
2835 | CTMS suite 2, Non-Viable Alert Level excursion for 45-CO2-11 | craig.stevenson@sanofi.com | 45NYA-187 | 2024-01-12 | 2024-01-12 | It is expected that non-viable results generate passing results in accordance with QU-SOP-0074750.
2834 | Missing delegate forms for quarterly review in 2023 | daniel.bati@sanofi.com | 8NYA | 2023-06-30 | 2023-12-29 | to find delegate forms in the binder
2833 | GEMBA for ABER Probe 1 not screwed into bellows correctly following autoclave | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2024-01-12 | 2024-01-12 | Probe is snug inside probe bellows.
2832 | Gemba-8NYA-Downstream Fz Post Viral Bag Fill- Recipe did not end after third bag fill | aidan.mula@sanofi.com | 8RM1313 | 2024-01-11 | 2024-01-11 | Recipe should end after third bag
2831 | GEMBA 51NYA: Temp excursion for vial shipment of ASM to patheon biologics | adam.rowe@sanofi.com | Patheon biologics | 2023-12-08 | 2024-01-11 | The temp loggers should be turned off before removing the vials form the shipper. It appears they were not. |  | Deviation
2830 | Tiny punch marks (indents) can be seen on JPN pouch label 849136 lot 23MN00412 | alan.rousseau@sanofi.com | Room 1009 at 74 NYA | 2024-01-11 | 2024-01-11 | Voids that render the lot, expiry or barcode illegible are a crtical defect and will result in failure of product and 100% inspection of all units labeled
2829 | 51NYA: BCP missing steps in area clearance | adam.rowe@sanofi.com | 51NYA/152 and 140 | 2024-01-11 | 2024-01-11 | Information in the BCP should match MES. |  | Deviation
2828 | Aseptic connection misaligned in CNC prior to transfer - Buffer 760730 | Nicholas.Grote@sanofi.com | 1220 | 2024-01-11 | 2024-01-11 | Good aseptic connection made. | Yes | No Deviation
2827 | A gemba is needed for: "RF1BRE002 Butyl Load filter tested with incorrect recipe". | jermaine.kellam@sanofi.com | 8RM1350 | 2024-01-10 | 2024-01-11 | That filter recipes are checked against the critical deviation interim control document for correct testing parameters.
2826 | Nanofilter clamped before Autoclaving - Nano NOT used yet | nicholas.grote@sanofi.com | 1372 | 2024-01-10 | 2024-01-10 | Nano tubing assembly clamped when it comes out of autoclave, not before.
2825 | Cz and Fz - Trace amounts of Lead (Pb) identified in Sartorius bags used in Fz and Cz processes | gregory.klotz@sanofi.com | NA | NA | NA | Expected that no lead would be identified in the study
2824 | Multiple Products- Lot-8 NYA- QCA Alarm from 1403 to neutral corridor is disabled. | amanda.sorel@sanofi.com | 8 NYA 1403 | 2020-12-01 | 2024-01-10 | It is expected that the QCA alarms are enabled.
2823 | 51NYA - Thyrogen APU- Room 152A- 6 wave bioreactors modified without as found data taken or updating of validated Metrology system | robert.bauch@sanofi.com | 51-152A | 2024-01-10 | 2024-01-10 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2822 | Fabrazyme End of Hold Lot: EGHBE001 P/N: 762729 Bioburden Volume | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2024-01-06 | 2024-01-10 | Per QU-SOP-0073109, a 10mL test volume must be filtered for sample 1010506 (FZ2_CB_762729_EOH_BIOB). | Yes | Deviation
2821 | 8NYA-Cz- X2240 Media Line not connected at time of tote switchover, leading to collapsed tubing and potential damage to skid pump | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-01-10 | 2024-01-10 | It is expected that media supply lines are welded to spply line wye, allowing for a transition from one supply tote to another
2820 | Water leak observed at 8NYA | edgar.vale@sanofi.com | 8NYA/ Corridor 1201 | 2024-01-10 | 2024-01-10 | Lab would be free of water leaks | Yes | Deviation
2819 | 8NYA- FZ suites GRADE D both MAL doors are stuck open to the CNC space | julianna.macconnell@sanofi.com | 8NYA Multiple Rooms | 2024-01-09 | 2024-01-09 | Differential pressure to be maintained. Doors to not be stuck open and to function as expected. | Yes | Deviation
2818 | 8NYA APU-CC8-room 1333-Flow indication transmitter FB-4898-048 and Flow Indicator FB-4898-049 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1333 | 2024-01-09 | 2024-01-09 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
2817 | Superseded Version of Manual Test Request Submitted for CZ MMC X2731001 CIP CV P/N: 716453; Lot: RF600E001 | abel.agnes@sanofi.com | 68NYA | 2024-01-07 | 2024-01-08 | Forms are submitted using the current effective version and are inclusive of all pertinent sample information.
2816 | Gemba-8NYA-Downstream SP- pH timer exceeded | aidan.mula@sanofi.com | 8NYA/RM1323 | 2024-01-08 | 2024-01-08 | Timer should not be exceeded
2815 | Gemba-8NYA-Downstream Fz Butyl-CNC Breakdown step was not opened by the record | aidan.mula@sanofi.com | 8NYA/ RM1323 | 2024-01-08 | 2024-01-08 | Record should open the steps that need to be performed. | Yes | Deviation
2814 | DO 100% standardization reading out of range given in batch record | james.letourneux@sanofi.com | 51 NYA Room 140 | 2023-11-09 | 2024-01-08 | Per Note below step 4.6.4 in QU-OPE-0525093, DO probe should be changed out if results are out of range.
2813 | FZ SP PH sample out of range. | alex.tello@sanofi.com | 8NYA RM1323 | 2024-01-08 | 2024-01-08 | path forward
2812 | AHU shutdown , need to small scale sample at risk | nicole.lufkin@sanofi.com | 51NYA 151 | 2024-01-07 | 2024-01-07 | N/A
2811 | RF1BRE001_002 Butyl Load filter tested with incorrect recipe | wilfredo.lopez@sanofi.com | 8RM1350 | 2024-01-06 | 2024-01-07 | That the filter recipes are checked against the critical deviation interim control document for correct testing parameters
2810 | 8NYA-Cz-Failed visual inspection of ATF Tubing set "Spare Permeate Line" prior to installation | john.lytle@sanofi.com | 8NYA Rm1403 | 2024-01-06 | 2024-01-06 | While not explicitly called out in the Inspection secdtion of the Single Use Tubing EWI, it is expected that no foreign material be found on sterile boundaries for Aseptiqwik G
2809 | Magnehelic Gauge 45 NYA Rooms 171 to 171A DP Within Action Level | jeff.hill@sanofi.com | 45 NYA Room 190 | 2023-12-24 | 2023-12-24 | QU-FOR-0091683, Attachment 1, DP must be above 0.04. |  | Deviation
2808 | 45NYA Plates EGFB2126 Expired | Cleiton.lopes@sanofi.com | Sample office | 2024-01-05 | 2024-01-05 | Samples are transported, received and tested within 24hr timeframe from collection
2807 | MCC X2560 went into hold due to end of run prompt not being answered. | Devin.Parks@Sanofi.com | 1403 | 2024-01-05 | 2024-01-05 | X2560 MCC should be able to finish collecting product without alarm due to end of run prompt on X2220.
2806 | 76 NYA PUR “Used of expired instrument to control homogenizer pressure during Vendor visit". | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2024-01-05 | 2024-01-05 | Per the Metrology Program, all instruments used at the manufacturing site, even for non GMP purposes, they must be calibrated befoe use
2805 | 36 Month Time Point was missed for the Bacterial Endotoxin testing of HA Lot #ABYHAM005 | Jonathan.edwards@sanofi.com | 76 NYA/Room 127/QC Biopolymers laboratory | 2023-12-07 | 2024-01-05 | Sample would be transported from the QCBP laboratory to the QC Micro laboratory where it would be tested for bacterial endotoxin within the stability window
2804 | 8NYA-Cz-RF639D022 Eluate Bag P/N 692168 failed visual inspection prior to installation | john.lytle@sanofi.com | 8NYA Rm1403 | 2023-12-26 | 2023-12-26 | Passing bag inspection per EWI-000015, Section 4 Bag Inspection
2803 | Cerezyme - RF639D019 - 8NYA - Day 0 Sampling Not Performed | michael.knapp@sanofi.com | 8NYA 1403 | 2023-10-21 | 2024-01-05 | Day 0 Sampling is performed per Sampling Plan.
2802 | Easy Cals due on 30Nov for FB-4784, 85, 86 and 87. Easy cal are in use. | Bhargav.Patel@sanofi.com | 1403 | 2023-11-30 | 2023-12-01 | Easy Cal send out for calibration by 30Nov
2801 | Gemba-8NYA- Cerezyme- Buffer 832327/DGHSL004- BOM update- Add 714049 and welding step at end of record | alejandro.martinez2@sanofi.com | 8NYA/CNC | 2024-01-03 | 2024-01-04 | Update BOM adding 714049 and step at end of EBR to weld manifold line, so it won't be missed. | Yes | No Deviation
2800 | 51NYA -Thyrogen APU-Room 108B F-14513-300 EMS temperature Indicator for F-14513 -80 Freezer found out of tolerance during a scheduled Calibration | robert.bauch@sanofi.com | 51-108B | 2024-01-02 | 2024-01-04 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results |  | Deviation
2799 | Performed by and verified by steps executed by the same operator during Thyrogen clarification DGF4B19A process at 51NYA. | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 139 | 2023-11-19 | 2023-12-22 | It is expected that the Performed by and Verified by steps are executed by two different operators per QU-SOP-0074684 step. 6.11.6
2798 | 68NYA Support APU-Metrology-Temperature Standard S-1695-010 and Decade Resistor S-1346 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2023-11-17 | 2024-01-04 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2797 | cold room 147 temp alarm | nicole.lufkin@sanofi.com | 51NYA 147 | 2024-01-04 | 2024-01-04 | alarm room is within temp
2796 | 51NYA Thyrogen APU-Room 140-F-6115-022 Agitator Indicator on Bioreactor E3 overdue for calibration | robert.bauch@sanofi.com | 51-140 | 2024-01-01 | 2024-01-03 | BQ-GOP-000046 Calibration Program Version 6 Section 5.7.5.5, unit should have been calibration by its due date, | Yes | Deviation
2795 | BIRF Attachment 1 of QU-SOP-0003344 missing entry | steven.servais@sanofi.com | 80 New York Av. Framingham, MA. | 2023-12-14 | 2024-01-03 | Birf Attachment 1 of QU-SOP-0003344 Portable Media Storage and usage log. should have been completed at the time the material was moved into the cold room. |  | Deviation
2794 | BIRF Attachment 1 of QU-SOP-0003344 missing entry | steven.servais@sanofi.com | 80 New York Av. Framingham, MA. | 2023-12-14 | 2024-01-03 | Birf Attachment 1 of QU-SOP-0003344 Portable Media Storage and usage log. should have been completed at the time the material was moved into the cold room.
2793 | Cz MCC skid X2831 Eluate SUM leaking | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2024-01-03 | 2024-01-03 | Bag does not leak | Yes | Deviation
2792 | 8NYA APU-PF8-Room 1503-4 GMP Critical Instruments overdue for calibration | robert.bauch@sanofi.com | 8-1503 | 2024-01-01 | 2024-01-03 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.7.3 Instrument should have been received into Metrology before it's due date
2791 | Gemba-8NYA-Downstream- Fz SP- Skid Disassembly not signed for at time of operation | aidan.mula@sanofi.com | 8NYA RM1323 | 2024-01-03 | 2024-01-03 | Operations should be signed at time of performance | Yes | Deviation
2790 | WFI Leak identifed at 8NYA Mech space | craig.stevenson@sanofi.com | 8NYA, Mechanical space room 1170, upper mezzanine | 2024-01-03 | 2024-01-03 | No WFI Leak
2789 | warm room temp alarm, two occaurances | nicole.lufkin@sanofi.com | 51NYA 153 | 2023-12-31 | 2024-01-03 | warm room temp 36-38 for 125mL flask stage |  | No Deviation
2788 | Gemba-8NYA-Downstream Butyl-Load filter was tested under incorrect recipe. | aidan.mula@sanofi.com | 8RM1350 | 2024-01-03 | 2024-01-03 | FIT information should be provided correctly and verified against the interim control.
2787 | Multiple Products - Multiple Lots - 8NYA - PCS Alarms not collected into MES | michael.knapp@sanofi.com | 8NYA 1403 | 2023-12-27 | 2024-01-02 | Quality/Critical Alarms are populated in MES as PCS alarms for review.
2786 | Overdue biennial requalification on stability Incubator F-15901 (Asset: 80-INCUBATOR-9) | kathy.tourville@sanofi.com | Room 148A in 80 NYA, Framingham MA | 2024-01-01 | 2024-01-02 | RQ would be performed within the required calendar month per FBL-SOP-000994.
2785 | 4 late PMs due 01Jan2024 0000 | kevin.southwick@sanofi.com | 8NYA (3) and 45NYA (1) - 2 separate deviations but 1 GEMBA | 2024-01-01 | 2024-01-02 | no late PMs
2784 | Weight Tolerance Issue during weigh out - Multiple Occurrence (RF557D048, RF608D010, RF547D017) | Nicholas.grote@sanofi.com | 1363 | 2023-12-06 | 2024-01-02 | complete weigh out within calibration range of scale.
2783 | CZ Bioburden Expired at 1049_01JAN24 | diane.collins@sanofi.com | 68 NYA ,4720 sample receipt | 2024-01-01 | 2024-01-01 | QU-SOP-0070558 step 6.2.5, samples required to be put on test 24 hours after being sampled.
2782 | Gemba-8NYA-Downstream Cz Thaw/Pooling Bag Leak noticed after HOFF unload | aidan.mula@sanofi.com | 8NYA/ RM1212 | 2023-12-31 | 2023-12-31 | Bags should not leak
2781 | 08 NYA UHU 8210 unexpectedly shutdown at approximately 2130 30DEC23 | Alex.Biskaduros@Sanofi.com | 08 NYA rooms 1400-throgh 1404 | 2023-12-30 | 2023-12-30 | AHU 8210 should not have shutdown unexpectedly
2780 | AHU 8210001 is not running for US Suite 2 | caleb.coppola@sanofi.com | 8NYA Rm 1403 | 2023-12-30 | 2023-12-31 | Expected result is the AHU- 8210001 is operational and would alarm and call out immediately if not operational.
2779 | 8NYA - RF543D009 705676 - filter was removed before clamp was in place. | bailey.gleason@sanofi.com | 1350 | 2023-12-30 | 2023-12-30 | Hold Gemba to see path forward.
2778 | GEMBA-8NYA-Downstream- Fz SP- Peak B values low out of spec | aidan.mula@sanofi.com | 8NYA/8RM1323 | 2023-12-30 | 2023-12-30 | Peak B should be within specification.
2777 | CZ MCC skid 2811 glycol waste filter change out | William.Christie@Sanofi.com | 8NYA room 1403 | 2023-12-30 | 2023-12-30 | Determine path forward and change out Glycol waste filter.
2776 | Fz Butyl- Mannifold Connection Hose had small leak | nathaniel.barnes@sanofi.com | 8 NYA 1323 | 2023-12-29 | 2023-12-29 | Determine path forward for processing
2775 | GEMBA-8NYA-Downstream-Disassembly step not signed at time of operation | aidan.mula@sanofi.com | 8NYA/ RM1323 | 2023-12-29 | 2023-12-29 | Steps should be signed at time of performance
2774 | Three freeze bags filled 7.299 and the limit is 7.3 | daniel.bati@sanofi.com | 8NYA | 2023-11-30 | 2023-12-28 | The bag weight should be within range.
2773 | 8NYA - Fabrazyme SP- DGHSE024 - Collection SUM self-taring during elution | stanley.durand@sanofi.com | 8NYA room 1313 | 2023-12-27 | 2023-12-27 | Collection should not tare during elution
2772 | Filter Serial Number Discrepancy Between MES and FIT Printout | craig.valery@sanofi.com | 51 NYA | 2023-12-10 | 2023-12-27 | accurate information expected to be recorded |  | Deviation
2771 | Missing Bulk Intermediate Product Reconciliation Logs | craig.valery@sanofi.com | 51 NYA | 2023-12-27 | 2023-12-27 | it is expected that all issued attachments be reconciled with QU-OPE-0529697 |  | Deviation
2770 | SP Buffer unable to chill during Production | Juan.Bonilla@sanofi.com | 8NYA/1323 | 2023-12-27 | 2023-12-27 | Chilled tote should be at temp.
2769 | Mixer did not mix during PH Adjust recipe. | Juan.Bonilla@sanofi.com | 8NYA/1323 | 2023-12-27 | 2023-12-27 | Agitator to mix during PH Adjust recipe. | No | Deviation
2768 | Incorrect iLims sequence used for 2023-51CV-0046 | craig.valery@sanofi.com | 51 NYA | 2023-12-26 | 2023-12-27 | It is expected that the corresponding iLims sampling sequence be used for T-2911 |  | No Deviation
2767 | Manual transition needed for Variable bleed due to capacitance not reach at 145 | bhargav.patel@sanofi.com | 8 NYA / 1403 | 2023-12-27 | 2023-12-12 | Bleed to switch over to | Yes | Deviation
2766 | 8NYA CZ Capto Butyl Eluate P/N 761217 DGHBN020 - Only 1 retain submitted instead of 2 | abel.agnes@sanofi.com | 68NYA | 2023-12-23 | 2023-12-24 | Two separate retains are submitted with one ID on each. | Yes | Deviation
2765 | 8NYA Gemba-Downstream- 2 Butyl Load SUM bags found defective | aidan.mula@sanofi.com | 8NYA RM1323 | 2023-12-25 | 2023-12-25 | Bags should pass inspection
2764 | Incorrect testing method for bioburden sample 2300019337 | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2023-12-21 | 2023-12-25 | Per Attachment 2 of QU-SOP-0073109, 100mL must be filtered on one plate then rinsed with 300mL Fluid A.
2763 | Unable to create stock in Athena for H14 harvest on D2 bioreactor | David.Loik@sanofi.com | 51 NYA | 2023-12-24 | 2023-12-24 | Harvest stock would be able to be created as normal
2762 | 8NYA - Fabrazyme - DGHDS001 - Unable to test shipping bag transfer filter due to part number not being in the interim control being in the | stanley.durand@sanofi.com | 8NYA room 1313 | 2023-12-22 | 2023-12-22 | Interim control "Table 2. Fabrazyme-ICB FIT Recipe Formula to be Used" should include FIT for Post Viral Dispensing part number 894747.
2761 | Fabrazyme post viral dispensing recipe did not end properly. | kenny.leland@sanofi.com | 8 NYA/ Room 1313 | 2023-12-22 | 2023-12-22 | Recipe should have ending and not been looking to fill a third barrel.
2760 | CZ Train 2 SUB X-2230 DGHBN020 | joseph.gonzalez@sanofi.com | 1403 | 2023-12-22 | 2023-10-15 | . ATF to function
2759 | 8NYA-Cz- X2250 pH probe 2 is not controlling as expected | john.lytle@sanofi.com | 8NYA Rm1403 | 2023-12-22 | 2023-12-22 | both pH probes controlling sparge rates as intended | Yes | No Deviation
2758 | 8NYA CZ Bioburden sample 974838 leaking upon receipt | abel.agnes@sanofi.com | 68NYA 4th Floor | 2023-12-22 | 2023-12-22 | Sample containers should not be cracked or leaking.
2757 | 8NYA CZ DGHBN020 3x Bioburden samples expired | cleiton.lopes@sanofi.com | 68NYA | 2023-12-21 | 2023-12-21 | Samples are received and tested within 24hrs from collection time
2756 | 6 daily WFI sites not sampled in 8NYA on 20Dec23 | marissa.lentoni@sanofi.com | 68 NYA, 4700, QC Micro Lab | 2023-12-20 | 2023-12-21 | All daily WFI sites would be collected at 8 NYA.
2755 | Product shipment temperature excursion | cleiton.lopes@sanofi.com | 45NYA QCC | 2023-12-18 | 2023-12-21 | Shipment maintains within acceptable range during transport
2754 | Freeze Bag under filled during post viral freeze bag dispensing record DGHDT020 | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1703 | 2023-12-21 | 2023-12-21 | To get a Phath forward to continue with record
2753 | Gemba: Multiple biob and settle plate paperwork reviewed without training on current version of SOP | judith.keller@sanofi.com | 68 NYA 4700 QCM | 2023-12-16 | 2023-12-20 | It is expected that analysts are trained on SOPs when executing.
2752 | 8NYA FZ DGHCE008 953285 duplicate sample submitted | cleiton.lopes@sanofi.com | 8NYA, 68NYA | 2023-12-20 | 2023-12-20 | No duplicate submissions to QC and no duplicate receipt | Yes | No Deviation
2751 | warm room 148 LOLO temp alarm | nicole.lufkin@sanofi.com | 148 51NYA | 2023-12-20 | 2023-12-21 | 35-37.5C
2750 | Document Approved without Doc Control Approval | kyle.lebherz@sanofi.com | 68 Nya | 2023-12-20 | 2023-12-20 | Doc Control would be a final authorizer on the document
2749 | Media equilibration not met for 24L scaleup | David.Loik@sanofi.com | 51 NYA | 2023-08-30 | 2023-12-19 | Witness step would be signed for.
2748 | Storage movement entry missing on Q fraction 7 BIRF lot DGF3B15A | craig.valery@sanofi.com | 51 NYA RM131 | 2023-12-14 | 2023-12-19 | it is expected that all applicable fields are filled in on QU-SOP-0074092 |  | Deviation
2747 | 8NYA-Cz-X2220 ATF A black particulate observed during wetting operations | john.lytle@sanofi.com | 8NYA Rm1403 | 2023-12-19 | 2023-12-19 | As the media is filtered by Solution Prep for final tranfer into a tote and the ATFs are certified clean sterile and gamma irradiated, no free floating particulates are expected to be observed during this operation
2746 | Investigate button sight glass on WFI still #2 is broken. See pictures. While doing walk-through at 05:40 19dec2023 was found water on floor around WFI still #2. after close look was found that button sight glass was leaking water from glass. Wfi still #2 is off due this issue. Wfi still #1 is still running. | jesus.acosta2@sanofi.com | 8 NYA 1170 | 2023-12-19 | 2023-12-19 | No broken sight glass, reference pictures attached
2745 | Fabrazyme - DGHCE005 - 8NYA - Eluate 2 Bioburden/Endotoxin samples not delivered to QC | michael.knapp@sanofi.com | 8NYA 1333 | 2023-09-12 | 2023-12-18 | Samples are delivered to the sample fridge at time of sampling for pickup and delivery to QC. | Yes | Deviation
2744 | No product-no lots-68NYA: Incubator maintenance log not completed within 6-8 day window as stated in step 6.4 of QU-SOP-0049227 | larissa.fairweather@sanofi.com | 68NYA/4735/QCM | 2023-12-15 | 2023-12-18 | Per QU-SOP-0049227 step 6.4 we are to verify incubator water level each week, 6-8 days from the last recording.
2743 | Leaking Ceiling in 45 NYA 190 | jeff.hill@sanofi.com | 45 NYA Room 190 | 2023-12-18 | 2023-12-18 | Room integrity. QU-SOP-0073883
2742 | CM eluate SUM/Drive mixer would not mix and also go into alarm | paul.devaney@sanofi.com | 8NYA/1503 | 2023-12-17 | 2023-12-17 | Sum should have mixed properly according to the recipe in DeltaV
2741 | 760726 - RF608D011 - DGHAI011 - 900 L batch - Tested high for pH and low for Condo | leo.gillespie@sanofi.com | 1350 - Solution prep | 2023-12-18 | 2023-12-17 | Sampling and testing instructions in FBL-EWI-000022 , Sampling plan in FBL-EWI-000009, and FBL-EWI-000019--expected acceptable results given attention to process details
2740 | X2240 ATF connections incorretly oriented | Ralston.najarro@sanofi.com | 8NYA1403 | 2023-12-15 | 2023-12-15 | Connecter should be oriented upright
2739 | SIP recipe running, one TI/trap will not heat up | paola.tello@sanofi.com | 8 NYA room 1333 | 2023-12-15 | 2023-12-15 | Recipe should run with no holds
2738 | D2 Harvest/Refeed Timer Exceeded | nicole.lufkin@sanofi.com | 51 NYA Room 140 | 2023-12-15 | 2023-12-15 | Per MSAT-QU-OPE-0172699, page 70, "Harvest Phase" PCE table, point 16: Re-feed interval should occur between 22-26 hours.
2737 | 76 R&D FM-RDSQO-23RQ0271 Autoclave RQ hood cleaning not recorded before QCM testing | judith.keller@sanofi.com | 68 NYA 4700 QCM | 2023-11-21 | 2023-12-14 | It is expected that hoods are cleaned before testing per QU-SOP-0053230 step 6.1.7/ QU-SOP-0075304.
2736 | 76 NYA Action Notices 2023-N-058, 2023-N-059, and 2023-N-060 released without QC results being completed | paul.consolmagno@sanofi.com | 76NYA | 2023-12-14 | 2023-12-15 | Validated results would be available prior to release
2735 | Incorrect part number used to test Filter during Formulation/Nanofiltration DGF3B11 process at 51NYA. | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 87 | 2023-11-19 | 2023-12-14 | It is expected that the correct filter part number is used during Formulated filter testing per QU-SOP-0072849 attachment 6
2734 | 2 AHUs at 51 NYA went down unexpectedy | Alex.Biskaduros@Sanofi.com | 51 NYA room 149 to 84 and additional rooms and rooms 152 ton 149 | 2023-12-15 | 2023-12-15 | AHU 7 should not have shutdown unexpectedly | Yes | Deviation
2733 | shipping bag transfer recipe did not end appropriately on record RF1E4D019 | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1703 | 2023-12-14 | 2023-12-14 | To get a Phath forward
2732 | Organism of Concern, Thyrogen (TSH) Maxisette Retentate Equilibration P/N 5151-00, Lot DGF3B15A | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2023-10-05 | 2023-12-14 | It is expected that no OoCs are in product
2731 | 8NYA-Cz- Train 2 Eluate #4, liquid observed in bottom of collection tote after transfer of 9.5kg | john.lytle@sanofi.com | Rm 1434 | 2023-12-13 | 2023-12-13 | it is expected that the system remain integral during eluate tranfer, no observable liquid outside of the SUT sterile boundry
2730 | Gemba – 8NYA – 3K tote bag built incorrectly | william.lee@sanofi.com | 8NYA/1290 | 2023-12-12 | 2023-12-12 | 3K tote bag should have had the 4 holes for installing hoist
2729 | Calc chem13 recovery failure | Jonathan.Ahiaba@sanofi.com | 8NYA | 2023-12-12 | 2023-12-12 | 95-105% recovery
2728 | 8NYA APU-Room 1323-Pressure indicator FB-2949-002 on system X-3091001 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1323 | 2023-12-12 | 2023-12-12 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2727 | 8NYA Upstream Harvest Vessel train 1 / V-2616 Overlay Actuator Valve failure | gary.dodakian3@sanofi.com | Suite 2 Upstream 1403 | 2023-12-11 | 2023-12-12 | Overlay does not fail and maintain proper CO2 levels
2726 | Production Media valve "BV-2040004" in CNC not open after switchover to Production. | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2023-12-12 | 2023-12-12 | Production Media valve should open by Delata V recipe.
2725 | Incorrect MO ID generated during Pool Formulated BULK DGFPB012 process at 51 NYA Thyrogen | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 87 | 2023-12-07 | 2023-11-11 | It is expected that the correct MO ID is generated during Pool Formulated BULK process
2724 | 76NYA Biosurgery APU-Room 352-F-4341 Autoclave with no as found data taken before replacement of the the PLC | robert.bauch@sanofi.com | 76-352 | 2023-12-11 | 2023-12-11 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before Replacement
2723 | Multiple CTMS Gowning OJTs not completed | craig.stevenson@sanofi.com | 45NYA, CTMS | 2023-06-01 | 2023-12-11 | It is expected that CTMS trainers follow GxP during OJT training.
2722 | 8NYA APU-Room 1503-PF8- FB-4311-007 Weight Indicator found out of tolerance during a Scheduled calibration. | robert.bauch@sanofi.com | 8-1503 | 2023-12-08 | 2023-12-11 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2721 | HOFF 2 Freeze Recipe went into hold for Overpressure Compressor. | Juan.Bonilla@sanofi.com | 8NYA/1365 | 2023-12-09 | 2023-12-09 | To resume recipe
2720 | Freezer bag overfilled during ROSS Opertions | Juan.Bonilla@sanofi.com | 8NYA | 2023-12-09 | 2023-12-09 | Freeze bag to fill within range.
2719 | Buffer 768506 0.1 M sodium hydroxide, 50% hexylene glycol Failed two FIT attempts | nicholas.grote@sanofi.com | 1290 | 2023-12-09 | 2023-12-09 | Pass FIT
2718 | 8NYA - Cz Seed Media - 760744/DGHAX009 - Bottle Stock Discrepancy | alejandro.martinez2@sanofi.com | 8NYA/RM: 1350 | 2023-08-21 | 2023-12-09 | Event needed to justify integrity check to terminate record, need to review this record before lot shipment Mon/Tuesday
2717 | DGHAQ010 - 760737 Samples taken in Biolims not recieved in QC | nicholas.grote@sanofi.com | 1350 | 2023-10-15 | 2023-12-08 | Samples received by QC for testing.
2716 | Missed witness step following refeed | David.Loik@sanofi.com | 51 NYA | 2023-12-08 | 2023-12-08 | Witness step would be signed for. |  | Deviation
2715 | QCA alarm for LoLo level on cryostat 11 at 45NYA | David.Loik@sanofi.com | 51 NYA | 2023-12-07 | 2023-12-07 | Level would be controlled within range.
2714 | Path forward with Cz CM record | zachary.rickard@sanofi.com | 8NYA/1503 | 2023-12-08 | 2023-12-08 | N/A
2713 | Missing Samples at time of batching DGHAD018 - Buffer 760720 | nicholas.grote@sanofi.com | 1350 | 2023-10-09 | 2023-12-08 | Take samples at time of batching
2712 | OneLims sample ID#: 965293 / Lot # DGHAW185 (ENDO_CHROM) submitted twice for testing to QC Micro | amy.roy@sanofi.com | 68NYA/4700/QC Micro | 2023-10-06 | 2023-10-08 | We would expect to receive only 1 vial for each sample
2711 | Incorrect Verifying Comment for Chem013 Information | zachary.rickard@sanofi.com | 8NYA/1503 | 2023-11-17 | 2023-12-07 | Verifier will make their own individual comment saying they verified the manual entry.
2710 | SP2 column leak during equilibration | craig.valery@sanofi.com | 51 NYA, RM 132 | 2023-12-08 | 2023-12-08 | It is expected that no leaks are observed |  | Deviation
2709 | Transfer Disassembly not signed for at time of occurrence. | zachary.rickard@sanofi.com | 8NYA 1212 | 2023-11-30 | 2023-12-08 | Steps will be signed for in real time.
2708 | Transfer Disassembly not signed for at time of occurrence. | zachary.rickard@sanofi.com | 8NYA 1212 | 2023-11-30 | 2023-12-08 | Steps will be signed for in real time.
2707 | 8NYA GEMBA- Downstream Cz Phenyl- Step was not signed for at time of operation. | aidan.mula@sanofi.com | 8NYA RM1503 | 2023-12-07 | 2023-12-07 | Event opened to terminate record and move forward with review.
2706 | Pompe-8NYA-Credo Coolers not Qualified for Sample Transport at 6-10C | julianna.macconnell@sanofi.com | 8NYA | 2023-12-07 | 2023-12-07 | Per QU-SOP-0044843 ‘Validation Program for the Massachusetts BioCampus’, transport solutions are to be prospectively qualified prior to initial use. The credo cubes have not been qualified for 6-10°C transport. | No | Deviation
2705 | X2230 Pharmed Tubing unable to be sealed | Ralston.najarro@sanofi.com | 8NYA1403 | 2023-12-07 | 2023-12-07 | Seal off pharmed tubing and cut tubing to perform ATF changeout | Yes | Deviation
2704 | 68NYA Support APU QCMicro-Room 4700- 3 TOC Analyzers found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-4700 | 2023-12-04 | 2023-12-07 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2703 | QC micro: Positive control results not entered for 8 NYA critical utilities gas sampling batch INC-08294 | megan.botti-marino@sanofi.com | 68 NYA, QC Microbiology Lab | 2023-09-28 | 2023-12-06 | Step 9.1.8 of MABC-QU-SOP-0023473 describes how to read positive control plates for anaerobic samples at the completion of incubation.
2702 | 8NYA GEMBA- CZ Q-Membrane EBR does not align with EWI. | alex.tello@sanofi.com | 8NYA RM1503 | 2023-12-06 | 2023-12-06 | path forward
2701 | 8NYA - RF601D009 760712 Tolerance violation during weigh and dispense out of calibration range. | bailey.gleason@sanofi.com | 8NYA | 2023-12-06 | 2023-12-06 | Hold Gemba to see path forward.
2700 | DGHAH024 - 1kL Chilled Tote Bag Installed Incorrectly | jermaine.kellam@sanofi.com | 8RM1350 | 2023-12-05 | 2023-12-06 | The tote bag should have been installed properly with the recirculation line located at the top of the bag
2699 | DGHBE062 - Fabrazyme Butyl Load Filter Tested with Incorrect Recipe | wilfredo.lopez@sanofi.com | 8RM1290 | 2023-11-30 | 2023-12-06 | That the filter recipes are checked against the critical deviation interim control document for correct testing parameters
2698 | GEMBA - 8 NYA - Leak on inlet bleed valve of nano filter | adam.keister@sanofi.com | 8 NYA, 1503 | 2023-12-05 | 2023-12-05 | To document leak in phenix event FRAD23E0908. Known issue that needs to be documented properly.
2697 | Cerezyme - RF639D022 - 8NYA - SUB inoculation with shortened DO control | michael.knapp@sanofi.com | 8NYA 1403 | 2023-12-05 | 2023-12-05 | SUB equilibration allows for proper DO control prior to inoculation.
2696 | Nano FIT Flow Rate unstable Error @480 Seconds (Testing duration 600 seconds) | nicholas.grote@sanofi.com | 1290 | 2023-12-05 | 2023-12-05 | Pass FIT
2695 | Fabrazyme - DGHCE008 - 8NYA - No Antifoam Dosing from 02DEC23 until 04DEC23 | michael.knapp@sanofi.com | 8NYA 1333 | 2023-12-04 | 2023-12-04 | Line is fully primed and scheduled dosing occurs starting on Day 5. | Yes | No Deviation
2694 | Incomplete documentation | patrick.hosegood@sanofi.com | 74 NYA room 1100 | 2023-12-04 | 2023-12-04 | Working copy completed at time of printing. Areas not used should be N/a'd and initialed and dated |  | No Deviation
2693 | During SP1 chromatogram review for lot DGF4B17C, pH was found to be out of range through out the process. | craig.valery@sanofi.com | 51 NYA | 2023-11-29 | 2023-12-04 | it is expected that the pH remains with in specified ranges |  | Deviation
2692 | Durning MES review it was found that the freezing manifold filter P/N:774182 was tested using an incorrect FIT program. | craig.valery@sanofi.com | 51 NYA | 2023-12-01 | 2023-12-04 | Per QU-SOP-0072849 and QU-OPE-0529697, step 4.4.3.6, the correct program should have been 6906-01.
2691 | 8NYA-Cz- Daily SUB sample roperations repeated after MSAT recommendation (based on sample results prior to pH adjustment) | john.lytle@sanofi.com | 8NYA, Rm 1403 | 2023-10-22 | 2023-12-04 | it is expected to GEMBA any out of specification result based on BGA analysis per EWI
2690 | 51 NYA autoclave alarm during bowie dick cycle | denise.mason2@sanofi.com | 51 NYA, room 129 | 2023-12-04 | 2023-12-04 | Bowie dick cycle would complete without alarm.
2689 | SUB bags failed visual inspections | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2023-12-01 | 2023-12-01 | SUB bag should pass visual inspection | Yes | Deviation
2688 | Missing Coliform Paperwork for CU site 409 at 76NYA | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2023-08-13 | 2023-12-04 | It is expected that all QCM paperwork be accounted for.
2687 | 705663 Storage Unit removed from tote during Cz MMC opts | robert.morawski@sanofi.com | 8NYA CNC | 2023-12-04 | 2023-12-04 | Storage Unit remained in tote where the buffer was used for production, and was available to be taken out into the correct EBR.
2686 | Cz Zinc- Chilled WFI unavailable for Dilution | nathaniel.barnes@sanofi.com | 8 NYA 1503 | 2023-12-02 | 2023-12-02 | Determine path forward for Zinc load preparation
2685 | High Pressure event on ATF guard filters X2220 | Ralston.najarro@sanofi.com | 8NYA1403 | 2023-12-02 | 2023-12-02 | Filters should not clog allowing flow through to the harvest bag |  | Deviation
2684 | Product(s) – Lot (or “Multiple Lots”) – 76 NYA – Late PMs at 76 NYA for the Month of November 2023 | kevin.southwick@sanofi.com | 76 NYA | 2023-12-01 | 2023-12-01 | SOP MABC-QU-SOP-0021310-Maintenance Operations, section 6.4.8.
2683 | Pressure event on X2220, DO probes diverging | Ralston.najarro@sanofi.com | 8NYA1403 | 2023-12-01 | 2023-12-01 | Probe should not be diverging
2682 | Late PMs at 8 NYA and 45 NYA for the Month of November 2023 | kevin.southwick@sanofi.com | 8 NYA, 45 NYA | 2023-12-01 | 2023-12-01 | SOP MABC-QU-SOP-0021310-Maintenance Operations, section 6.4.8.
2681 | 8NYA - RF525D042 Nanofilter was autoclaved with the incorrect gaskets. | bailey.gleason@sanofi.com | 8NYA | 2023-12-01 | 2023-12-01 | Hold Gemba to see path forward.
2680 | Ineffective CAPA FRAD22A0163 Identified During Effectiveness Review FRAD22R0125 | kyle.tesorero@sanofi.com | 8NYA | 2023-10-03 | 2023-11-30 | Equipment with failed calibration would remain in the Out of Service state until a passing calibration result is achieved
2679 | 51NYA Thyrogen APU-Room 140-FB-1560-019 Pressure Gauge overdue for calibration | robert.bauch@sanofi.com | 51-140 | 2023-12-01 | 2023-12-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
2678 | Simulated high pressure event on X2220 due to pendotech being disconnected. | Devin.Parks@Sanofi.com | 1403 | 2023-12-01 | 2023-12-01 | Pendotech should not have been disconnected causing a high pressure event. This event led to pH and DO being out of range.
2677 | Incorrect connection used during buffer transfer | nicholas.grote@sanofi.com | CNC | 2023-11-30 | 2023-11-30 | Connect to transfer port on tote bag. Complete transfer successfully.
2676 | 705663 storage unit taken out on unused buffer tote | robert.morawski@sanofi.com | 8NYA CNC | 2023-11-30 | 2023-11-30 | Storage Unit remained in tote. No buffer has been used, and should not have been taken out.
2675 | During Blue Split Filtration if was identified that the TSK column capacity will be exceeded | craig.valery@sanofi.com | 51 NYA | 2023-11-30 | 2023-11-30 | It is expected loading capacity is not exceeded
2674 | 8NYA-Cz-RF658D021- Single MiniCapt utilized for US processing will expire calibration at midnight, leaving Upstream without a calibrated option for production activities after midnight | john.lytle@sanofi.com | 8nya Rm 1343 | 2023-11-30 | 2023-11-30 | calibrated equipment for production activities
2673 | GEMBA for Bioburden Count Not Entered into BioLIMS | jessica.kychler@sanofi.com | QC Microbiology, 68 NYA floor 4 room 4700 | 2023-11-28 | 2023-11-29 | Per QU-OPE-0329679 STEP 9.2.11
2672 | RODI Level Sensor Was Bumped Into by Project Contractor | MA-BioCampus_F&E-Compliance@sanofi.com | 68NYA RODI room | 2023-11-28 | 2023-11-28 | no leaks or loose clamps on the RODI system | Yes | Deviation
2671 | 76NYA Biosurgery APU-Room 355-F--14980 Balance found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 76-355 | 2023-11-29 | 2023-11-30 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2670 | Cz Nano Failed FIT Attempt 2 | nicholas.grote@sanofi.com | 1503 and 1290 | 2023-11-29 | 2023-11-30 | Pass FIT
2669 | 8NYA-Cz-DGHCG020 Wave Equil hold due to Common Trouble alarm, automation unable to restart recipe post rocker EM hold | john.lytle@sanofi.com | 8NYA Rm1403 | 2023-11-29 | 2023-11-29 | After completion of the media transfer, the wave rocker acquires the associated EMs and begins rocking, without physical interference
2668 | Missed bioburden count | anastacia.hachem@sanofi.com | 68 NYA QC Microbiology | 2023-11-28 | 2023-11-29 | Per QU-OPE-0329679 STEP 9.2.11
2667 | GEMBA - 8 NYA - Leak from top inlet bleed valve during CZ nano ops | adam.keister@sanofi.com | 8 NYA, 1503 | 2023-11-29 | 2023-11-29 | Document according to RBIA
2666 | GEMBA - 8 NYA - Automation Intervention during Nano Equil Fill Time | adam.keister@sanofi.com | 8 NYA, 1503 | 2023-11-29 | 2023-11-29 | Investigate and adjust equil prime recipe as needed.
2665 | No data recorded at time of use for Eluants 1 and 2 for QU-OPE-0202442, missing page 5 of 5 for worksheet QU-FOR-0052577 | sarah.adams@sanofi.com | 45 NYA | 2023-11-19 | 2023-11-29 | Page to be printed and reagent recorded.
2664 | 8NYA - Downstream - EZ Cals to be put OOS. | stanley.durand@sanofi.com | 8NYA Room 1503 | TBD | 2023-11-28 | It's expected that for GMP crit to be calibrated on schedule.
2663 | Incorrect SP recipe launched Post PM | Juan.Bonilla@sanofi.com | 8NYA | 2023-11-29 | 2023-11-29 | Recipe to include air purge
2662 | 76NYA Biosurgery APU-Room 231-F--18066 Balance found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 76-231 | 2023-11-28 | 2023-11-29 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2661 | No reportable results for Cz samples in batches FBL1327-000000303 and FBL1327-000000304 | kathryn.lentoni@sanofi.com | QC Micro 4700 | 2023-11-20 | 2023-11-28 | For there to be counts entered for samples 890933 and 909185
2660 | SIP failure caused discard of microcarriers. D2 was reWFI for next batch of carriers | jesse.welch@sanofi.com | 51-140 | 2023-11-27 | 2023-11-27 | operate as designed |  | Deviation
2659 | 76 NYA During Routine EM of Common Areas, Static or Dynamic Not Checked for 3 Sites, 220-1, 220-2, 220-3 | megan.botti-marino@sanofi.com | 76 NYA | 2023-11-14 | 2023-11-28 | Per step 6.7 of QU-SOP-0074634, the sampling analyst records the activity level, static or dynamic
2658 | No product - Not Lots - 51NYA - No reportable result for routine EM site 130B Viable Air (866479) on 25Sep2023 | craig.stevenson@sanofi.com | 51 NYA Room 130B | 2023-09-25 | 2023-11-28 | It is expected that Viable Air samples are taken during rourinte EM
2657 | HA - DBYHAS005 - NBO - HA product was found to be OOT during transit to customer Alcon | Juan.Ramos-Torres@sanofi.com | NBO | 2023-11-07 | 2023-11-20 | from temp logger profile, ideal range: -20.0 C to 8.0 C
2656 | QU-SOP-0033695 (Non-Routine Sample Request Procedure) Revised to Version 8.0 without a Final Authorizer | brian.mcgee@sanofi.com | N?A | 2023-11-22 | 2023-11-27 | It is expected for a Doc controller to be selected as a signatory on SOP approvals.
2655 | Seprafilm - 76NYA RM233 - DBYSEP108B - incorrect IPA 70% used | fatima.fragaalvarez@sanofi.com | 76 NYA RM233 | 2023-11-21 | 2023-11-27 | expected to use IPA PN GZ4528-01 but operators used IPA PN GZ5528-02
2654 | Missing samples for Pompe DGHDR003X Shipment | ana.salichs@sanofi.com | 8 NYA | 2023-07-11 | 2023-11-27 | Samples are received and tested by QC.
2653 | 45NYA-Support APU-FFO2-Room 228-TOC FB-2469 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 45-228 | 2023-11-20 | 2023-11-27 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2652 | 51 NYA Thyrogen Q lot DGF1B13B Fraction samples identified to be leaking during sample pickup by MLC | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 108C | 2023-11-26 | 2023-11-26 | It is expected that samples show no sign of leaking during process and transport to sample QC. |  | No Deviation
2651 | 8NYA - Cerezyme - Shipping observed to have a damaged seal during post fill inspection. | stanley.durand@sanofi.com | 8NYA Room 1220 | 2023-11-26 | 2023-11-26 | Seals should not appear broken
2650 | Solution Prep - Leak during Sealing of transfer line on Media Tote | nicholas.grote@sanofi.com | CNC Space outside solution prep | 2023-11-26 | 2023-11-26 | Good seal, no leaking.
2649 | Warm room 153 QCA for LoLo temperature | David.Loik@sanofi.com | 51 NYA | 2023-11-26 | 2023-11-26 | Temperature would be controlled within range.
2648 | Incorrect incubation paramater for pompe sample 23MS02349 | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2023-11-17 | 2023-11-26 | Per QU-SOP-0036604, "Plates are incubated at 20-25C for 3-4 days and then at 30-35C for 48-72 hours"
2647 | 8NYA - Cerezyme - Phenyl scale out of calibration. | stanley.durand@sanofi.com | 8NYA room 1503 | 2023-11-21 | 2023-11-25 | Scales to be calibrated prior to use
2646 | Chromatogram review, pH out of range SP2 elution (reading 8.97 range 8.1-8.9) | craig.valery@sanofi.com | 51 NYA RM132 | 2023-11-12 | 2023-11-25 | it is expected that the pH remains with in specified range
2645 | Cryo in alarm for lolo level | David.Loik@sanofi.com | 51 NYA | 2023-11-24 | 2023-11-25 | LN2 would be replaced before liquid levels dropped below alarm levels. |  | Deviation
2644 | Leak on inlet clamp of Cz nano filter | zachary.rickard@sanofi.com | 8NYA/1703 | 2023-11-25 | 2023-11-25 | No leaks will be present in the system
2643 | 8NYA Fabrazyme - DGHBE061 - Manifold tube connection not signed for on the record. | stanley.durand@sanofi.com | 8NYA room 1323 | 2023-11-24 | 2023-11-25 | Step performed by operator should be signed for on the record.
2642 | Pompe - RF1NAD002 - 8NYA - Missing verification of manual entry for CEDEX data | michael.knapp@sanofi.com | 8NYA 1403 | 2023-08-14 | 2023-11-11 | All manual corrections have a verifier.
2641 | Verifier comment missing on a manual takeout on Pompe Copper CNC | zachary.rickard@sanofi.com | 8NYA/CNC | 2023-09-18 | 2023-11-24 | If an auto step fails and a manual takeout is needed, the takeout should be performed and a second verifier should make a comment as well. | Yes | No Deviation
2640 | SOLUTION PREP - Media tote transfer line leak during sealing | nicholas.grote@sanofi.com | CNC Space outside solution prep | 2023-11-24 | 2023-11-21 | Good seal, no leaking.
2639 | Cz mannitol 757792 Inventory | zachary.rickard@sanofi.com | 8NYA/1503 | 2023-11-24 | 2023-11-24 | Mannitol will not be expired and any additional stock should be within expiry. | Yes | Deviation
2638 | calculated TSK loading capacity out of range. | Luis.Dominguez@sanofi.com | 51NYA room 132 | 2023-11-23 | 2023-11-23 | Per QU-OPE-0295249 the expected loading capacity is 0.2-0.5 |  | Deviation
2637 | 8NYA - Record needed for Pompe Shipment RF1YDD004, RF1YCD003. Exceptions do not have quantities used for each HU to allocate to the BOM in SAP. | bailey.gleason@sanofi.com | 8NYA | 2023-11-22 | 2023-11-22 | Hold Gemba to see path forward. | Yes | Deviation
2636 | 8 NYA-GEMBA PUMP issue on SP skid | stephen.olszewski@sanofi.com | 8 NYA 8-1323 | 2023-11-22 | 2023-11-22 | Looking for a path forward | No | No Deviation
2635 | 8NYA - RF608D009 PI failed to pull the correct value for Glycol. Glycol was trended, automation was contacted, Additional take out is necessary. | bailey.gleason@sanofi.com | 8NYA | 2023-11-22 | 2023-11-22 | Hold Gemba to see path forward. | No | No Deviation
2634 | 250mL flask equil performed in room 8RM1343 while under action notice | Brandon.Steff@sanofi.com | 8RM1343 8NYA | 2023-11-22 | 2023-11-22 | No operations performed in room while under action notice
2633 | GEMBA - 8 NYA - WFI recirculation line not connected during 705688 buffer batch | nicholas.grote@sanofi.com | 1350 | 2023-11-21 | 2023-11-21 | Connect the line. | Yes | Deviation
2632 | Supplier Devation required - Isolator TSA leaking in bags prior to opening | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2023-11-21 | 2023-11-21 | It is expcted that all media is in a solid form and not have any excess media in an unopened bag.
2631 | 8NYA-Cz- Rf639D018 - Failed seal at the conclusion of ATF A wetting operations | john.lytle@sanofi.com | 8NYA Rm1403 | 2023-11-21 | 2023-11-21 | Use of Sealers result in integral weld | Yes | No Deviation
2630 | GEMBA - 8 NYA - Mannitol Not Scanned into Nano Record | adam.keister@sanofi.com | 8 NYA, 1323 | 2023-11-13 | 2023-11-16 | For all buffers to be accounted for in records, for inventory to be accurate.
2629 | 8NYA Cerezyme - DGHBP026 - MMC POU sample found in sample airlock fridge leaking and beyond expiry. | stanley.durand@sanofi.com | 8NYA room 1410 | 2023-11-17 | 2023-11-21 | Sample should remain intact and not damaged. Sample should also be picked up and delivered to QC for testing.
2628 | HiHi O2 QCA on E4 bioreactor | David.Loik@sanofi.com | 51 NYA | 2023-11-21 | 2023-11-20 | The prompt would be acknowledged real time.
2627 | 8NYA-Cz- MCC X2811 continues to hold for Hi Hi buffer Flow Flow Controller alarms, placing the system in hold. Troubleshooting on the floor identifies no potential cause | john.lytle@sanofi.com | 8NYA Rm1403 | 2023-11-10 | 2023-11-10 | Buffer flow should stabilize after air is worked from the system, and maintain flow rate throughout the completion of production activities
2626 | Chromatogram review, pH out of range SP2 elution (reading 9.14 range 8.1-8.9) | craig.valery@sanofi.com | 51 NYA RM132 | 2023-11-02 | 2023-11-20 | it is expected that the pH remains with in specified range
2625 | Chromatogram review, pH out of range Q regen 1 (reading 8.16 range 7.1-7.9) | craig.valery@sanofi.com | 51 NYA RM131 | 2023-11-11 | 2023-11-20 | it is expected that the pH remains with in specified range |  | Deviation
2624 | Post formulation Osmo value was below target value | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-11-20 | 2023-11-20 | Osmo reading will fall within the range during testing.
2623 | Media Cx value wrongly set as expired | Jacob.Roman@sanofi.com | 8RM1333 | 2023-11-20 | 2023-11-20 | Normally the record correctly sets the media expiry, reads it, and passes the check.
2622 | Nano Filter inlet leak | Juan.Bonilla@sanofi.com | 8NYA/1323 | 2023-11-20 | 2023-11-20 | No leaks | Yes | Deviation
2621 | 8NYA - RF1BHD164 - Compromised bioburden sample | bailey.gleason@sanofi.com | 8NYA | 2023-11-19 | 2023-11-19 | Hold Gemba to see path forward.
2620 | 8NYA - RF1BHD162 - P/N 714043 SUT changed to below target released | bailey.gleason@sanofi.com | 8NYA | 2023-11-19 | 2023-11-19 | Hold Gemba to see path forward.
2619 | 45NYA TSH - SLED Date not populating GZ1411-54 | cindy.guerra@sanofi.com | 45NYA Rm. 190 | 2023-11-17 | 2023-11-17 | SLED date should be populated automatically based on the date of manufacture |  | No Deviation
2618 | Harvest SUM PBM2716001 for Train 2 is leaking material | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2023-11-17 | 2023-11-17 | SUM bag should not be leaning Harvest Material
2617 | 8NYA - RF557D044 Weight received does not match any of the defined weighing range. | bailey.gleason@sanofi.com | 8NYA | 2023-11-17 | 2023-10-27 | Hold Gemba to see path forward.
2616 | Air Supply Lines to BRX Inlet Valves Crossed/Swapped | james.letourneux@sanofi.com | 51 NYA Room 140 | 2023-11-17 | 2023-11-17 | Valves should be connected to their corresponding air supply lines, and actuate in reality as indicated on DeltaV. | Yes | Deviation
2615 | 45 NYA Action Notice 2023-T-044 unable to be located | vincent.palumbo@sanofi.com | 45 NYA | 2023-11-17 | 2023-11-17 | It is expected that stop signs would remain in place at impacted areas until taken down for release. |  | Deviation
2614 | TSH 45NYA Daily Magnahelic Reading not performed | cindy.guerra@sanofi.com | 45NYA | 2023-11-16 | 2023-11-17 | Daily magnehelic reading for rooms would be performed in log sheet |  | Deviation
2613 | MCC Skid 2560 Buffer Flow Controller Low Low alarm | marcelo.dossantos2@sanofi.com | 8NYA/1403 | 2023-11-16 | 2023-11-16 | Recipe should run without holding events.
2612 | Low level QCA received for cryostat 101 | David.Loik@sanofi.com | 45 NYA | 2023-11-15 | 2023-11-15 | Adequate liquid nitrogen would be supplied to the cryo to maintain level |  | Deviation
2611 | BIRF entry missed on material 808708 batch DGHDR001 | michael.montiverdi@sanofi.com | JIT 8NYA/ CR 1131 | 2023-11-14 | 2023-11-16 | BIRF entry filled out in realtime
2610 | Chromatogram review, pH out of range Q regen 1 (reading 8.22 range 7.1-7.9) | craig.valery@sanofi.com | 51 NYA RM131 | 2023-11-09 | 2023-11-16 | it is expected that the pH remains with in specified range |  | Deviation
2609 | GEMBA - 8 NYA - Step not Signed for in EBR | adam.keister@sanofi.com | 8 NYA, 1503 | 2023-11-16 | 2023-11-16 | Path Forward to close out record
2608 | Pompe 3M Sodium Acetate Buffer made without Mixing PQ | stephanie.aceto@sanofi.com | 8NYA Room 1350 | 2023-06-01 | 2023-11-16 | Solutions made for GMP operations are prepared in qualified equipment with validated mixing per QU-SOP-0044843
2607 | WFI Fill Prior to QA Release of AN 2023-I-057 | nicholas.grote@sanofi.com | 1350 | 2023-11-16 | 2023-11-15 | Wait for AN release before starting WFI fill
2606 | 8NYA Fabrazyme - DGHSE021 - Steps not signed for in CIP record. | stanley.durand@sanofi.com | 8NYA room 1323 | 2023-11-12 | 2023-11-15 | It is expected from operator to answer all steps.
2605 | Floor scale allocation issue for Pompe UF-A | zachary.rickard@sanofi.com | 8NYA/1503 | 2023-09-01 | 2023-11-14 | Equipment should be able to be allocated to the record upon being scanned.
2604 | 68NYA Support APU-Metrology-Pressure Standard FBS-0089 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2023-09-09 | 2023-11-15 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2603 | UF TMP/Membrane recovery calculations out of range | craig.valery@sanofi.com | 51 NYA RM88 | 2023-11-15 | 2023-11-15 | It is expected that calculation results fall within range
2602 | DBYFRM051 Fermentation wrong temperature setpoint entered 76NYA Room 364 | dianne.moustafa@sanofi.com | 76 NYA/364 | 2023-11-13 | 2023-11-14 | Step 7.8.8.11 Temperature setpoint expected to be set at 35C.
2601 | RF601D009 DGHAB009-760712--pH meters not calibrated to required standard | leo.gillespie@sanofi.com | 1350 - Solution prep | 2023-11-14 | 2023-11-14 | pH and condo should be tested shortly after mix timer is competed
2600 | TSH - Buffer Prep - Sled date not populating for Goods receipt of Process order RG054D294 | jeremy.wooden@sanofi.com | 45 NYA room 190 | 2023-11-14 | 2023-11-14 | MABC-QU-SOP-0032609-General Manufacturing Athena Transactions
2599 | Post viral Pinch Clamp failed to close on Ross | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1703 | 2023-11-14 | 2023-11-14 | To get a Phath forward
2598 | Framingham Biosurgery-Gaps discovered for Provisc and Discovisc for Japan M1 Dossiers | paul.ward@sanofi.com | 76 NYA, Biosurgery. | 2023-11-01 | 2023-11-14 | See above | Yes | Deviation
2597 | 76-AHU 2A fan failure | john.beaule@sanofi.com | 76-224,227 and 229 | 2023-11-14 | 2023-11-14 | QU-SOP-0051602 | No | Deviation
2596 | DBYHAH006 discrepancy in WFI amount added to V-3501 Pre-purification 76 NYA HA | dianne.moustafa@sanofi.com | 76 NYA/3101 | 2023-11-09 | 2023-11-13 | Discuss, no requirement or calculation. This is an expected amount.
2595 | Fz x2510 mcc skid column 3 outlet leak and mcc wye connection leak to SUM tote | paola.tello@sanofi.com | 8NYA RM 1333 | 2023-11-13 | 2023-11-13 | fbl-ewi-00030
2594 | Gemba – 8NYA – Items not issued to Athena | william.lee@sanofi.com | 8NYA/1350 | 2023-11-13 | 2023-11-13 | Materials should have been consumed in Athena
2593 | 8NYA GEMBA - CZ Nanofiltration low low differential pressure alarm | alex.tello@sanofi.com | 8NYA RM1503 | 2023-11-13 | 2023-11-13 | path forward
2592 | 51 NYA, Approved paperwork used to execute aseptic challenge 09 OCT23 | megan.botti-marino@sanofi.com | 51 NYA | 2023-10-09 | 2023-11-13 | It is expected that effective documents be used.
2591 | 76 NYA PUR DBYHAH006 “Amount of alcohol added is out of specifications (Ist precipitation)" | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2023-11-11 | 2023-11-11 | Per the QU-OPE- 0434886, step 7.7.4.6 the amound of alcohol to be added for first precipitation is +/- 1% of HA solution weight. From DBYHAH006, the weight of the HA is measured to be 1748.5kg. Based on this the range of alcohol to be added should be 1731.0kg - 1766.0kg. However, the tank capacity maxed out at 1730.0kg. This is out of expected range by 1.0 kg.
2590 | 68NYA Support APU-Metrology Vendor certifications may not be meeting 21CFR 211 for proper signatures | robert.bauch@sanofi.com | 68-1830 | 2023-09-01 | 2023-11-13 | it is expected that the person taking the data to have an electronic signature with the date of signature
2589 | WFI was charged into a SUM bag inadvertently. | jermaine.kellam@sanofi.com | 8RM1350 | 2023-11-12 | 2023-11-12 | HES should have been charged into the SUM per procedure.
2588 | MOD Enzyme record (DGHAS011) associated with shipped Cerezyme lot #DGHDS009 is still in draft status and is not included in genealogy report. | ana.salichs@sanofi.com | N/A | 2023-11-12 | 2023-11-12 | Record should be reviewed prior to shipment and be included in the genealogy report.
2587 | GEMBA - 8 NYA - MES BOM does not match EWI step | adam.keister@sanofi.com | 8 NYA, 1503 | 2023-11-12 | 2023-11-12 | Path Forward, which SUT to use, also for EWI and BOM to align
2586 | Cerezyme - RF639D019 and RF658D017 - 8NYA - Eluate and End of Run Bioburden Samples Delivered Expired | michael.knapp@sanofi.com | 8NYA 1403 | 2023-11-09 | 2023-11-11 | Samples are not expired prior to arriving for QC receipt.
2585 | 8NYA-705661 50L. Recipe calls for 100L SUM but BOM calls for 200L SUM bag | Tanner.Raiburn@sanofi.com | 8NYA-1350 | 2023-11-11 | 2023-11-11 | Find a path forward, Updated BOM, New MO released
2584 | Cz Phenyl Additional SUT comment missing HU | zachary.rickard@sanofi.com | 8NYA/1503 | 2023-10-11 | 2023-11-10 | When making a comment on the record about additional SUT being needed all information must be presented.
2583 | DGHBY015 Cerezyme CM Load Filter FIT Recipe Failed to Launch in MES | wilfredo.lopez@sanofi.com | 8NYA Room 1290 | 2023-11-10 | 2023-11-10 | It is expected that the correct filter testing recipe launches through MES to DeltaV based on the CX values pulled in and that the recipe is found in the system.
2582 | 8NYA-Cz-X2220 SUB Bag (P/N 753288) failed Visual Inspection prior to installation | john.lytle@sanofi.com | 8NYA Rm1403 | 2023-11-10 | 2023-11-10 | All SUT used for processing be integral
2581 | GEMBA 51NYA: failed FIT | adam.rowe@sanofi.com | 51NYA/140 | 2023-11-10 | 2023-11-10 | Test should pass
2580 | Cz Q-membrane manual entry missing verification signature | zachary.rickard@sanofi.com | 8NYA/15 | 2023-10-12 | 2023-11-10 | Any time a manual entry is required a second operator will put a verification comment on the record.
2579 | Operator experiencing nosebleed while in clean room space. | jacob.kallon@sanofi.com | 51 NYA/room 90 | 2023-11-10 | 2023-11-10 | Per QU-SOP-006952 Step 6.1.2: • All personnel including visitors and contractors shall report to supervisory personnel any health conditions (i.e. apparent illness or open lesions) that may have an adverse effect on drug products or cause the shedding of abnormal numbers or types of contaminants (see QU-SOP-0065573, OHS-SPI-044 Communicable Disease Policy). • Individuals identified with any such health condition shall be excluded from entering the controlled areas. This exclusion must remain in effect until the condition is corrected.
2578 | 45 NYA Action Notice 2023-T-042 released without all of the signs returned. | vincent.palumbo@sanofi.com | 45 NYA | 2023-11-04 | 2023-11-06 | It is expected that the same number of signs issued would be found at their posted locations to be taken down by QA during release process. |  | Deviation
2577 | Cz Q Membrane- Flowrate could not be achieved | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2023-11-09 | 2023-11-09 | Determine path forward for Q processing
2576 | Issue with Athena to move H12 (LOT DGF2B16) from Blocked to unrestricted. | Subhashini.Satishkumar@Sanofi.com | 51NYA | 2023-11-09 | 2023-11-09 | It is expected Athena code would allow to move H12 from Blocked to Unrestricted per the SOP
2575 | 76 NYA PUR DBYHAH005 “SWECO screen found damaged before installation". | munish.kumar2@sanofi.com | 76 NYA/Room#3103/Purification | 2023-11-09 | 2023-11-09 | Per FBS-SOP-000651 step# 6.7.4 the screen should not have any dents, creases or tears.
2574 | GMP Use of ineffective batch record (706930) | kyle.tesorero@sanofi.com | 8NYA, 1333 | 2023-10-03 | 2023-11-09 | v7 would have been aborted and v8 released by Scheduling and utilized for operations | Yes | Deviation
2573 | 8NYA GEMBA FZ Small Scale DGHSD010 EM Plate Not Scanning During Initial Equil | gary.dodakian3@sanofi.com | 8NYA inoc prep room | 2023-11-08 | 2023-11-08 | Initial plate should be able to scan and document plate.
2572 | 768307, DBYFRM050, QU-OPE-0309296, pH on chart out of range | dianne.moustafa@sanofi.com | 76 NYA/364 | 2023-11-08 | 2023-11-08 | pH stay within range of plus or minus 0.3 of setpoint. Setpoint was 7.50 | Yes | Deviation
2571 | Mixer,1500 failed to initiate initial WFI fill | Evans.kamau@sanofi.com | 8NYA1350 | 2023-11-08 | 2023-11-08 | Initial WFI would fill after tare prompt is answered.
2570 | HA - No lots - 76 NYA QA Records Room - Missing GMP documentation | charlie.houle@sanofi.com | 76 NYA; QA record room (M18) | 2020-01-01 | 2023-11-08 | It is expected that Attachment 3 should be maintained in the logbook. It is expected that issued papework is returned and recorded appropriately on the logbook.
2569 | Late entries for the trainer observation | nicole.lufkin@sanofi.com | 51NYA 140 | 2023-09-04 | 2023-11-07 | a Trainer Observation can only be made as a late entry if there is a verifiable GMP source to support it
2568 | 1x Facility cleaning incorrectly documented at 51NYA | nicholas.shoenfelt@sanofi.com | 51NYA rooms 149, 152, 152B, 153, 130, 131, 130A, 130B | 2023-11-07 | 2023-11-08 | It is expected that 1x cleanings are documented in the special cleaning logbook and that all cleaning solutions prepared are documented in the solution preparation worksheet. |  | Deviation
2567 | Gemba-8NYA-Downstream-Manual entry of material information missing HU and no verification | aidan.mula@sanofi.com | 8NYA RM1323 | 2023-11-07 | 2023-11-08 | SOP should be followed for manual entry of SUT data with verification from second operator.
2566 | Testing volume for sample 836736 was not recorded during analysis. | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2023-09-03 | 2023-11-03 | Per QU-SOP-0073109, a 10mL test volume must be filtered for sample 836736 (FZ2_CB_762729_EQUIL_BIOB).
2565 | Cerezyme - RF639D018 - 8NYA - Pull tab tear during ATF Changeout | michael.knapp@sanofi.com | 8NYA 1403 | 2023-11-08 | 2023-11-08 | There are no tears on the ATF pull tab membranes.
2564 | Missing BioLims sample | thomas.loik@sanofi.com | 8RM1350 | 2023-10-01 | 2023-11-02 | Utilize Down Stream testing
2563 | Power outage 51NYA - Temporary loss of AHU-7 | nicholas.shoenfelt@sanofi.com | 51 NYA Small Scale, and Purification | 2023-11-07 | 2023-11-07 | It is expected that AHU will remain running.
2562 | GEMBA 51NYA: failed FIT | adam.rowe@sanofi.com | 51NYA/140 | 2023-11-07 | 2023-11-07 | 1 of the tests should have passed
2561 | Incorrect weight used in the daily check of Balance 85-QC-0327 | paul.consolmagno@sanofi.com | 76NYA, Room 107 | 2023-10-25 | 2023-11-07 | 52g would have been used
2560 | Nano filter leak during Nanofiltration | Juan.Bonilla@sanofi.com | 8NYA/1503 | 2023-11-07 | 2023-11-07 | For the Nano filter to not leak during processing.
2559 | Ingoing Q column equilibration buffer resampled and retested. | craig.valery@sanofi.com | 51 NYA | 2023-11-04 | 2023-11-07 | It is expected that resampling is not to be performed without approved justification. |  | Deviation
2558 | Buffer lot number not recorded in SP column packing Athena transactions | craig.valery@sanofi.com | 51 NYA | 2023-11-28 | 2023-11-06 | It is expected that all buffers used be recorded to operation process order.
2557 | 45 R&D: worksheets pulled of saved from Gibraltar used to execute EM | megan.botti-marino@sanofi.com | 45 NYA | 2023-10-16 | 2023-11-07 | That documents and assay worksheets be pulled from Qualipso.
2556 | 8NYA APU-Room 1333-Speed Indicator Controller FB-2936-002 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1333 | 2023-11-07 | 2023-11-07 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2555 | 8NYA APU-Room 1333-Weight Indicator FB-2937-004 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1333 | 2023-11-03 | 2023-11-07 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2554 | 51NYA GEMBA: Low pressure on E3 | adam.rowe@sanofi.com | 51NYA/140 | 2023-11-06 | 2023-11-06 | Pressure should be back at setpoint post-refeed. There should be no alarms or QCA callouts.
2553 | 8NYA Cerezyme - CM DGHBY013 - NRSR for pre-filtered CM eluate not taken. | stanley.durand@sanofi.com | 8 NYA 1503 | 2023-11-01 | 2023-11-06 | All NRSRs for MSAT should be taken per FBl-EWI-000009/ QU-WIN-0046060 | Yes | No Deviation
2552 | 8NYA - Pompe UF-A - DGHUA004 - Manual entry of SUT performed without second verifier. | stanley.durand@sanofi.com | 8 NYA Room 1503 | 2023-08-30 | 2023-11-06 | It's expected per FBL-SOP-001320/QU-SOP-0019707 that all manual entry in EBR requires a second verifier.
2551 | 8 NYA - Media POU DGHAW148 sample expired | cleiton.lopes@sanofi.com | 68NYA QCSM 4720 | 2023-11-05 | 2023-11-06 | Bioburden to be tested within 24hrs
2550 | No product - No ltos - 8 NYA - EM/CU samples inadvertently assigned to incorrect plan | craig.stevenson@sanofi.com | 8NYA | 2023-11-04 | 2023-11-04 | It is expected that 8NYA FG does not schedule QCM EM samples
2549 | Gemba – 8NYA – Missing samples for CZ during record review | william.lee@sanofi.com | 8NYA/1350 | 2023-11-05 | 2023-11-05 | Samples should have been taken and received by QC.
2548 | 8NYA GEMBA - FZ Butyl BOM items not used in process. | alex.tello@sanofi.com | 8NYA RM1323 | 2023-11-05 | 2023-11-05 | path forward
2547 | 8NYA GEMBA- FZ Butyl Eluate SUM out of calibration | alex.tello@sanofi.com | 8NYA RM1330 | 2023-11-05 | 2023-11-05 | calibration to be within range
2546 | Gemba – 8NYA - Fz 500ml sampling result/Possible contamination | Yves.Antoine@sanofi.com | 8NYA, Room 1333 | 2023-11-04 | 2023-11-04 | We expected the Fz 500ml sampling result to be within range. FBL-EWI-00009, QU-OPE-0028959.
2545 | Q membrane Load sum not mixing | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1703 | 2023-11-03 | 2023-11-03 | Recipe shoul mix for 15 minutes prior to sample
2544 | Gemba 8NYA-Pompe Downstream UFDF-B-SUT not taken out in process | aidan.mula@sanofi.com | 8NYA RM1503 | 2023-11-03 | 2023-11-03 | Determination if part required
2543 | Cz Freeze Dispensing Record comparing fill weight to fill volume range | catherine.kuzdzal@sanofi.com | 8 NYA Suite 2 Post Viral | 2023-10-17 | 2023-11-02 | It is expected that volume and/or weight ranges are compared accurately.
2542 | No initial incubation time documented for routine EM at 76NYA R&D | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2023-10-14 | 2023-11-03 | It is expected that the incubation time is recorded temporaneously at the time of when the action is performed.
2541 | CSE vial Lot # differs from Packaging label | lauren.ruffing@sanofi.com | 68NYA 4700 QC Micro | 2023-11-02 | 2023-11-02 | CSE reagent in CSE box matches what teh box label's claim
2540 | 8NYA-DOWNSTREAM-Fz-DGHDT011/RF1W5D011- material expires | MARC.LATTA@SANOFI.COM | 8NYA/1700 | 2023-11-03 | 2023-11-03 | SU should not be expired prior and per operation
2539 | Skid X3451 went on hold in Delta V at end of post use Nano filter wetting for pressure differential. | jim.ramjattan@sanofi.com | 8NYA, Room 1503 | 2023-11-03 | 2023-11-03 | Per FBL-EWI-000021 recipe should end successfully.
2538 | Cz Freeze bag underfilled | zachary.rickard@sanofi.com | 8NYA/1703 | 2023-10-17 | 2023-11-02 | Freeze bags will be filled within the acceptable range of 7.3-7.7kg
2537 | Leak on bleed valve of CZ nanofilter during equilibration and nanofiltration | zachary.rickard@sanofi.com | 8NYA/1503 | 2023-11-02 | 2023-11-02 | Leaks will not occur
2536 | GEMBA 51NYA: calculation error | adam.rowe@sanofi.com | 51NYA/152 | 2023-10-25 | 2023-11-02 | Calculation should = 1615
2535 | Cerezyme - RF639D020 - 8NYA - SUB Weight Alarm | michael.knapp@sanofi.com | 8NYA 1403 | 2023-11-02 | 2023-11-02 | Weight does not drift | Yes | No Deviation
2534 | GEMBA - 8 NYA - Nano Recipe on Hold Due to SUM EM Failure | adam.keister@sanofi.com | 8 NYA, 1503 | 2023-11-02 | 2023-11-02 | Path Forward
2533 | No takeout performed for part number 813903 (3/8" Mannifold). | zachary.rickard@sanofi.com | 8NYA/1703 | 2023-10-17 | 2023-11-02 | Material that is required for processing will be taken out against the BOM.
2532 | 76 NYA PUR DBYHAM006 “Amount of precipitate remaining after first alcohol wash out of specification ". | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2023-11-02 | 2023-11-02 | Per QU-OPE-0202445, step 7.12.29 the amount of precipitate remining should be between 280-320 kg
2531 | 8 NYA - DGHSE019 - Eluate Adjust Buffer Clamped During Elution of FZ SP Production. | stanley.durand@sanofi.com | 8 NYA Room 1323 | 2023-11-01 | 2023-11-01 | Eluate pH adjust should be unclamped during elution.
2530 | The Integrity Test for the vent filters located in 76 NYA R&D Tank and the Mueller Still were not performed | brizaida.medina-vega@sanofi.com | 76 NYA R&D Tank and the Mueller Still (currently decommissioned). | 2021-06-28 | 2023-11-01 | Integrity Test performed and report attached to the WO 2747142.
2529 | 76 Mfr. WFI Tank 305 Vent Filter Post Integrity Test Failed | brizaida.medina-vega@sanofi.com | 76-WFI Manufacturing Tank 305 | 2023-08-22 | 2023-11-01 | Vent Filter pass the integrity test and the bubble point must be burst at 18 psig or more.
2528 | 1 Year Preventive Maintenance of 8 NYA Chromatography Skid X-3050001 was not performed by Extended Due Date | kevin.southwick@sanofi.com | 8 NYA, Room 1323 | 2023-11-01 | 2023-11-01 | Infor EAM PM Code 80164
2527 | Cerezyme - RF639D019 - 8NYA - Bleed Pressure Indicator ModBad | michael.knapp@sanofi.com | 8NYA 1403 | 2023-11-01 | 2023-11-01 | There is no ModBad alarm on the bleed system
2526 | 8NYA-Fz-SUB X2200 pCO2 meters trend as having potential errors during in-process calibrations | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-11-01 | 2023-11-01 | pCO2 adjustment executed per ewi
2525 | 68NYA- Support APU-QCM Department-Room 4700-FB-2703 Hygrometer with no as found data | robert.bauch@sanofi.com | 68-4700 | 2023-07-28 | 2023-11-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2524 | 76NYA- Biosurgery APU-HAFM Room 364-Pressure Gauge F-18691 overdue for calibration | robert.bauch@sanofi.com | 76-364 | 2023-11-01 | 2023-11-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date
2523 | 8NYA APU-Room 1333-6 items overdue for calibration on 2 product mixer bag skids FB-2936 and FB-2937 | robert.bauch@sanofi.com | 8-1333 | 2023-11-01 | 2023-11-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date | Yes | Deviation
2522 | 8NYA APU-Room 1403-41 items on MCC Chromatography Skid X-2811001 overdue for calibration | robert.bauch@sanofi.com | 8-1403 | 2023-11-01 | 2023-11-01 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date | Yes | Deviation
2521 | 8NYA GEMBA- FZ Butyl ammonium sulfate SUM displaying incorrect weight | alex.tello@sanofi.com | 8NYA RM1330 | 2023-10-31 | 2023-10-31 | path forward | Yes | Deviation
2520 | defective Shipping Bags | Christian.Venegas@Sanofi.com | 8NYA Downstream suite 1 | 2023-10-31 | 2023-10-31 | To find a Phath forward to continue with post viral operations
2519 | Cold Room 1322: Floor storage location near freezer door temperature was outside the acceptance criteria of 2 to 8C | ernest.annancy@sanofi.com | Cold Room 1322 | 22-26 Sep 2023 | 2023-09-26 | Per 2023 RQ Protocol 18-0685-P, the temperature must be between 2-8C.
2518 | 76NYA Biosurgery APU-Room 3103-F-6938-027 Vacuum indicator on panel view found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 76-3103 | 2023-10-31 | 2023-10-31 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2517 | 8NYA- GEMBA- Nano Osmo out of range | alex.tello@sanofi.com | 8NYA RM1503 | 2023-10-30 | 2023-10-30 | Osmo result to be within range
2516 | 51NYA TSH - Q - DGF4B10A Percent Recovery Out of range | cindy.guerra@sanofi.com | 51NYA - 131 | 2023-10-30 | 2023-10-30 | Recovery result would be between 72% to 110%
2515 | 51NYA TSH - Q pH Regen 1 out of range - DGF4B10A / DGFB10B | cindy.guerra@sanofi.com | 51NYA - 131 | 2023-10-20 | 2023-10-30 | pH should be in range during process 7.1-7.9
2514 | Incorrect entry made for WFI rinse on logbook QU-FOR-0027702 | Juan.Ramos-Torres@sanofi.com | 76 NYA room 3104 | 2023-09-07 | 2023-10-30 | Select pass or fail after CIP. If performing a different activity select N/A
2513 | Weekly cleaning not performed due to autoclave failuer | valeria.moreno@sanofi.com | 76-355 | 2023-10-27 | 2023-10-27 | Step 6.1.1 from QU-SOP-0051603. When no HA Fill is performed in Room 355, only weekly Sanitization is required.
2512 | 8NYA- Fz- X2200 pCO2 probe 1 diverging from probe 2, causing system hold | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-10-29 | 2023-10-29 | It is expected that the 2 probes used to monitor gas mix in the sub do not diverge by more than their allowed setpoints
2511 | Fabrazyme - DGHCE005 - 8NYA - Media Line SUT not properly identified and taken out | michael.knapp@sanofi.com | 8NYA 1333 | 2023-08-18 | 2023-10-28 | The correct part is taken out against BOM immediately prior to use and is identified. Manual entries have a verification.
2510 | GEMBA - 8 NYA - SUM hi temp alarm during processing | adam.keister@sanofi.com | 8 NYA, 1323 | 2023-10-28 | 2023-10-28 | SUM to remain in temp during transfer | Yes | Deviation
2509 | Unable to release product intermediate in ERP | timothy.ryley@sanofi.com | 08 NYA | 2023-10-27 | 2023-10-27 | QA should perform Section 6.1 of MABC-QU-SOP-0025901 | Yes | Deviation
2508 | 8NYA- Cz Post Use Nanofilter wetting recipe going into hold | alejandro.martinez2@sanofi.com | 8NYA, Downstream PS Ballroom | 2023-10-28 | 2023-10-28 | Considering 500L has been flushed through filter, proceed with FITing | Yes | Deviation
2507 | 51NYA TSH SP2 Column operation pH out of range during load and elution | Bruno.darosa@sanofi.com | 51NYA/ RM132 | 2023-10-21 | 2023-10-27 | MABC-QU-SOP-0061107 attachment 6 load range 5.4-6.0/elution range 8.1-8.9
2506 | GEMBA - 8 NYA - leak from pump insert on nano SUT | adam.keister@sanofi.com | 8 NYA, 1503 | 2023-10-27 | 2023-10-27 | To not have SUT leak. Replace SUT and restart process.
2505 | 8NYA-Cz- DGHBN012 ATF B "A2B Vent Filter Line" Vent Filter found installed with incorrect flow direction | john.lytle@sanofi.com | 8NYA Rm1403 | 2023-09-26 | 2023-10-18 | It is expected that the orientation of the A2B Vent filter is such that allows suitable airflow during wetting operations while remaining integral. The A2B Vent Filter utilizes unidirectional flow |  | No Deviation
2504 | 8NYA CZ CV MCC Skid X2560 DGHAA030 bioburden expired | judith.keller@sanofi.com | 68 NYA 4700 QCM | 2023-10-26 | 2023-10-26 | It is expected that bioburden samples are tested before their 24 hour expiry
2503 | Weigh and Dispense HU would not scan into formulation record--error message | leo.gillespie@sanofi.com | 1350 - Solution prep | 2023-10-26 | 2023-10-26 | Record should accept the HU scan--unclear what the root cause of the error was
2502 | Butyl skid pH and Condo readings out of range during equal. | paul.devaney@sanofi.com | 8NYA/1323 | 2023-10-27 | 2023-10-27 | pH reading should be between 6.8 and 7.5 and Conductivity should be between -1mS/cm and 1 mS/cm.
2501 | 8NYA GEMBA - FZ Butyl Eluate SUM over temperature range | alex.tello@sanofi.com | 8NYA RM1323 | 2023-09-09 | 2023-10-26 | for operating temperature to be 16-24
2500 | GEMBA for Cz MCC skid X2811 column high temperature | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-10-26 | 2023-10-26 | column temperature stays below alarm limit of 24C
2499 | Two Monthly PMs Not Completed by Due Date of 30Sep2023 (Autoclave WO 3017168; Parts washer WO 3017139) | tanni.arif@sanofi.com | 45NYA | 2023-10-01 | 2023-10-24 | Complete all PM tasks prior to PM closure (otherwise approved by QA) or a deviation is required per site specific SOP.
2498 | TSH Blue DGF2B12 - Incorrect Wash 2 Buffer total | cindy.guerra@sanofi.com | 51-138 | 2023-09-24 | 2023-10-25 | Total volume would be between the min and max of step 4.5.3.1
2497 | 68NYA Support APU-Room 4700-QCM68-two TOC Analyzers found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 68-4700 | 2023-10-23 | 2023-10-26 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2496 | 8NYA - RF1YID030 Insufficient second verify. | bailey.gleason@sanofi.com | 8NYA | 2023-10-26 | 2023-10-26 | Hold Gemba to see path forward.
2495 | Multiple unused RODACs discovered with growth. | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2023-10-25 | 2023-10-25 | It is expected that RODACs do not have growth on them prior to use, fresh out of the sterile bag.
2494 | 51 NYA Recirculation Apparatus not procured following QU-SOP-0011047 | stephen.peglow@sanofi.com | 51 NYA / 45 NYA Solution Prep | 2023-10-25 | 2023-10-25 | Materials used in GMP manufacturing are onboarded per Sanofi requirements and procedures
2493 | GEMBA for Harvest Vessel PBM-2916001 not mixing or controlling temp (Cz) | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-10-24 | 2023-10-25 | Harvest Mixes at 16CPM, controls to 17.5C
2492 | 51NYA -Thyrogen APU-Cold Room 138-EMS F-3933-300 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 51-138 | 2023-10-24 | 2023-10-25 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2491 | 8NYA- Cerezyme ROSS fill bag underfilled | stanley.durand@sanofi.com | 8NYA Room 1703 | 2023-10-24 | 2023-10-24 | It's expected during ROSS fill that all bags reach the minimum filled weight 7.400kg
2490 | Cerezyme - DGHBN018 - Harvest SUM Leak | michael.knapp@sanofi.com | 8NYA 1403 | 2023-10-24 | 2023-10-24 | There is no leak from the harvest bag.
2489 | Equilibration temperature out of range found, during chromatogram review. | craig.valery@sanofi.com | 51 NYA | 2023-10-11 | 2023-10-23 | Temperature expected to remaing with int 6-10C
2488 | GEMBA - 8 NYA - Shipping Bag Inspection Failure | adam.keister@sanofi.com | 8 NYA, 1703 | 2023-10-24 | 2023-10-24 | Bags to pass inspection
2487 | 8NYA-Fz- Incorrect pCO2 adjustment to SUB caused system hold | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-10-23 | 2023-10-23 | pCO2 adjustment executed per ewi
2486 | GEMBA - 8 NYA - Slow leak inlet bleed valve of nano filter | adam.keister@sanofi.com | 8 NYA, 1503 | 2023-10-23 | 2023-10-23 | Decision on path forward, to forward process with product or not
2485 | Incubation Batch INC-08500 (202303856-1) missing some TAC/Mold results for counting data in biolims | amanda.lavalley@sanofi.com | QC Micro Lab 68 NYA Room 4700 | 2023-10-11 | 2023-10-23 | Per Step 6.3.5 in QU-SOP-0019490, the data for plate counts should be entered and saved at the time of counting.
2484 | 8 NYA GEMBA Request- Non managment reviewed department walk through | richard.manasian@sanofi.com | 8 NYA | 2023-10-23 | 2023-10-23 | Delegate forms would be on file or walkthrough would be signed by management
2483 | RQ - F076-VSL-V-3511-76NYA-CQ-0002 - Failing Stage 3 Conductivity Results | charlie.houle@sanofi.com | 76 NYA, Room 364 (HA Fermentation); QC Micro | 2023-10-20 | 2023-10-20 | It is expected the conductivity result would meet specification. The specification called for is "Meets Requirements of USP <645>", which is 2.1 us/cm at pH 6.6.
2482 | Thyrogen - SP1 Lot DGF3B15A - Low pH found during chromatogram review | craig.valery@sanofi.com | 51 NYA | 2023-10-22 | 2023-10-22 | It is expected that pH should remain in range per attachment 3 of QU-SOP-0061107 |  | Deviation
2481 | Air Manifold | Jonathan.ahiaba@sanofi.com | 8 NYA | 2023-10-22 | 2023-10-22 | SUT should have come with zip tie attached
2480 | pH adjust prompt answered as no incorrectly in delta v when further ph adjust was required. | Brendan.Dickhaut@sanofi.com | 8NYA/Room 1503 | 2023-10-21 | 2023-10-21 | Expected results are that operators correctly answer prompt to align with the needs of recipe.
2479 | GEMBA - 8 NYA - Failure on Q Membrane Recipe due to WFI running out | Brendan.Dickhaut@sanofi.com | 8NYA/Room 1503 | 2023-10-21 | 2023-10-21 | Expected result that enough buffers remain to finish the process being run
2478 | TSH - Buffer Prep - Lot number (DGFB2065) and tank used (V-1501) for the process do not align. | Jeremy.Wooden@Sanofi.com | 45 NYA Room 190 | 2023-10-14 | 2023-10-21 | It is expected that V-1502 would be used in conjunction with lot number DGFB2065. MABC-QU-FOR-0103150 dictates that a lot number containing B1 would be used for tank V-1501 and B2 would be used for tank V-1502. |  | Deviation
2477 | 8NYA-DOWNSTREAM-Cz-Q load-DGHBV020-Agitator did not turn on during mixing time after WFI dilution | marc.latta@sanofi.com | 1503 | 2023-10-21 | 2023-10-21 | mixer should turn on and run through the mixing time until prompt is answered
2476 | Q membrane Record and EWI 34 does not align | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1503 | 2023-10-21 | 2023-10-21 | Need Phath forward to continue with production record
2475 | MCC skid 2543 failed to change status at end of run | marcelo.dossantos2@sanofi.com | 8NYA/1340 | 2023-10-21 | 2023-10-21 | MCC status should have changed from sterile in use to soiled.
2474 | 8 NYA GEMBA - CZ MMC Buffer line disconnected prior to storing skid | alex.tello@sanofi.com | 8NYA RM1503 | 2023-10-21 | 2023-10-21 | Line to remain connected to the skid then answering "Remove devices" prompt
2473 | 8NYA-Cz-Downstream Zinc-Eluate SUM bag failed visual inspection | aidan.mula@sanofi.com | 8NYA RM1503 | 2023-10-20 | 2023-10-20 | Determination at Gemba, new bag was sourced to proceed | Yes | No Deviation
2472 | Wood pallets discovered in hallways between FBS and R&D | craig.stevenson@sanofi.com | 76 | 2023-10-19 | 2023-10-19 | No untreated wood in the facility
2471 | 8NYA - RF555D007 705681 Tote chill step was not executed | bailey.gleason@sanofi.com | 8NYA | 2023-10-20 | 2023-10-20 | Hold Gemba to see path forward. | Yes | No Deviation
2470 | 8NYA-Cz-Downstream MMC- Leak on eluate bag at elbow | aidan.mula@sanofi.com | 8NYA RM1503 | 2023-10-20 | 2023-10-20 | Determination on path forward | Yes | Deviation
2469 | NRSR are not commented in the Batch records | daniel.bati@sanofi.com | 8NYA all department | 2023-10-19 | 2023-02-22 | FBL-SOP-000010
2468 | 8NYA-Cz-damage found prior to install/ ATF Permeate Line Extension during ATF Setup operations | john.lytle@sanofi.com | Rm 1403/ 8NYA | 2023-10-19 | 2023-10-19 | SUT parts are recieved integral from the vendor, with no visible signs of damage
2467 | Determine if Supplier event is needed, related to CAPA FRAD23E0285/ FRAD23A0363 | kyle.tesorero@sanofi.com | 8NYA, 1323 | 2023-10-19 | 2023-10-19 | N/A
2466 | 8 NYA GEMBA Request- DGHDT014 Trainer signed as trainer and verifier for same process | richard.manasian@sanofi.com | Phase 1 Downstream ballroom | 2023-09-14 | 2023-10-18 | Verifier would be trained on step that they are signing for | Yes | Deviation
2465 | Live house fly was found inside the men's gowning Room-76-205. | munish.kumar2@sanofi.com | 76-205 men's gowning room | 2023-10-19 | 2023-10-19 | : Refer to QU-SOP-0005611 steps 6.2.2 and 6.3.1. Per the SOP, the classified area should be free of pest activities.
2464 | 8NYA-Gemba- Clean Equipment Prep 8RM1373- Leak in ceiling observed | alejandro.martinez2@sanofi.com | 8RM1373 | 2023-10-19 | 2023-10-19 | Find root cause of leaking medium pressure steam line in the interstitial space, no more dripping in room 1373 | Yes | Deviation
2463 | Remediation steps not documented in batch record | timothy.ryley@sanofi.com | 8NYA1503 | 2023-10-16 | 2023-10-19 | Remediation activities need to be documented at the time of execution.
2462 | QCA Refeed Timer exceeded on E4 | david.loik@sanofi.com | 51 NYA | 2023-10-19 | 2023-10-19 | Refeed would be started within the time limits
2461 | GEMBA - 8 NYA - Pump Failure on Q Membrane Recipe. | adam.keister@sanofi.com | 8 NYA, 1503 | 2023-10-18 | 2023-10-18 | Path forward to complete recipe. Skid is all disposable SUT, no store required. There is one filter attached to skid that needs to get tested.
2460 | Incorrect Stability Concentrations used for Determination of Aggregate Assay for newGAA | sarah.adams@sanofi.com | 68 NYA, QCC | 2023-09-20 | 2023-10-18 | Correct concentrations used for assay execution.
2459 | Record is looking to install media during drain and skipping POU sampling; media already installed and | Jeffrey.boldizar@sanofi.com | 8 NYA 1403 | 2023-10-16 | 2023-10-19 | Installed into the appropriate record
2458 | 8NYA-Cz-X2220 Perfusion failure/ Investigation | john.lytle@sanofi.com | 8NYA Rm1403 | 2023-10-18 | 2023-10-18 | continuous perfusion
2457 | GEMBA - 8 NYA - Stock Creation Error | adam.keister@sanofi.com | 8 NYA, 1503 | 2023-10-18 | 2023-10-19 | To have SU for Q stock correctly populated with 124.2KG (SUM weight) and CX value updated accordingly.
2456 | DGHAS012 - 760739 MOD Enzyme Buffer Incorrect Stock Creation Expiry Assigned by MES | wilfredo.lopez@sanofi.com | 8NYA Room 1350 | 2023-10-18 | 2023-10-18 | Based on the specification sheet (FBL-SPC-000202), the MOD Enzyme material has a storage condition of <24 hours at ambient temperature. It was expected that the assigned expiry would be 24 hours from the time of batching completion. Having the incorrect expiry date and time assigned could lead to the use of expired material.
2455 | Fz Butyl- Eluate Inventory sample incorrectly signed for | nathaniel.barnes@sanofi.com | 8NYA 1323 | 2023-10-17 | 2023-10-18 | BioLIMS support needed to recreate sample as well as cancel the original sample. QC also will need to be contacted as they likely have a bag that is mislabeled as an eluate sample when it is in fact an End of Hold sample
2454 | Atypical fraction 4 designation of do not pool. A280/ELISA results trending low | craig.valery@sanofi.com | 51 NYA | 2023-10-18 | 2023-10-18 | Typically fractions 2-7 are pooled
2453 | Thyrogen sample located in 8NYA sample fridge | ana.salichs@sanofi.com | 8NYA RM 1410 | 2023-06-27 | 2023-10-16 | Samples should be adequately stored and transported to appropriate locations. | Yes | No Deviation
2452 | Thyrogen - Q Lot DGF1B09B - pH found out of range high during Q chromatogram review | craig.valery@sanofi.com | 51 NYA | 2023-10-18 | 2023-10-18 | All chromatogram parameters are expected to be within defined ranges per attachment 3 of QU-SOP-0061107
2451 | 8NYA-Gemba Mod Enzyme Cocktail Hold Date Not Okay | Tanner.Raiburn@sanofi.com | 8NYA-1350 | 2023-10-18 | 2023-10-18 | Find a path forward
2450 | 8NYA-Fz- discuss switching to back up vent filter set on SUB X2200 due to foaming | John.Lytle@sanofi.com | Rm 1333/ 8NYA | 2023-10-18 | 2023-10-18 | It is expected that foam stay below 2" per FBL-EWI-000011 SUB, Daily Monitoring
2449 | 8NYA APU-Room 1323-Pressure indicator FB-2950-002 on system X-3091001 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 8-1323 | 2023-10-18 | 2023-10-18 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2448 | 8NYA: Approved forms used for EM execution by Manufacturing on 28SEP23 | megan.botti-marino@sanofi.com | 8 NYA | 2023-09-28 | 2023-10-18 | It is expected that effective versions of documents be printed and used.
2447 | Phenyl Skid CIP record steps not executed contemporaneously. | timothy.ryley@sanofi.com | 8RM1503 | 2023-10-16 | 2023-10-17 | Rinse samples and Skid should be visually inspected.
2446 | material label found in 8RM 1323 with no recorded take out of item | Daniel.Sciotti@sanofi.com | 8NYA room 1323 | 2023-10-16 | 2023-10-16 | FBL-SOP-000942-Good Documentation Practices, FBL-SOP-001320 Inova Execution at 8NYA
2445 | Braided hose failure causing leak during Media transfer. SUT part number 713975 | Brendan.Dickhaut@sanofi.com | 8NYA, Room 1350 | 2023-10-17 | 2023-10-17 | Tubing remains intact and does not fail during process.
2444 | Stock creation did not populate in Athena from MES | Brendan.Dickhaut@sanofi.com | 8NYA/ Room 1703 | 2023-10-16 | 2023-10-16 | Expected results is that matching stock is created in Athena whenever stock is created in MES.
2443 | 8NYA CzNano Downstream Nano filter has a small leak from the bleeder valve | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1503 | 2023-10-16 | 2023-10-16 | To have a Phath forward to continue Nano process
2442 | 8NYA GEMBA - CZ Phenyl Skid CIP requiring probes | alex.tello@sanofi.com | 8NYA RM1503 | 2023-10-16 | 2023-10-16 | path forward
2441 | 8NYA-DOWNSTREAM-Cz-Nanoformulation-DGHDS010- Missing TEM MONITORING TUBE SET | marc.latta@sanofi.com | 1503 | 2023-10-16 | 2023-10-16 | COnnect to post viral SUM must have temperature monitoring tube set per FBL-EWI-000029
2440 | Pipette FB-3494 Failed as found calibrations | Nicholas.grote@sanofi.com | 8 NYA | 2023-10-11 | 2023-10-13 | Pass as founds; last passing calibration was 25Jul23. Pipette has not been used for mfg operations since 16May23 per 2023LBG001
2439 | TSH buffer - 1411 hood batch - operator scanned Right and Left Operations Passive air monitoring settle plate labels in error when monitoring set up. | Jeremy.Wooden@Sanofi.com | 45 NYA Room 190 | 2023-10-16 | 2023-10-16 | That the correct set up labels would have been scanned.
2438 | 8 NYA- Point of use step skipped. | Jose.lugo2@sanofi.com | 8 NYA CNC | 2023-10-16 | 2023-10-16 | N/A
2437 | 51NYA - TSH Fraction 7 Lot DGF3B11B ID 2300016235 rejected and submitted to QC | cleiton.lopes@sanofi.com | 68/45 QC | 2023-10-13 | 2023-10-13 | Rejected samples not to be submitted for testing and issues identified at receipt
2436 | SUB 2220 Trend review and result review from weekend samples | Yves.Antoine@sanofi.com | Room 1403 | 2023-10-14 | 2023-10-16 | It is expected LAL samples results to be normal
2435 | Negative control RODAC not submitted for aseptic requal (2300016035) performed on 07Oct23 | marissa.lentoni@sanofi.com | 68 NYA, 4700, QC Micro Lab | 2023-10-07 | 2023-10-14 | A rodac negative control would be submitted with the EM plates for the aseptic requal.
2434 | Cz Phenyl skid floating matter on the Bubble trap | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1503 | 2023-10-16 | 2023-10-16 | Gemba was call for Phath forward | Yes | Deviation
2433 | 8NYA Cerezyme - DGHDS010 - Low pressure alarm due to Nano load SUM running empty | stanley.durand@sanofi.com | 8NYA room 1503 | 2023-10-15 | 2023-10-15 | Per FBL-EWI-000029/MABC-QU-WIN-0046642 it's expected that chase buffer profibus cable is properly connected to load SUM
2432 | 8NYA -Issues with VI skid testing buffer records needed for today 16OCT23 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-10-16 | 2023-10-16 | RF1NID010 - needs to be 1k SUT and 200L barrel, RF1N8D006 - manual weigh out for liquid conversion GENDOC_203020, RF1NFD004 - asking for 1000L SUT 200 is under min mix timer.
2431 | 8NYA-Gemba 706210 Cancelled in BioLims | Tanner.Raiburn@sanofi.com | 8NYA-1350 | 2023-10-13 | 2023-10-13 | Find a path forward
2430 | 8NYA GEMBA - Stock creation volume issue. | jack.laplante@sanofi.com | 8NYA RM1350 | 2023-10-15 | 2023-10-15 | Looking for path forward and someone that can change the stock in athena
2429 | 8NYA GEMBA: Conductivity meter #2009970 incorrectly standardized | lolav.ahmed@sanofi.com | 8NYA room 1403 | 2023-10-13 | 2023-10-15 | Conductivity meter should be standardized using standard 12.88mS/cm prior to performing offline sample of buffer.
2428 | Bleed Setup step not executed | joseph.gonzalez@sanofi.com | 1403 | 2023-10-16 | 2023-10-15 | step should be executed during the completion of setup.
2427 | 8 NYA - DGHDS010 Cerezyme Nano Formulation leak observed from pump head during equil. | stanley.durand@sanofi.com | 8NYA Room 1503 | 2023-10-15 | 2023-10-15 | Pump head should not leak during equil.
2426 | Mouse in the JIT | michael.montiverdi@sanofi.com | JIT 8NYA | 2023-10-15 | 2023-10-15 | No pests in the JIT are expected
2425 | 8NYA- Cz- Media Wave bag for P/N 760744 found with sample bag already filled due to partially open pinch clamp | john.lytle@sanofi.com | Rm 1403, 8NYA | 2023-10-15 | 2023-10-15 | The sample bag is filled during point of use sampling, immediately prior to use in Equilibration operations
2424 | 8NYA GEMBA - Media tote scale not zeroed before | jack.laplante@sanofi.com | 8NYA CNC | 2023-10-14 | 2023-10-14 | Looking for path foward.
2423 | Steps not executed in 705665 record. | robert.morawski@sanofi.com | 8NYA CNC | 2023-10-13 | 2023-10-13 | Buffer record executed until completion, and buffer stock created. | Yes | No Deviation
2422 | Take Out issue with 760712 | robert.morawski@sanofi.com | 8NYA CNC | 2023-10-13 | 2023-10-10 | Tote volume should be correct in MES to match the physical volume left in the tote.
2421 | Pump head on Q skid was set up incorrectly. Flow direction was going the wrong way. | paul.devaney@sanofi.com | 8NYA/1503 | 2023-10-13 | 2023-10-13 | Pump head should have properly installed the correct way. pump should be pulling buffer correctly from CNC.
2420 | 8NYA-GEMBA-FZ Butyl EBR not executed after performing step. | alex.tello@sanofi.com | 8NYA 1323 | 2023-10-11 | 2023-10-12 | connection from cnc buffer to the skid to be completed on the batch record. | Yes | Deviation
2419 | Bradford assay sample bag failed to fill during TSH harvest sample collection. | james.letourneux@sanofi.com | 51 NYA Room 140 | 2023-10-12 | 2023-10-12 | When pinch clamps are opened on sample lines during harvest, the sample bags should fill with harvest fluid.
2418 | FB-1254 digital clock used at 51 in Room 127 Material Airlock is missing | brian.mcgee@sanofi.com | 51 Room 127 | 2023-03-31 | 2023-10-11 | It is expected for calibrated items to be properly removed from service if no longer used, and for items that cannot be located to be assessed for impact.
2417 | TSH Buffer - pH is out of range for DGFB1044 GZ1291-50 | Jeremy.Wooden@Sanofi.com | 45 NYA Room 190 | 2023-10-12 | 2023-10-12 | It is expected that the pH would be in range per QU-SOP-0059509 Attachment 6 and QU-OPE-0184556 step 4.20 |  | No Deviation
2416 | 76NYA- Biosurgery APU-HAFM Room 364-Mixer F-8090 was found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-364 | 2023-10-12 | 2023-10-12 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2415 | Low percent recovery for rhTSH SP2 DGF4B10A at 51NYA | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 131 | 2023-10-12 | 2023-10-12 | It is expected that the recovery range is between 67% to107% per QU-OPE-0487499 Step. 4.8.3
2414 | Meeting to discuss GZ1287-52 buffer leak in 200L barrel during SP1 DGF3B15B process. | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 138 | 2023-10-11 | 2023-10-11 | Per MABC-QU-OPE-0079501 buffer volume should be greater the 95 Liters
2413 | 8NYA-DOWNSTREAM-Fz-BUTYL -DGHBE051-LOAD ADJUST BAG HAS MANUFACTURER DEFFECT | marc.latta@sanofi.com | 1323 | 2023-10-11 | 2023-10-11 | fbl-ewi-000033
2412 | 8-NYA GEMBA - Tote leaking 0.1M Sodium Hydroxide | alex.tello@sanofi.com | 8NYA CNC | 2023-10-11 | 2023-10-11 | Path forward
2411 | Bioburden buffer sample 872612, lot DGHAS010 was received within 1 hour of expiry | Stefanie.Erickson@sanofi.com | QC Micro bioburden room | 2023-10-10 | 2023-10-11 | for the analyst to have enough time to prepare for testing, sanitizing reagents and biosafety cabinet, creating biolims batch and testing for bioburden sample.
2410 | Butyl regen stored devices not removed during cleaning recipe | Christian.Venegas@Sanofi.com | 8NYA Downstream suite 1 room 1323 | 2023-10-11 | 2023-11-17 | Devices should be removed prior the end of regen stored recipe
2409 | 8NYA - Charging 760726 to RF643D021 in SAP | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-10-11 | 2023-10-11 | Get Deviation number to reference in record for manual entry and manual calculations
2408 | Unable to locate SIP being performed on MCC skid prior to use | steven.erban2@sanofi.com | 8NYA Upstream | 2023-07-23 | 2023-10-11 | It is expected that the pre use MCC CIP & SIP would be performed prior to use | Yes | Deviation
2407 | Operator did not sign, when entering room 355 (HA core) | valeria.moreno@sanofi.com | 76-357 | 2023-10-10 | 2023-10-10 | The operator right after leaving the aseptic area signed the log sheet with a late entry and the supervisor cosigned as verifier.
2406 | Cerezyme - DGHBN018 - 8NYA - D0-D2 Sampling performed for incorrect stage in EBR; QC samples for Day 0 not taken | michael.knapp@sanofi.com | 8NYA 1403 | 2023-10-06 | 2023-10-10 | Sampling is performed in the correct loop. Day 0 QC samples (Bioburden and Endotoxin) are collected.
2405 | Buffer Storage Unit Issues in MES | robert.morawski@sanofi.com | 8NYA CNC | 2023-10-10 | 2023-10-10 | The Storage Unit should have remained in the tote.
2404 | Leak on Pre-Guard Filter During Nano Filtration | adam.keister@sanofi.com | 8 NYA, 1323 | 2023-10-10 | 2023-10-10 | N/A
2403 | 8NYA- Fz- Fabrazyme Production recipe does not control pH prior to Perfusion start. Investigation needed per MSAT request | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-10-10 | 2023-10-10 | pH control of the system at the appropriate time
2402 | 8NYA- Pressure sensor ruptured when ramping up during transfer 760741/RF1BHD130 | alejandro.martinez2@sanofi.com | 8NYA/8RM1350 | 2023-10-10 | 2023-10-10 | Discard of media batch
2401 | Thyrogen - SP1 Lot DGF4B14C - Low pH found during chromatogram review | craig.valery@sanofi.com | 51 NYA | 2023-10-10 | 2023-10-10 | Low pH found during chromatogram review. pH= ~4.5 however per attachment 3 of QU-SOP-0061107, range is 4.8-5.6
2400 | GEMBA - 8 NYA - Nano Filter leaking at clamped connection | adam.keister@sanofi.com | 8 NYA, 1323 | 2023-10-10 | 2023-10-10 | Path Forward
2399 | MMC Column Bottom Mobile Phase leaking | Juan.Bonilla@sanofi.com | 8NYA/ 1503 | 2023-10-09 | 2023-10-09 | Column should not be leaking
2398 | Approved version of paperwork was printed without the working copy signed | rajnee.bains@sanofi.com | QC Micro | 2023-10-08 | 2023-10-09 | Its expected that the working copy was signed and the effective form was filled out correctly.
2397 | 8NYA-Fz- ATF B wetting step in EBR not performed at time of execution | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-10-08 | 2023-10-09 | reql time documentation of GMP activities
2396 | 8NYA Fabrazyme - DGHSE015 - Gemb anot initially held for bag failure. | stanley.durand@sanofi.com | 8NYA room 1323 | 2023-09-15 | 2023-10-07 | Per FBL-EWI-000015 it is expected that procedure is followed in the event of an SUT failure.
2395 | Walk in cold room (F-115591) went out of temp range of 2-8C | amy.roy@sanofi.com | 68NYA/4700/QC Micro | 2023-10-08 | 2023-10-08 | Cold room Temperature must remain within 2-8C
2394 | During CZ Zinc WFI dilution the recipe in Delta V continues to go into Hi Hi alrarm. | paul.devaney@sanofi.com | 8NYA/1503 | 2023-10-09 | 2023-10-09 | During WFI dilution the cold 2-8 degree filtered WFI should be transfered into the SUM while maintaining temp under the 10 degree threshold.
2393 | DO high high alram on bioreactor X2200 | Brendan.Dickhaut@sanofi.com | 8NYA, Room 1333 | 2023-10-08 | 2023-10-08 | FBL-CST-000028 table 7.5 DO parameter is 30-120%. FBL-EWI-000011 SUB operation/trouble shooting allows for manual addition of O2 during low events. | Yes | Deviation
2392 | 8NYA - RF618D010 760739 Thaw record, incorrect Activity level entered, and corrected per CofA. Calculations need MES and MSAT verification | bailey.gleason@sanofi.com | 8NYA | 2023-10-08 | 2023-10-08 | Hold Gemba to see path forward.
2391 | 8NYA Fz x2200 CO2 sparge line needs to increase so that pH can do down and innoculation could occur. | paola.tello@sanofi.com | 8NYA RM1333 | 2023-10-08 | 2023-10-08 | N/A
2390 | 8NYA - Incorrect stock volume created for RF608D005 - 760726 SU 131843 (Follow up) | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-10-08 | 2023-10-08 | Hold Gemba to see path forward.
2389 | 8NYA - Incorrect stock volume created for RF608D005 - 760726 SU 131843 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-10-08 | 2023-10-08 | Hold Gemba to see path forward.
2388 | Gemba to discuss Missing verified by on BIRF on UF/DF DGF2B08. | Luis.Dominguez@sanofi.com | 51NYA/88 | 2023-10-04 | 2023-10-07 | Per QU-OPE-0074092 it is expected that the section C be verified during processing.
2387 | Walk in cold room went out of temp range of 2-8C, Incubator FB-2057 went out of temp range | amy.roy@sanofi.com | 68NYA/4700/QC Micro | 2023-10-08 | 2023-10-08 | Cold room Temperature must remain within 2-8C, incubator must remain within 20-25C
2386 | 8NYA-Fabryzyme Downstream UFDF- Aseptiquik Connection had from Zinc Eluate SUM outlet has broken seal | aidan.mula@sanofi.com | 8NYA RM1323 | 2023-10-07 | 2023-10-07 | Determination on path forward
2385 | Pump went into hold during pre elution buffer prime phase | paul.devaney@sanofi.com | 8NYA/1503 | 2023-10-08 | 2023-10-08 | Pump should have started during prime buffer phase
2384 | GEMBA for Cz Endotoxin sample found to be leaking upon sample receipt | Brandon.Steff@sanofi.com | 8RM1343 | 2023-10-07 | 2023-10-07 | Sample doe not leak prior to testing
2383 | GEMBA for F-15594 Warm Room OOS Briefly | jessica.kychler@sanofi.com | QC Microbiology, 68 NYA floor 4 room 4700 | 2023-10-07 | 2023-10-07 | It is expected that incubators stay within their specified range
2382 | 8NYA-Fabryzyme Downstream SP-Issues with 2 exceptions | aidan.mula@sanofi.com | 8NYA RM1323 | 2023-10-07 | 2023-10-07 | Procedures need to be followed at time of operation
2381 | 8NYA-Fabryzyme Downstream Zinc- CNC breakdown step not performed at time of operation | aidan.mula@sanofi.com | 8NYA RM1323 | 2023-10-06 | 2023-10-06 | Steps should be performed at time of operation.
2380 | Unable to verify manually loaded Delta V Recipes during QA batch record review | steven.erban2@sanofi.com | 8NYA Downstream | 2023-08-22 | 2023-10-06 | It is expected that if a Delta V recipe is manually loaded, it can be found and verified the correct recipe is loaded. |  | No Deviation
2379 | 8NYA-8NYA APU-Room 1403-CC8-FB-4221-048 and FB-4221-140 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 8-1403 | 2023-10-05 | 2023-10-06 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results |  | Deviation
2378 | Sartorius Biosealer overdue for 12 month PM is unaccounted for. | joel.henchar@sanofi.com | 8NYA Upstream | 2023-10-01 | 2023-10-01 | Equipment outside of PM/calibration should not be used in manufacturing.
2377 | Action notice not initiated for sampling of 51-DI-1 | marissa.lentoni@sanofi.com | 51 NYA, room 129A (parts washer mechanical space) | 2023-09-23 | 2023-10-04 | Facilities would be contacted to initiate an action notice prior to performing sampling.
2376 | Unable to verify manually loaded Delta V Recipe during QA batch record review | steven.erban2@sanofi.com | 8NYA Downstream | 2023-08-23 | 2023-10-05 | It is expected that if a Delta V recipe is manually loaded, it can be found and verified the correct recipe is loaded. | Yes | No Deviation
2375 | MCC X2831 failed visual inspection after CIP | Jeffrey.boldizar@sanofi.com | 8 NYA 1403 | 2023-10-05 | 2023-10-05 | Clean Surface | Yes | Deviation
2374 | Gemba - 8NYA - Pompe Nano -DGHDA004X- Material takeout failed and DOM CX for SU 132630 did not populate | stanley.durand@sanofi.com | 8NYA Room 1703 | 2023-10-04 | 2023-10-04 | Material needs to be taken out and Cx values properly populated. | Yes | Deviation
2373 | Lower than expected A280 results following diafiltration | alyson.Belleville@sanofi.com | 51 NYA | 2023-10-04 | 2023-10-04 | Historically the results are greater than 2.0
2372 | HA Lot - CBYFRM037; incorrect program selected for filter integrity test | charlie.houle@sanofi.com | 76 NYA; Room 364 / 3105 (HA Fermentation) | 2022-11-22 | 2023-09-29 | Per FBS-MBR-000058 (QU-OPE-0309296) v15.0; step 7.11.22 , a forward flow test integrity test using air as a test gas will be performed per FBS-SOP-000594 (QU-SOP-0046034).
2371 | Multiple EM discrepancies in HA Media Fill | craig.stevenson@sanofi.com | 76NYA, Room 355 | 2023-09-19 | 2023-10-04 | Please see attached
2370 | Multiple instances of CTMS not using OneLIMS when BCP is not in place | craig.stevenson@sanofi.com | 45NYA | 2023-09-19 | 2023-10-04 | It is expected that all dynamic EM sampling should be performed using OneLIMS when BCP isnt in place.
2369 | Out of stock of mophead needed for internal wipe of bioreactor | david.loik@sanofi.com | 51 NYA | 2023-10-04 | 2023-10-04 | Mopheads would be available for use.
2368 | Incorrect biolims build for endotoxin raw material 871321 Lot 22GN01533 | johny.lisitu@sanofi.com | LAL Bench QC Micro | 2023-10-03 | 2023-10-03 | For bioLIMs to allow a chrom batch be created.
2367 | Incorrect biolims build for endotoxin raw material 871321 Lot 22GN01533 | johny.lisitu@sanofi.com | LAL Bench QC Micro | 2023-10-03 | 2023-10-03 | For bioLIMs to allow a chrom batch be created.
2366 | Failed SUM container - DS Ops | paul.devaney@sanofi.com | Room 1324 - Downstream Phase 1 | 2023-10-03 | 2023-10-03 | Container expected to maintain seam integrity
2365 | 8NYA - RF1N5D002 Cannot charge acetic acid to the BOM in SAP, Needed for shipment. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-10-03 | 2023-10-03 | Hold Gemba to see path forward.
2364 | 8NYA-Pompe-Downstream ANX-Part not used as required on BOM | aidan.mula@sanofi.com | 8NYA RM1503 | 2023-10-03 | 2023-10-03 | Determination at Gemba, likely removal of item from BOM | Yes | Deviation
2363 | 8 NYA-Downstream-Fz-Zinc-DGHZE016 Record Issue | christina.burkhart@sanofi.com | 8NYA, RM 1324 | 2023-10-03 | 2023-10-03 | Record should be started day of performance
2362 | Plates INC-08407 Batch: 202303758-1 not incubated in OneLims | sarah.adams@sanofi.com | 68 NYA, QCM | 2023-09-27 | 2023-10-03 | Incubation time and date recorded in OneLIMS
2361 | 8 NYA-Downstream-Fz-Zinc-DGHZE016 Record Issue | christina.burkhart@sanofi.com | 8NYA, RM 1324 | 2023-10-03 | 2023-10-03 | Record should be started day of performance | Yes | No Deviation
2360 | X2250 CO2 probe AIT2250029 not reading | Brandon.steff@sanofi.com | 8NYA1403 | 2023-10-03 | 2023-10-03 | Probe should be giving values on the transmitter
2359 | X2240 | Brandon.steff@sanofi.com | 8NYA1403 | 2023-10-02 | 2023-10-02 | line should not be pinhed allowing gases to flow freely
2358 | Non-contemporaneous data entered for Cz SDS-PAGE in OneLIMS, Lot: AGHDS005, 36M | aimee.park@sanofi.com | 68NYA, QC Chemistry 2nd floor | 2023-08-16 | 2023-10-02 | The analyst would have entered all assay data contemporaneously
2357 | TAC/Mold CFU results not recorded on negative control and settle plate paperwork during counting on 25SEP23 | amanda.lavalley@sanofi.com | 68 NYA, QC Micro Lab | 2023-09-25 | 2023-10-01 | Results are expected to be recorded on the Settle Plate paperwork at the time counting is completed.
2356 | Gemba for 8NYA EM plates removed from incubator before the minimum incubation period (INC-08298). And 8NYA gas batch without a positive control result (INC-08294). | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2023-09-28 | 2023-09-30 | Per QU-FOR-0118296 samples must be incubated in the 30-35C for a minimum of 48 hours. And per QU-FOR-0118296 anaerobic conditions have to be documented.
2355 | MCC Skid X-2831 failed visual inspection during CIP swab sample | Yves.Antoine@sanofi.com | Room 1403 | 2023-10-02 | 2023-10-02 | Bubble trap Sensor to pass Visual inspection
2354 | 45NYA-Thyrogen APU-Room 169A-AutoClave #6, 4 GMP Critical instruments overdue for calibration | robert.bauch@sanofi.com | 45-169A | 2023-10-01 | 2023-10-02 | QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5.5 unit should have been calibrated before its due date |  | Deviation
2353 | No reportable results for TOC for site 76-WFI-2 on 30Sep23 | kathryn.lentoni@sanofi.com | 68NYA 4700, QCM | 2023-10-02 | 2023-10-02 | For there to be uncomprised TOC results
2352 | 8NYA- FZ SP. SUT item in EWI not in EBR BOM | alex.tello@sanofi.com | 8NYA RM1323 | 2023-10-02 | 2023-10-02 | path foward
2351 | Failed shaker incubator in Cell culture. Currently using contingency plan for tCCR at status 4 | jesse.welch@sanofi.com | 51-140 | 2023-10-01 | 2023-10-01 | operate as designed |  | Deviation
2350 | RF1NAD002 Pompe PPQ1 record needs review finished up | dana.hopkins@sanofi.com | 8 NYA room 1403 | 2023-10-01 | 2023-10-01 | MFG review should be completed
2349 | TSH Buffer - SLED dates not populating in Athena for DGFB1103 GZ1384-59 | Jeremy.Wooden@Sanofi.com | 45 NYA Room 190 | 2023-09-28 | 2023-09-28 | It is expected that the SLED date in Athena would be prepopulated when performing Goods Receipt transactions. |  | No Deviation
2348 | Plates INC-08359 Batch: 202303703-1 not incubated in OneLims | sarah.adams@sanofi.com | 68 NYA, QCM | 2023-09-25 | 2023-09-30 | Incubation time and date recorded
2347 | DGHBE049, 762729 Resampling of Butyl Eluate Endotoxin in BioLims due to Original Sample Taken at the Incorrect Point | wilfredo.lopez@sanofi.com | 8NYA, Room 1323 Downstream. | 2023-10-01 | 2023-10-01 | It is expected that all samples are taken and labeled correctly per the sampling plan.
2346 | TSH Formulated Bulk Thawing in Cold Room for Shipment DGFPB009 Found in Ambient Hallway | alyson.belleville@sanofi.com | 51 NYA | Unknown | 2023-10-01 | It is expected that the product would remain in the cold room that it was last documented to be stored in. No undocumented movements should take place and product should remain within temp constraints or only be ambient for allotted TOR.
2345 | 8NYA Fabrazyme - DGHDT016- Shipping bag may not meet inspection passing criteria. | stanley.durand@sanofi.com | 8NYA rm 1313 | 2023-09-30 | 2023-09-30 | All welds on shipping bags should meet visual inspection criteria.
2344 | 8NYA-Incorrect SUM scanned into MES record for 705688/RF1AUD012 | alejandro.martinez2@sanofi.com | 8NYA/8RM1350 | 2023-09-30 | 2023-09-30 | MES to give pathforward if record steps can be re-executed and will launch recipe to correct SUM (PBM1515)
2343 | 51NYA, verification missing for placing flask back into warm room | adam.rowe@sanofi.com | 51NYA, room 153 | 2023-08-09 | 2023-09-30 | Step should have been verified
2342 | 8NYA - Pompe UF-B - DGHUB003 - SUT used prior to expiry could not be transacted against process order. | stanley.durand@sanofi.com | 8NYA Rm 1503 | 2023-08-24 | 2023-09-28 | It's expected for SUT to be transacted against process orders prior to their expiry.
2341 | 8NYA- Incorrect SUT used for transfer per BOM for 805010 batches | alejandro.martinez2@sanofi.com | 8NYA/8RM1350 | 2023-09-29 | 2023-09-29 | MES to update BOM
2340 | 8GEMBA: Wave reactor X-2141 held for gas flow LO LO alarm | Lolav.Ahmed@sanofi.com | 8 NYA 1403 | 2023-09-29 | 2023-09-29 | MABC-QU-OPE-0519399 table 7.5, overlay flowrate (non-KPP) PAR 0.10 - 0.50LPM no impact within PAR. Flowrate dropped down to 0.0LPM for approximately one minute.
2339 | Ceiling tile observed to open in room 51 51NYA room 130B | jacob.kallon@sanofi.com | 51NYA 130B | 2023-09-29 | 2023-09-29 | It is expected that ceiling tiles remain integral within the clean rooms at 51 NYA
2338 | 8NYA GEMBA: Sterile connection on A2B Straight Line in ATF assembly kit is misaligned | joseph.gonzalez@sanofi.com | 8NYA room 1403 | 2023-09-29 | 2023-09-29 | paper edge is non existent on seal connection
2337 | GEMBA - 8 NYA - BioLims Not available for DGHDR003X | adam.keister@sanofi.com | 8 NYA, 1703 | 2023-09-29 | 2023-09-29 | Path Forward, either labels generated or we use MTR's to follow sample plan
2336 | D1's level sensor used in D2 | david.loik@sanofi.com | 51 NYA | 2023-09-27 | 2023-09-29 | D2 level sensor would continue to work properly
2335 | GEMBA - 8 NYA - Short PN 843138 Shipping Bags for Pompe Refiltration | adam.keister@sanofi.com | 8 NYA, 1703 | 2023-09-29 | 2023-09-29 | Permission to use alternative PN to move process forward
2334 | 8NYA DGHAW123 Bioburden sample Expired | cleiton.lopes@sanofi.com | 8NYA | 2023-09-28 | 2023-09-28 | Bioburden samples must be delivered to QC Micro with enough time for testing.
2333 | 8NYA FZ DGHBE048 Bioburden sample leak | cleiton.lopes@sanofi.com | QC Micro | 2023-09-28 | 2023-09-28 | compromised samples not to be submitted, picked up, or received into the lab
2332 | CV 2023-45CV-0001 Received but not Returned to QA | paul.mclaughlinjr@sanofi.com | 45NYA | 2023-09-28 | 2023-09-28 | All CV paperwork signed out would be returned to QA
2331 | Pompe UFDF A/B Skids Module Failure DeltaV Alarms were not enabled during the Cz to Pompe Suite Changeover | seyda.dumrul@sanofi.com | 8NYA Suite 2 Downstream | 2023-08-22 | 2023-09-28 | MABC-QU-SOP-0016987-Product Changeover at 8 NYA, moving equipment into the Suite
2330 | GEMBA for Mised BSC Clean Documentation | jessica.kychler@sanofi.com | QC Microbiology, 68 NYA floor 4 room 4700 | 2023-09-08 | 2023-09-27 | It is expected analysts will document pre and post cleans of BSC per procedure
2329 | 148 warm room temp alarm | nicole.lufkin@sanofi.com | 51 NYA warm room 148 | 2023-09-27 | 2023-09-28 | range per MABC-QU-SOP-0003344-Portable Media Storage and Usage Log (v5.0), is 35-39 degrees
2328 | 50L wave DGHCG014 EOR samples taken, after which it was determined to take 50L wave out another day, retains need to be resampled | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-09-27 | 2023-09-27 | 50L EOR samples taken on final day of wave stage, retains do not need to be duplicated | Yes | No Deviation
2327 | Fz Butyl recipe went into hold during cleaning prime phase | edgar.vale@sanofi.com | 8NYA/ 1324 | 2023-09-27 | 2023-09-27 | It's expected that would not go into hold
2326 | Cz SUB X2240 media DO low following pH control | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-09-27 | 2023-09-27 | DO reading about 50% prior to innoculation
2325 | pompe campaign set up using dummy pendotech plugs | dana.hopkins@sanofi.com | 8NYA room 1403 | 2023-09-27 | 2023-09-27 | Should not have gotten bleed pressure alarm before bleed hooked uo
2324 | Incorrect Recipe parameter for DO stabilization X2240 | Brandon.steff@sanofi.com | 8NYA1403 | 2023-09-27 | 2023-09-27 | 15 min stabilization time
2323 | Incorrect Cleaning Recipe Launched on UFDF B | zachary.rickard@sanofi.com | 8NYA/1503 | 2023-09-26 | 2023-09-26 | The correct recipe will launch from MES. | Yes | Deviation
2322 | ELISA samples tested using wrong HRP | kristin.gothier@sanofi.com | QCM 4700 | 2023-09-23 | 2023-09-26 | Current lot/dilution of HRP would be used to test
2321 | HA Lot - DBYFRM041; Sample port valve not operational | charlie.houle@sanofi.com | 76 NYA; Room 364 (HA Fermentation) | 2023-09-26 | 2023-09-26 | The expected result would have flow from the sample port when opened so a sample can be collected and tests can be run.
2320 | Cerezyme - DGHBN017 - 8NYA - SUB bag failed visual inspection | michael.knapp@sanofi.com | 8NYA 1403 | 2023-09-25 | 2023-09-25 | The seal around the bag is of sufficient width to ensure integrity.
2319 | Leak Observed Pre and Post Filter During Form Nano Equilibration | wilfredo.lopez@sanofi.com | 8RM1503 | 2023-09-26 | 2023-09-26 | It is expected that no leaks are observed during this oparation. | Yes | Deviation
2318 | 8NYA Pompe - UFDF-B product pump stop due to speed error | stanley.durand@sanofi.com | 8NYA room 1503 | 2023-09-25 | 2023-09-25 | Pump should not stop due to speed error
2317 | DeltaV is reading inaccurate pH reading during production. | kenny.leland@sanofi.com | 8 NYA/1503 | 2023-09-25 | 2023-09-25 | pH should reach range.
2316 | Inproper gowning observed at 51NYA Room 139 | craig.stevenson@sanofi.com | 51 NYA Room 139 | 2023-09-25 | 2023-09-25 | All trained personnel to follow gowning procedure
2315 | 8NYA- EM Expired INC-08350 Collection Batch :202303698-1 | john.lytle@sanofi.com | Rm 1343 8NYA, QC Sample management | 2023-09-25 | 2023-09-25 | It is expected that EM plates be incubated in the allowable time | Yes | Deviation
2314 | Missing Entries for Negative Control Volume | krystal.masseri@sanofi.com | 68NYA QCM 4700 | 2023-09-29 | 2023-10-03 | The negative control volume should be documented
2313 | 51NYA, 12L wave placed on wrong gas mixture | adam.rowe@sanofi.com | 51NYA, room 152 | 2023-09-23 | 2023-09-25 | CO2 should have been left at 10%
2312 | Path forward needed for UFB membrane unistall | zachary.rickard@sanofi.com | 8NYA/1503 | 2023-09-25 | 2023-09-25 | Any step that occurs on a piece of manufacturing equipment should have a procedure for it.
2311 | 76 NYA PUR DBYHAS010. "SWECO drying outlet moisture reading is out of range." | munish.kumar2@sanofi.com | 76 NYA/Room#3103/Purification | 2023-09-25 | 2023-09-25 | Per QU-OPE-0275728 step # 7.7.8 the dew point for SHMW should be between -18C to -26C (recommended -22C). The reading of the moisture probe is -80C
2310 | Cerezyme - DGHBN017 - 8NYA - SUB bag tear located following passing integrity test | michael.knapp@sanofi.com | 8NYA 1403 | 2023-09-24 | 2023-09-25 | It is expected that the SUB bag is integral, and that any defects or tears are located through the inspection and integrity test activities.
2309 | Bioburden plates placed in incorrect bag resulting in incorrect incubation for 2 batches | Stefanie.Erickson@sanofi.com | QC Micro bioburden room | 2023-09-21 | 2023-09-25 | For the sample plates to be placed in the correct bag that matched the biolims batch to be put in the correct incubator
2308 | 8NYA-8NYA APU-Room 1503-Weight Indicating Controller FB-4229-027 found out of tolerance during an unscheduled calibration for suite change over. | robert.bauch@sanofi.com | 8-1503 | 2023-09-25 | 2023-09-25 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2307 | 8NYA- Fz- RF500D007 scale up to 1L operations started prior to Daily Sample activities | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-09-25 | 2023-09-25 | Scale up operations proceed after required sampling performed per MABC-QU-OPE-0028959-Fz-ICB Sampling Plan for 8 NYA
2306 | D2 bioreactor Micro carriers exceeded 24 hour expiry | adam.hildebrand@sanofi.com | 51-140 | 2023-09-23 | 2023-09-23 | It is expected that the pressure hold would be integral and the Microcarriers would not exceed expiry
2305 | incorrect batch ID on EM sample labels | ryan.devanna@sanofi.com | Innoc Prep at 8 NYA | 2023-09-24 | 2023-09-24 | correcting the batch IDs on sample labels with the accurate batch ID
2304 | pH out of range on 12L wave | nicole.lufkin@sanofi.com | 51 NYA 152 | 2023-09-24 | 2023-09-24 | pH should be in range
2303 | Butyl Eluate Hold tote Bag failed inspection | Christian.Venegas@Sanofi.com | 8NYA Downstream suite 1 | 2023-09-24 | 2023-09-24 | Phath forward to used new bag and return damaged bag to the vendor
2302 | SUB bag found with hold | Ralston.najarro@sanofi.com | 8NYA1403 | 2023-09-24 | 2023-09-24 | No holes in bag
2301 | Diafiltration buffer runout | Christian.Venegas@Sanofi.com | 8NYA Downstream suite 1 | 2023-09-23 | 2023-09-23 | Need Phath forward to continue process
2300 | Speaker in gowning room 120 fell out of ceiling | lindsey.hymer@sanofi.com | 51NYA/120 | 2023-09-23 | 2023-09-23 | rooms should remain integral with no potential breaches
2299 | Conductivity spec not being achieved during blank run of ANX skid | Brendan.Dickhaut@sanofi.com | 8NYA/Room1503 | 2023-09-22 | 2023-09-22 | MABC-QU-OPE-0306912, FBL-CST-000031
2298 | 8NYA Fabrazyme UF/DF Downstream EWI doesn't match with the Batch Record. | Christian.Venegas@Sanofi.com | 8NYA Downstream suite 1 room 1323 | 2023-09-22 | 2023-09-22 | to alinged EWI and the Record | Yes | Deviation
2297 | 8NYA-GEMBA, Abnormal observation during swabbing of UFDF- A | alex.tello@sanofi.com | 8NYA RM1503 | 2023-09-22 | 2023-09-22 | Cleaning skid during swabbing, QU-FOR-0090360
2296 | missed calibrations for autoclave and parts washer at 45NYA | lindsey.hymer@sanofi.com | 45NYA/169 | 2023-09-22 | 2023-09-22 | Calibrations should be able to be performed on time
2295 | Fz SP- Missing pH adjust buffer | nathaniel.barnes@sanofi.com | 8 NYA, 1323 | 2023-09-22 | 2023-09-22 | Determine a path forward for processing. Product is currently in contact time with polysorbate buffer (max 24hrs) and the column has already been equilibrated and needs to remain chilled for processing.
2294 | 76 NYA PUR DBYHAS010 “Failed to deliver samples on time for Bioburden and Endotoxin to QC ". | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2023-09-20 | 2023-09-21 | QU-SOP-0050801 step# 6.2.9; QU-SOP-0037624 step# 6.35.5; QU-OPE-0202445 step #7.8.6-7.8.6.4. Per the abovementioned SOPs and respective steps, the samples must be stored at 2-8 °C and must be delivered immediately to QC.
2293 | Expired EM Plates Incubation Batch 2023003635-1, FB GZ5157-00 DGF3B07 | alyson.belleville@sanofi.com | 51-88 | 2023-09-21 | 2023-09-21 | It is expected that the plates would be delivered to the laydown area at 51 and then picked up and delivered to QC with enough time to incubate the plates.
2292 | Warm room 148 QCA, low temp | adam.rowe@sanofi.com | 51NYA, room 148 | 2023-09-19 | 2023-09-20 | Temp should remain within warm room control parameters 35C-39C
2291 | 6-10 Pompe Sample Fridge - Vendor completed work may effect validated state | Nicholas.grote@sanofi.com | 1410 | 2023-09-15 | 2023-09-20 | NA
2290 | 45NYA-Support APU-QCC45-Room 1130-Balance F-17180 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 45-1130 | 2023-09-20 | 2023-09-21 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2289 | Bioburden began leaking after sample receipt. | sarah.adams@sanofi.com | 68 NYA | 2023-09-20 | 2023-09-20 | Sample to be sealed to delivery to QCM
2288 | Pompe Copper Blank Run Wash Buffer Volume too low | zachary.rickard@sanofi.com | 8NYA 1212 | 2023-09-21 | 2023-09-21 | Buffer will be batched to the correct size needed to run a process.
2287 | 8NYA-Cz- RF639D012 SUB X2220 pH Offset incomplete and incorrect 1 point pH standardization performed | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-09-19 | 2023-09-20 | pH offset and 1 point Std performed per FBL-EWI-000011
2286 | Pompe UF B Membrane Install CNC Record Incomplete | zachary.rickard@sanofi.com | 8NYA/1212 | 2023-09-09 | 2023-09-20 | Buffer takeouts will be performed in real time and all steps in batch record will be signed for when they occur. | Yes | Deviation
2285 | 8NYA-Pompe-DGHAE006X- Two samples received by QCM with same number. | stanley.durand@sanofi.com | 8 NYA 1503 | 2023-09-20 | 2023-09-20 | Two samples should not contain same sample ID
2284 | Pompe UF A Buffer Supply Pump Failure | zachary.rickard@sanofi.com | 8NYA/1503 | 2023-09-20 | 2023-09-20 | Pump should not stop working while process is running.
2283 | Leaking Bottle of Serum | stephen.roberts2@sanofi.com | 51NYA Dock | 2023-09-19 | 2023-09-19 | It is expected that this material arrives in uncompromised containers.
2282 | Multiple "Change Approval and Implementation Form" not returned | Charlie.Houle@sanofi.com | 76 NYA Biosurgery | 2020-10-21 | 2023-08-11 | It is expected the form would be completed and returned to QA. | No
2281 | 8NYA- Pompe- DGHUA008X- No back-up Periodate buffer available for Pompe UF-A Oxidation. | stanley.durand@sanofi.com | 8NYA rm 1503 | 2023-09-20 | 2023-09-20 | Per FBL-EWI-000047 an unexpired Back-up Periodate buffer is required to execute oxidation.
2280 | Bioburden received without any labeling. | sarah.adams@sanofi.com | 68 NYA | 2023-09-19 | 2023-09-19 | Samples to be labeled with LIMS or test request.
2279 | GEMBA for Cz SUB X2240 DO calibrated incorrectly (FBL-EWI-000010 has incorrect instruction) | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-09-19 | 2023-09-19 | DO is calibrated to 100% air prior to charging media, pCO2 and pH are calibrated just before inoculation, not DO
2278 | Small leak at the "T" connection on UF/DF A skid, 8NYA phase 2 downstream room 1503 | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1503 | 2023-09-19 | 2023-09-19 | need Phath forward to continue with UF\DF operations | Yes | Deviation
2277 | GEMBA: Onelims Spreadplate batch build calculates total bioburden incorrectly | judith.keller@sanofi.com | 68 NYA 4700 QCM | 2023-09-19 | 2023-09-19 | MABC-QU-SOP-0019490 | Yes | No Deviation
2276 | 68NYA-Support APU-room 4700-QCM68-TOC Analyzer F-19932 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 68-4700 | 2023-09-14 | 2023-09-19 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2275 | TSH Buffer - bottom parameter for low range drop check for scale F-7143 is set to high in MES for 1409 batch. | Jeremy.Wooden@Sanofi.com | 45 NYA Room 190 | 2023-09-19 | 2023-09-19 | MES would allow us to use a lower value weight for drop check without creating an exception. |  | Deviation
2274 | 8 NYA - Pompe ANX Load concentration failure | alex.tello@sanofi.com | 8NYA RM1503 | 2023-09-17 | 2023-09-18 | path forward
2273 | cracked bovine serum jug used in batch of media 609615 | lindsey.hymer@sanofi.com | 51NYA/125 | 2023-09-16 | 2023-09-18 | per QU-OPE-0192519 step 4.3.16 calculated volume of serum is 120.0L | Yes | Deviation
2272 | CNC BUFFER MANIFOLD HAS FAILD VISUAL INSPECTION | marc.latta@sanofi.com | 1503 | 2023-09-18 | 2023-09-18 | the manifold should not have any damage;
2271 | Live Centipedes was found in 76 NYA Room 3105 (HA Purification)". DBYHAS010 | munish.kumar2@sanofi.com | 76 NYA/Room#3105/Purification | 2023-09-18 | 2023-09-18 | Refer to QU-SOP-0005611 steps 6.2.2 and 6.3.1. Per the SOP, the classified area should be free of pest activities.
2270 | 8NYA-Fabryzyme Downstream SP-Issues with exceptions in record | aidan.mula@sanofi.com | 8NYA RM1323 | 2023-09-18 | 1050 | Determination of impact at Gemba
2269 | Low Viability on Daily Vicell Sampling for Fz DGHCE005 | Yves.Antoine@sanofi.com | 8NYA, Room 1333 | 2023-09-18 | 2023-09-18 | Expect the Viability to stay above the 80% | Yes | Deviation
2268 | 8NYA APU-Room 1503-PF8-no As founds taken on FB-4282-001 Mixer due to the unit being non-functional | robert.bauch@sanofi.com | 8-1503 | 2023-09-15 | 2023-09-18 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2267 | Dry Shipper Received with Alarmed Dataloggers and Missing Vial Receipt Logbook | Jonathan.Lucas@sanofi.com | 45 NYA / Security Desk | 2023-09-15 | 2023-09-16 | MABC-QU-SOP-0070997
2266 | Pompe Copper Eluate below target % Recovery | zachary.rickard@sanofi.com | 8NYA/1503 | 2023-09-17 | 2023-09-17 | %Recovery will be greater than or equal to the target stated in the batch record.
2265 | 8NYA- Fz- Eluate SUM temperature indicates potential issues with the TCU | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-09-17 | 2023-09-17 | Per MABC-QU-OPE-0524531-Control Strategy for Fabrazyme Integrated Continuous Biomanufacturing (Fz-ICB) Process, "The Capto Adhere eluate is continuously collected into jacketed single-use mixers where it is kept at 2–10 °C"
2264 | 8NYA, FZ. Butyl parking spot temperature control not working | joseph.gonzalez@sanofi.com | 8 NYA, Room 1333 | 2023-09-16 | 2023-09-16 | TCU control should operate as intended per reicpe
2263 | GEMBA for Leaking Bioburden Sample MEDIA_760741_BIOBURDEN Sample # 853455 | jessica.kychler@sanofi.com | QC Microbiology, 68 NYA floor 4 room 4700 | 2023-09-16 | 2023-09-16 | It is expected that the bioburden sample will be intact to maintain sterility for testing
2262 | 8NYA GEMBA: Product vials identified to be in WMS01 location rather than Cryostorage rack in MES | Lolav.ahmed@sanofi.com | 8NYA 1280 | 2023-09-16 | 2023-09-16 | N/A
2261 | Pompe Copper Pooling SUM leaking from pH probe. | zachary.rickard@sanofi.com | 8NYA 1503 | 2023-09-16 | 2023-09-16 | SUM bags will not leak.
2260 | FZ SP CNC Buffer Setup Error | robert.morawski@sanofi.com | 8NYA CNC | 2023-09-15 | 2023-09-15 | Buffer transfer line connected to the correct port. | Yes | Deviation
2259 | Work/cleaning not signed for on AN 2023-F-062 | david.loik@sanofi.com | 51 NYA | 2023-09-14 | 2023-09-15 | Work performed would be signed for as it is completed |  | Deviation
2258 | Leaking Bioburden Sample Delivered to QC Sample Management | sarah.adams@sanofi.com | 68 NYA | 2023-09-15 | 2023-09-15 | Sample to be sealed closed correctly | Yes | No Deviation
2257 | Late entry for Manifold Tube Connection | zachary.rickard@sanofi.com | 8NYA/CNC | 2023-08-12 | 2023-09-14 | Steps in batch record will be signed for in real time when the operation is performed. | Yes | Deviation
2256 | 8 NYA- GEMBA FZ, 3K tote bag manufactured incorrectly | christain.galvez-tello@sanofi.com | 8RN 1202 | 2023-09-15 | 2023-09-15 | Inspection results to be made correctly | Yes | Deviation
2255 | LAL vial for endotoxin testing cracked | lauren.ruffing@sanofi.com | 68NYA 4700 QC Micro | 2023-09-15 | 2023-09-15 | vial is not cracked
2254 | 160 L Bioprocess Container received defective | jeffrey.mei@sanofi.com | 51 NYA 140 | 2023-09-13 | 2023-09-13 | The bioprocess container is expected to come from the supplier with all connections secured.
2253 | FBL-EWI-000011 approval without final authorizer checked | janet.langon@sanofi.com | 10CAL | 2023-09-14 | 2023-09-14 | Document would have Final authorizer checkbox checked | Yes | No Deviation
2252 | Anaerobic gas chamber latch not secured leading to non-anaerobic conditions. | anastacia.hachem@sanofi.com | 68 NYA QC Microbiology | 2023-09-08 | 2023-09-14 | It is expected that the anaerobic bioburden chamber is latched securely and anaerobic conditions are met per QU-SOP-0023473
2251 | Step missed in MFG quarterly walkthrough at 51NYA | nicholas.shoenfelt@sanofi.com | 51-143 | 2022-05-07 | 2023-05-25 | It is expected that all steps in the PM would have been performed.
2250 | No product-no lots-68NYA: Incubator maintenance log not completed within 6-8 day window as stated in step 6.4 of QU-SOP-0049227. | larissa.fairweather@sanofi.com | 68NYA/4735/QCM | 2023-09-08 | 2023-09-11 | Per QU-SOP-0049227 step 6.4 we are to verify incubator water level each week, 6-8 days from the last recording.
2249 | 68NYA-Support APU-Room 2700-QCC68-Balance FB-2466 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 68-2700 | 2023-09-13 | 2023-09-14 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2248 | 51NYA-Support APT-Room 146-F-10964-012 and F-10964-015 Resistivity Meters on Clean Steam Generator #1 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 51-146 | 2023-09-13 | 2023-09-14 | QU-SOP-0069490 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2247 | 8NYA-DOWNSTREAM-DGHBE045 BUTYL SKID SETUP RECIPE PROMPT NEED TO BE BACKUP | marc.latta@sanofi.com | 1323 | 2023-09-14 | 2023-09-14 | CNC set up should be finished prior to answering the prompt "sut install complete" | Yes | Deviation
2246 | 8NYA - GEMBA FZ . Sample bag was not removed during sampling | alex.tello@sanofi.com | 1313 | 2023-09-13 | 2023-09-13 | Remove the sample bag | Yes | Deviation
2245 | Trainer comment not documented during formulation training | jacob.kallon@sanofi.com | 51 NYA/room 88 | 2023-08-24 | 2023-09-13 | Per QU-SOP-0074684 step 6.11.5, if individual performing task is completing records as part of a training process, those steps in record must be countersigned by a trainer who has been qualified on that specific task.
2244 | 8NYA - Post Autoclave activities were performed prior to autoclave. | bailey.gleason@sanofi.com | 8NYA RM1373 | 2023-09-13 | 2023-09-13 | Hold Gemba to see path forward.
2243 | 8NYA -Operator did not take two sets of rinse samples for FZ UFDF CV | Pwynes51@gmail.com | 8NYA 1323 | 2023-09-13 | 2023-09-13 | its expected that CV procedure will be followed by operators conducting the sampling
2242 | Post filtration pressure did not maintain between 5.0-10.0 DBYFRM039 768307 | dianne.moustafa@sanofi.com | 76 NYA Room 364 Fermentation | 2023-09-12 | 2023-09-12 | Pressure maintains
2241 | Filter elements were no replaced and missing step verification in HA Ferm - DBYFRM034 | Juan.Ramos-Torres@sanofi.com | 76 NYA Room 364 | 2023-08-01 | 2023-09-13 | filters should have replaced when re-doing SIP
2240 | Cz ICB CGHBN004 First Harvest Day Sampling Not Performed | mark.lindblad@sanofi.com | 8 NYA Phase 2 Upstream | 2022-09-07 | 2023-09-12 | First Harvest Day samples are required per QU-OPE-0044689, Table 8. FHD Endotoxin and Bioburden are IPCs per QU-OPE-0519399.
2239 | Gemba 8NYA-Pompe Downstream Nanofiltration- Material take out failed and DOM CX failed to populate | aidan.mula@sanofi.com | 8NYA RM1503 | 2023-09-12 | 2023-09-12 | Material needs to be taken out and Cx values properly populated. |  | Deviation
2238 | Breach of ceiling due to water leak in room 45-169 | nicholas.shoenfelt@sanofi.com | 45-169 | 2023-09-12 | 2023-09-12 | The integrity of the room is expected to be maintained per FBL-SOP-000367.
2237 | Cracked bovine serum jug leaving us short on materials for media batch 609615 | lindsey.hymer@sanofi.com | 51NYA/109 | 2023-09-13 | 2023-09-13 | materials should have no cracks or leaks prior to use
2236 | 8NYA- Nanofilter wetting automation recipe changed, not captured in EWI | alejandro.martinez2@sanofi.com | 8RN1323 | 2023-09-12 | 2023-09-12 | Pathforward on nanofilter
2235 | Fabrazyme - DGHSD004 -8NYA - Wave 2132 Utilized after 2130 Identified as Out of Service | michael.knapp@sanofi.com | 8NYA 1403 | 2023-07-16 | 2023-09-12 | Correct wave is identified and setup per EBR. Any wave out of service by metrology should be out of service in MES. | Yes | Deviation
2234 | Fabrazyme - DGHSD004 -8NYA - No Verification Steps for Atrium Manual Entries | michael.knapp@sanofi.com | 8NYA 1343 | 2023-07-06 | 2023-09-12 | Consistency with documentation of atrium information, including verification.
2233 | Fabrazyme - DGHSD004 -8NYA - No Wave Tray Selection EWI instruction for Activity | michael.knapp@sanofi.com | 8NYA 1333 | 2023-07-12 | 2023-09-12 | All EWI references point to Effective EWI instructions. | Yes | Deviation
2232 | 8NYA Fabrazyme - UFDF Load Filter Over Pressurizing Due To Clogging | stanley.durand@sanofi.com | 8 NYA Room 1323 | 2023-09-09 | 2023-09-10 | Pump is expected to stay on and not over pressurize | Yes | Deviation
2231 | Column Cleaning CV, late entry comment needs improved justification | zachary.rickard@sanofi.com | 8NYA/1303 | 2023-07-30 | 2023-09-11 | Steps should be signed for in real time and if a late entry is to be made the justification should not be questionable. | Yes | Deviation
2230 | 8NYA-Fz-Takeout of emergency Bleed replacement SUT not performed | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-08-02 | 2023-09-11 | All SUT used for processing is expected to be taken out of inventory using authorized transactions | Yes | Deviation
2229 | 8NYA GEMBA: Discussion on EWI clarifications for SUB probe autoclave and install probe steps | Lolav.Ahmed@sanofi.com | 8NYA | 2023-09-11 | 2023-09-11 | N/A
2228 | Insufficient flow rate for CIP | Jeffrey.boldizar@sanofi.com | 8 NYA 1333 | 2023-09-11 | 2023-09-11 | For flow to maintain above 330L and properly clean the skid
2227 | 8NYA Pompe - DGHDR001 - Missing procedure for Pompe DS refiltration in EWI | stanley.durand@sanofi.com | 8NYA rm1703 | 2023-09-10 | 2023-09-11 | It's expected to have procedure in FBL-EWI-000029 to perform refiltration of Pompe DS.
2226 | ANMS Alarms not generating exceptions into batch record. | timothy.ryley@sanofi.com | 08 NYA | 2023-08-31 | 2023-09-11 | The ANMS alarms would generate exceptions in the batch records and those exceptions would need to be explained and acknowledged.
2225 | Personnel incomplete on CORE GMP Trainings and are currently executing GMP work | stephanie.aceto@sanofi.com | 8 NYA | 2023-06-30 | 2023-07-11 | Individuals complete all training on time in the training period
2224 | Performed by missing from 1 page in shipper log | adam.rowe@sanofi.com | 74NYA, room 166 | 2023-04-25 | 2023-09-11 | Step should be signed for when performed
2223 | Unable to Acquire CNC valves for CNC setup. | Juan.Bonilla@sanofi.com | 8NYA | 2023-09-11 | 2023-09-11 | Open CNC valves to perform CNC setup.
2222 | 45NYA-Room 1139-QCC- No As Found Data collected before repair of GMP Critical EMS sensor on FB-1464-300 | robert.bauch@sanofi.com | 45-1139 | 2023-09-08 | 2023-09-11 | BQ-GOP-000046 Calibration Program Version 6 Section 6 as found should have been collected before the repair
2221 | GEMBA for Fz MCC X2510 Post Use CIP Low Buffer Flow | Brandon.Steff@sanofi.com | 8RM1333 8NYA | 2023-09-10 | 2023-09-10 | MCC CIP completes caustic flush before buffer supply runs out, typically consumes ~400L of NaOH
2220 | At 8 NYA mechanical room 1170, On WFI still #1, Sight glass is broken. Unit is down until get fixed. | jesus.acosta2@sanofi.com | 08-1170 mechanical room | 2023-09-10 | 2023-09-10 | No sight glass broken. No leaks present. System should be running. | Yes | Deviation
2219 | Gemba-8NYA-Downstream Fz UFDF-Permeate PCV went into hold | aidan.mula@sanofi.com | 8NYA RM1323 | 2023-09-10 | 2023-09-10 | Determination at Gemba
2218 | 8NYA GEMBA- FZ UFDF going into hold due to pressure sensor | alex.tello@sanofi.com | 8NYA RM1323 | 2023-09-10 | 2023-09-10 | pump not to stop
2217 | 8NYA - Pre use Nano wetting SUT leaked from the Pump insert. | bailey.gleason@sanofi.com | 8NYA | 2023-09-10 | 2023-09-10 | Hold Gemba to see path forward. Get a new pre use SUT and restart recipe.
2216 | 8NYA - GEMBA FZ UFDF recipe did not open valve to install SUT | alex.tello@sanofi.com | 8NYA RM1323 | 2023-09-09 | 2023-09-09 | open up valve to be able to install SUT
2215 | Facility personnel observed entering grade C gowning area without properly adhering to the gowning procedure. | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 116, 88,89 | 2023-09-08 | 2023-09-08 | It is expected that the all personnel entering the manufacturing areas following MABC-QU-SOP-0006952
2214 | Incorrect column volume used during TSK lot DGF2B08B processing | jacob.kallon@sanofi.com | 51 NYA room 132 | 2023-09-06 | 2023-09-08 | It is expected that the correct volume be used to calculate TSK loading capacity and to determine target fraction volumes
2213 | Sartorius mixer X-4201012 asset # 2012466 needs to be shipped to vendor for repair | douglas.dubois@sanofi.com | 8NYA | 2023-06-14 | 2023-06-14 | none
2212 | 76 NYA PUR DBYHAS009 “Missing data on Nitrogen cool down Chart recorder ". | munish.kumar2@sanofi.com | 76 NYA/Room#3103/Purification | 2023-09-07 | 2023-09-08 | QU-SOP-0048453. Step# 6.9.2, 6.9.5, 6.10.3. Due to missing information and unreadable data there is not a way to ensure that the pressure of the system was maintained.
2211 | 51 NYA rhTSH Purification-Operator observed entering grade C area without following proper gowning procedure | jacob.kallon@sanofi.com | 51 NYA/room 88 | 2023-09-08 | 2023-09-08 | It is expected that gowning procedure ( QU-SOP-0006952) be adhered to when entering controlled manufacturing areas at all times. |  | Deviation
2210 | 8NYA - GEMBA Pompe - Incorrect recipe launched for CIP | alex.tello@sanofi.com | 8NYA RM1503 | 2023-09-08 | 2023-09-08 | Run recipe with CV
2209 | 8NYA DOWNSTREAM POMPE NANOFILTRATION EWI DISCREPENCY | marc.latta@sanofi.com | 1503 | 2023-09-08 | 2023-09-08 | 3 CHASE BUFFERS NEED TO BE CONNECTED PER FBL-EWI-000029 | Yes | Deviation
2208 | GEMBA Problem Solving- Small qty of shipper GZ7313-01 used @74NYA are sealed w/2" tape on box bottom | alan.rousseau@sanofi.com | Room 1017 74 NYA | 2023-09-07 | 2023-09-07 | Expected result is that the bottom of shipper would be sealed with 3" tape
2207 | Osmometer EWI weekly standardization ask for 50, 850 and 2000 standard, however we don't have a 2000 standard in place. | maria.palmar@sanofi.com | 8 New York Ave | 2023-09-08 | 2023-09-08 | FBL-EWI-000024
2206 | Weight for liquid weigh and dispense activities was off. | jermaine.kellam@sanofi.com | 8RM 1350 | 2023-09-07 | 2023-09-07 | Material transactions should have processed.
2205 | GEMBA - 8 NYA - PN 808710 (Pompe Nano Pre-Guard) Spinning Clamps at Pump | adam.keister@sanofi.com | 8 NYA, 1503 | 2023-09-07 | 2023-09-07 | Path Forward with Functioning SUT
2204 | GEMBA - 8 NYA - Unable to draw A280 sample from Retentate Vessel | adam.keister@sanofi.com | 8 NYA, 1503 | 2023-09-07 | 2023-09-07 | Path forward
2203 | LAL sample auto-released without valid result | kristin.gothier@sanofi.com | QCM lab 68NYA 4700 | 2023-09-06 | 2023-09-07 | Lot will not auto-release an invalid result
2202 | Need a verifiable sample time for stability Bioburden testing | Kathy.tourville@sanofi.com | 68 NYA QC Micro lab | 2023-09-06 | 2023-09-06 | Expectation is that it should be clear for the timeframe to test the Bioburden within
2201 | Action Notice 2023-I-40 was released based on non-viable environmental Monitoring Results | jennifer.marion@sanofi.com | 8 NYA, Room 1403 | 2023-09-07 | 2023-09-07 | QU-SOP-0022403 Action Notice Procedure Attachment #3 and Section 6.7.2 Release Requirements and QU-FOR-0118296 Step Environmental Monitoring Requirements for Room 1403
2200 | 51 NYA GEMBA: last count of zero filter not zero for dynamic EM performed on 22AUG23 | megan.botti-marino@sanofi.com | 51 NYA | 2023-08-22 | 2023-09-07 | Per step 6.4.1, table 2, of MABC-QU-SOP-0030227, the final count of the zero filter for a grade A space must be zero.
2199 | Fabrazyme-DGHSE013-8NYA- Manual entry of conductivity solution into the EBR missing second verifier. | stanley.durand@sanofi.com | 8NYA rm 1323 | 2023-08-11 | 2023-09-07 | When the conductivity solution lot number is manually entered a second verifier is required per MABC-QU-SOP-0019707.
2198 | Cerezyme - DGHBN013 - 8NYA - DO Stabilization Timer Not Configured Properly for Phase 3 Recipe | michael.knapp@sanofi.com | 8NYA 1403 | 2023-09-06 | 2023-09-06 | Correct stabilization timer is attached to recipe.
2197 | 8NYA-Fz-X2543 SUM 1 Profibus Loose connection is placing MCC production on hold | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-09-06 | 2023-09-06 | SUM connection/ profibus connection should remain in a Connected state during active MCC operations following FBL-EWI-000032 eluate collection
2196 | 'As Founds' not performed prior to removal of D2 agitator | david.loik@sanofi.com | 51 NYA | 2023-09-06 | 2023-09-06 | As founds would be performed prior to the removal of the agitator. |  | Deviation
2195 | Expired Bioburden Sample | sarah.adams@sanofi.com | 68 NYA | 2023-09-05 | 2023-09-06 | Sample to be submitted within 24 expiry
2194 | X2250 has signed for QA contained Equipment execute GMP activity (SUB set up and Probe install) | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2023-09-05 | 2023-09-06 | Equipment should have release for use prior to start GMP activity
2193 | 8NYA-Pompe CNC space room 1212 Buffer spill | stanley.durand@sanofi.com | 8NYA room 1212 | 2023-09-06 | 2023-09-06 | There should not be leaks observed from SUM bags.
2192 | 80NYA-Thyrogen APU-room 150, -80 Freezer unplugged waiting for repair, unable to collect as found data. | robert.bauch@sanofi.com | 80-150 | 2023-09-05 | 2023-09-06 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2191 | MES launched incorrect recipe version for Cz SUB X2250 | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-09-05 | 2023-09-05 | MES launches UP-SUB-PROD-P3 variant of SUB production recipe
2190 | QC micro GEMBA: Bench cleaning performed with expired IPA on 22AUG23 | megan.botti-marino@sanofi.com | 68 NYA | 2023-08-22 | 2023-09-05 | It is expected that cleaning agents be within expiry when used.
2189 | 8NYA Gemba - MCC Eluate SUM temperature out of range | Alex.tello@sanofi.com | Rm 1333 | 2023-09-05 | 2023-09-05 | to be within 16 - 24 temperature
2188 | GEMBA to discuss cold room 138 at 51NYA high temperature reading | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 138 | 2023-09-05 | 2023-09-05 | It is expected that the room temperature is between 6-10C per FBL-SOP-000790.
2187 | Out of sequence step performed during Butyl set up recipe | Christian.Venegas@Sanofi.com | 8NYA Downstream suite Room 1323 | 2023-09-05 | 2023-09-05 | Phath forward to be able to continues with CNC and skid set up for Butyl
2186 | GEMBA for Cz X2250 pCO2 probes used as dummy probes | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-09-05 | 2023-09-05 | Dummy probes consist of a set of probes that is not actually used for processing, bellows remain integral for probes used for processing.
2185 | 8 NYA - 820323 DGHFA001 Biolims not taken | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-09-05 | 2023-09-05 | Hold Gemba to see path forward.
2184 | No reportable result for site 210, HA fill routine static EM (no operations performed), 09AUG23 | megan.botti-marino@sanofi.com | 76 NYA | 2023-08-09 | 2023-09-05 | It is expected that all non-viable sites 210-214 and 217-222 be sampled on a weekly basis per step 6.2.1 of QU-SOP-0074634.
2183 | PQR missing Bioburden and LAL results. CZ CGHBN00406 | Devin.Parks@Sanofi.com | 1403 | 2023-09-05 | 2023-09-05 | The four once per eluate samples should have been taken.
2182 | Buffer GZ1402-57 3x FIT Failure (DGFB2040) | cindy.guerra@sanofi.com | 45NYA Rm. 190 | 2023-09-04 | 2023-09-04 | FIT would have passing results per MABC-QU-SOP-0072849
2181 | 45NYA Thyrogen APU-Room 169A-4 items on 45 parts washer overdue for calibration | robert.bauch@sanofi.com | 45-169A | 2023-09-01 | 2023-09-05 | GF-QU-SOP-0069490 Calibration Program Version 6 Section 5.7.5 unit should have been calibrated before its due date
2180 | Pest found in 76NYA DSMB Room 388 on 03Sep2023 | craig.stevenson@sanofi.com | 76NYA, Room 388 | 2023-09-03 | 2023-09-03 | No pests in the classified rooms
2179 | 8NYA APU-Room 1403-28 items on MCC Chromatography Skid X-2831001 overdue for calibration | robert.bauch@sanofi.com | 8-1403 | 2023-09-01 | 2023-09-05 | BQ-GOP-000046 Calibration Program Version 6 Section 5.7.5 unit should have been calibrated before its due date
2178 | MCC Skid 2560 Collection Monitor Failure | sean.gormley@sanofi.com | 8RM1403 | 2023-09-04 | 2023-09-04 | Per EWI-000032 it is expected that lines from MCC to collecting totes will be setup appropriately.
2177 | 804995 - Delta V recipe will not proceed | leo.gillespie@sanofi.com | 1350 - Solution prep | 2023-09-04 | 2023-09-04 | batch WFI draw does not exceed tolerances
2176 | One (1) Late Preventative Maintenance (PM) Work Order Not Completed by Extended Due Date of 31Aug2023 for Bioreactor Pump Asset # 2009192 | kevin.southwick@sanofi.com | 8NYA, Room 1333 | 2023-09-01 | 2023-09-04 | Complete all PM tasks prior to PM closure (otherwise approved by QA) or a deviation is required per site specific SOP | No | No Deviation
2175 | 8NYA Gemba - Incorrect auto recovery value. | alex.tello@sanofi.com | 8NYA RM1503 | 2023-09-04 | 2023-09-04 | Value to be in range
2174 | 8 NYA - 820323 DGHFA002 Sample plan has 2 post filter samples. No Biolims are available and previous record had no samples taken. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-09-04 | 2023-09-04 | Hold Gemba to see path forward.
2173 | Leak on collection sum near P/H probe on Copper flow thought. | kelvin.glean@sanofi.com | 8 NYA | 2023-09-04 | 2023-09-04 | No leaks on bag.
2172 | FZ bioburden sample 836737 lot DGHBE041 found leaking during verification | Stefanie.Erickson@sanofi.com | QC Micro bioburden room | 2023-09-03 | 2023-09-03 | for the sample to remain sterile
2171 | Fabrazyme – DGHUP013– 8NYA – UFDF skid pump leaked during Equil recipe. | stanley.durand@sanofi.com | 8NYA room 1323 | 2023-09-02 | 2023-09-02 | No leaks should be observed during UFDF operations. | Yes | Deviation
2170 | 8NYA- Fz- SUB X-2210 pressure reading below expected values. Potential faulty Pendotech sensor on Production Bioreactor. | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-09-02 | 2023-09-02 | The SUB is expected to maintain positive pressure of </= 0.3psi during process operations. Increased foam levels require additional process monitoring utilizing a functional pressure indicator | Yes | No Deviation
2169 | Fabrazyme – DGHZE013– 8NYA – Expired WFI used during Zinc setup | stanley.durand@sanofi.com | 8NYA room 1210 | 2023-09-02 | 2023-09-02 | It's expected to use unexpired WFi during operations. | Yes | Deviation
2168 | Pompe – DGHUA004– 8NYA – UFDF-A stock creation CX value error | stanley.durand@sanofi.com | 8NYA rm 1503 | 2023-09-01 | 2023-09-01 | It is expected that the CX value would propagate onto the storage unit | No | Deviation
2167 | 8NYA-Fz-SUB X2200 Annual PM could not be completed by FAC by date due | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-09-01 | 2023-09-01 | It is expected that PMs be executed by due dates listed, and that equipment be available for PM to be executed during scheduled dates listed |  | Deviation
2166 | Sample valve on E3 unable to enable | david.loik@sanofi.com | 51 NYA | 2023-09-01 | 2023-09-01 | Enabling of the sample valve would be allowed.
2165 | Pompe ANX Buffer Take Out Issue | robert.morawski@sanofi.com | 8NYA CNC | 2023-09-01 | 2023-09-01 | Tote starting volumes would have been updated with new tote information. |  | No Deviation
2164 | 8NYA GEMBA: Unfiltered 1K batched WFI bag leaking for Post CIP on MCC Skid X2831 | Lolav.Ahmed@sanofi.com | 8 NYA room 1403 | 2023-09-01 | 2023-09-01 | n/a |  | No Deviation
2163 | 8NYA Pompe _ UFDF-A UV sensor not in MES | stanley.durand@sanofi.com | 8NYA Room 1503 | 2023-08-18 | 2023-08-31 | It's expected for UV sensors to be in MES in order to be allocated by record | Yes | Deviation
2162 | Pompe bioburden EM samples for spread plates were removed from incubation early. | Alyssa.Barrett@sanofi.com | 68 NYA/Room 4700 | 2023-08-05 | 2023-08-31 | Per QU-SOP-0036327 step 6.6.5, settling and RODAC plates are to be incubated at 20-25 C for a minimum of 5 days and a maximum of 7 days.
2161 | Pompe bioburden sample 807175 and negative control were removed from incubation early. | Alyssa.Barrett@sanofi.com | 68 NYA/Room 4700 | 2023-08-13 | 2023-08-31 | Per QU-SOP-0036604 step 6.2.4, plates are to be incubated at 30-35 C for 48-72 hours.
2160 | Pompe UF/DF A recipe went into hold due to pump failure | edgar.vale@sanofi.com | 8NYA, room 1503 | 2023-09-01 | 2023-09-01 | Troubling shooting the pump error and access any impact to the batch | Yes | Deviation
2159 | Hold on bioreactor X-2200 due to pressure related event with vent filters. | Brendan.Dickhaut@sanofi.com | 8NYA, Room 1333 | 2023-08-31 | 2023-08-31 | Vent filters do not become wetted out leading to pressure related events | Yes | Deviation
2158 | 8NYA DOWNSTREAM ANX DGHAE002 PH ADJUST ISSUE | marc.latta@sanofi.com | 8NYA/1503 | 2023-08-31 | 2023-08-31 | EXPECTED THAT RECORD , EWI AND DELTA V PROMPT WILL ALIGN
2157 | Multiple routine QC EM/CU sites canceled in LIMS by MFG for all MA BioCampus | craig.stevenson@sanofi.com | 45, 51 and 8NYA | 2023-08-28 | 2023-08-31 | It is expected that MFG does not cancel QC's routine EM/CU sites.
2156 | Unable to locate original document | jacob.kallon@sanofi.com | 51 NYA/room 104 | 2023-08-31 | 2023-08-31 | It is expected that the original column qualification chart be retained with the column packing batch record
2155 | 68 NYA 4700 BSC Vacuums not working 30AUG23 PM | judith.keller@sanofi.com | 68 NYA 4700 QCM | 2023-08-30 | 2023-08-30 | Hoods operate as expected
2154 | UF-A Equil recipe hold due to WFI lockout | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2023-08-31 | 2023-08-31 | Determine if any further impact at Gemba | Yes | No Deviation
2153 | ANX Load- protein is below MES target range | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2023-08-30 | 2023-08-30 | MSAT input needed to determine a path foward and what impact this will have on product and associated recipes | Yes | Deviation
2152 | GEMBA for production media did not switch CNC valve for SUB X2220 BV-002 | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-08-30 | 2023-08-30 | BV-002 in the CNC would open when switching from growth media to production media to allow media to be pumped into SUB. | Yes | Deviation
2151 | Leak at clamp of sterile filter during media filtration | lindsey.hymer@sanofi.com | 51NYA/125 | 2023-08-30 | 2023-08-30 | there should be no leaks in filtration line
2150 | Verified step not signed for confirming load pattern on parts washer at 51NYA | nicholas.shoenfelt@sanofi.com | 51-129 | 2023-08-29 | 2023-08-30 | It is expected that the setup of the load on the parts washer rack be verified and signed for by a second operator prior to starting the run.
2149 | 51 NYA room 116 QCA not commented on | jacob.kallon@sanofi.com | 51 NYA/Room 116 | 2023-07-24 | 2023-08-29 | Per MABC-QU-SOP-0070733 alarms must have impact statement added by the area owner.
2148 | Nanofilter leaked at filter outlet triclamp connection post autoclave | caleb.coppola@sanofi.com | 8RM1373 | 2023-08-29 | 2023-08-29 | Nanofilter would not have any leaks
2147 | Pompe ANX buffer line leak | zachary.rickard@sanofi.com | 8NYA/CNC | 2023-08-29 | 2023-08-29 | Buffers will not leak past clamps and flood buffer manifold.
2146 | 8NYA-Cz-X2220 requires additional sampling/ testing per MSAT request for investigational purposes | john.lytle@sanofi.com | Rm 1343 8NYA | 2023-08-29 | 2023-08-29 | n/a- refer to MSAT trend analysis
2145 | 8NYA ASM NRSR 22SR8053 appears to be sampled in triplicate | judith.keller@sanofi.com | 68NYA/4700; 8NYA upstream | 2022-11-24 | 2023-08-27 | It is expected that NRSRs are executed per the NRSR request and delivered in a timely manner.
2144 | GEMBA - 8 NYA - Leak fin Nano Filter Discovered during Skid Setup | adam.keister@sanofi.com | 8 NYA, 1323 | 2023-08-29 | 2023-08-29 | Path Forward
2143 | Missed Signature Entries | Jonathan.Lucas@sanofi.com | 51 NYA / rm 152 | 2023-08-28 | 2023-08-29 | Steps should have performed by/verified by signatures
2142 | Column failed visual inspection | joseph.gonzalez@sanofi.com | 1403 | 2023-08-29 | 2023-08-29 | Column housing intact from supplier
2141 | HA - DBYHAS005 - 76NYA - Excursion obtained during HA Fill without followup | craig.stevenson@sanofi.com | 76NYA, Room 355 | 2023-07-25 | 2023-08-29 | It is expected that the sampling analysts follows the SOP to reduce the NV levels if an excursion has occured.
2140 | Non-viable site 211 not re-sampled after alert level obtained, 76 NYA, Routine HA fill monitoring (no Fill operations occruing) | megan.botti-marino@sanofi.com | 76 NYA | 2023-08-02 | 2023-08-29 | Per MABC-QU-SOP-0074750 (Use of the Climet CI-1054) step 6.5.8, when an alert level non-viable excursion is obtained, the site is to be re-sampled per FBL-SOP-000192. It is expected that a re-sample would have been taken for this site.
2139 | Fz SP Low pH timer exceeded 210 minutes | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-08-28 | 2023-08-28 | Low pH timer would be ended before reaching the maximum 210 minutes.
2138 | No reportable viable or rodac results for NBO routine EM performed on 22Aug2023 | mlentoni@gmail.com | 68 NYA 4700 QCM lab | 2023-08-22 | 2023-08-28 | Incubation workflow would be initiated for all EM samples.
2137 | 8NYA GEMBA: MCC X2831 holding post EOR acknowledged prompt for col 1 flush | lolav.ahmed@sanofi.com | 8 NYA room 1403 | 2023-08-28 | 2023-08-28 | N/A
2136 | Pompe UF A equipment visual inspection failure | zachary.rickard@sanofi.com | 8NYA/1503 | 2023-08-28 | 2023-08-28 | Cleaning will properly remove any particles from the skid.
2135 | 45NYA - Room 172 - Incorrect gowning by MFG | craig.stevenson@sanofi.com | 45NYA Room 172 | 2023-08-28 | 2023-08-28 | • Don a reusable or disposable frock. • Don an additional pair of gloves. • Don a pair of low-top shoe covers, putting on one shoe cover stepping over the line then repeat for second shoe cover. • Refer to FBL-FRM-000961 “Secondary Gowning for Sanitization Suites”
2134 | 8NYA-Fz-Eluate SUM unable to reach temperature setpoints. Investigation required | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-08-27 | 2023-08-27 | Per MABC-QU-OPE-0524531, The Capto Adhere eluate is continuously collected into jacketed single-use mixers, where it is kept at 2-10deg C
2133 | 8NYA-Pompe- Last ROSS freeze bag underfilled. | stanley.durand@sanofi.com | 8NYA room 1212 | 2023-08-27 | 2023-08-27 | It's expected per FBL-EWI-000043 "ROSS Fill Freeze Bags" for all freeze bags to reach a target weight of 7.3-7.7 kg.
2132 | GEMBA for leaking harvest bioburden sample for Pompe eluate 25 | Brandon.Steff@sanofi.com | 8NYA 8RM1410 | 2023-08-26 | 2023-08-26 | Caps are tight on sample bag, bioburden sample is not opened prior to testing at QC.
2131 | 2 Expired Pompe Bioburden samples submitted to QC | kathryn.lentoni@sanofi.com | 8NYA | 2023-08-27 | 2023-08-27 | Expect the samples to be delivered before the 24 hour expiry
2130 | 8 NYA GEMBA - Pump tubing leak from G-connection | alex.tello@sanofi.com | 8NYA RM1330 | 2023-08-26 | 2023-08-26 | No leak
2129 | 8NYA - Post Nano filter Wetting recipe calls for 225 flush volume. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-08-26 | 2023-08-26 | Hold Gemba to see path forward.
2128 | 8NYA - Post Nano filter Wetting recipe calls for 225 flush volume. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-08-26 | 2023-08-26 | Hold Gemba to see path forward.
2127 | Gemba-8NYA-Pompe PPQ 1 UF-A CV Downstream- Visual inspection failed after 3rd CIP | aidan.mula@sanofi.com | 8NYA RM1503 | 2023-08-25 | 2023-08-25 | Determination on path forward
2126 | Gemba-8NYA RM1503- Pompe PPQ 1 Downstream Nano-Incorrect info in FBL-EWI-000029 | aidan.mula@sanofi.com | 8NYA RM1503 | 2023-08-25 | 2023-08-25 | EWI should be correct
2125 | EWI instructions for Nano Filter wetting do not match EBR tasks | christopher.arnold@sanofi.com | 8RM1323 | 2023-08-24 | 2023-08-24 | That the EWI and EBR will be in agreement over parts required to perform the operation.
2124 | Leak on Pompe Nano filter | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1503 | 2023-08-25 | 2023-08-25 | get path foward to continue operations
2123 | Pompe Nanofiltration- Incorrect info in FBL-EWI-000029 | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2023-08-25 | 2023-08-25 | EWI needs to be corrected to correctly align instructions to our process parameters
2122 | Leak during UFDF B load on discharge of P-3441004 load pump. | kenny.leland@sanofi.com | 8NYA 1503 | 2023-08-24 | 2023-08-24 | No leaks should be observed.
2121 | A Gemba is needed for a High High alarm for UV reading during UFDF B concentration. | kenny.leland@sanofi.com | 8 NYA 1503 | 2023-08-24 | 2023-08-24 | less than .2 OD
2120 | HighHigh alarm for UV reading duringUFDF B concentration. | kenny.leland@sanofi.com | 8NYA 1503 | 2023-08-24 | 2023-08-24 | less than .2 OD
2119 | Leak during UFDF B load on discharge of P-3441004 load pump. | kenny.leland@sanofi.com | 8NYA 1503 | 2023-08-24 | 2023-08-24 | No leaks should be observed.
2118 | No Documentation of IPA Contact Time for Enhanced Bench Cleaning | joseph.napoli@sanofi.com | 68NYA QCM | 2023-08-22 | 2023-08-23 | for the bench to be cleaned with IPA after a minimum of 5 minute contact time of sporklenz, and for the time to be recorded
2117 | Incorrect incubation time recorded for 2-hour alkylation for FBL-MTD-000023 | aimee.park@sanofi.com | 45 New York Avenue, QC Chemistry | 2023-07-19 | 2023-08-24 | A 2 hour incubation for alkylation would have been performed, not 3 hours.
2116 | Fabrazyme - DGHCE004 - 8NYA - Low Viability on Daily Vicell Sampling | michael.knapp@sanofi.com | 8NYA 1333 | 2023-08-19 | 2023-08-24 | SUB remains >=80% viability
2115 | Incorrect volume of water added to sample dilution for FBL-MTD-000012 TSH SDS-PAGE | aimee.park@sanofi.com | 45 New York Avenue | 2023-07-24 | 2023-08-24 | The correct volume of water would have been added for a total volume of 80uL.
2114 | 8NYA-Gemba Request- Solution Transfer- Delta V recipe RF1N0D002#201 finished transferring but recipe stuck on "Ready for Pressure Control-percent xfer" | alejandro.martinez2@sanofi.com | 8NYA Solution Prep, 1350 | 2023-08-24 | 2023-08-24 | Abort recipe as transfer is complete
2113 | 8 NYA GEMBA Request- 692321 missing in BOM for 820323 | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-08-23 | 2023-08-23 | All required SUT would be included in the BOM
2112 | 8NYA UFDF-A skid failed visual inspection. | kenny.leland@sanofi.com | 8NYA 1503 | 2023-08-23 | 2023-08-23 | System should be visually clean during visual inspection.
2111 | No- Product - No Lots - 68NYA - No reportable results for 4 EM Viable Air samples at 68NYA and negative control | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2023-08-06 | 2023-08-23 | Per MABC-QU-SOP-0073284, it is expected that Viable Air samples be taken for weekly EM samples in the BSC at 68NYA.
2110 | Fz Butyl retain samples mis-labeled | christopher.arnold@sanofi.com | 8RM1323 | 2023-08-21 | 2023-08-22 | That the correct labels will go on the correct samples.
2109 | Pompe Copper yield is below target | zachary.rickard@sanofi.com | 8NYA/1503 | 2023-08-23 | 2023-08-23 | Yield would be greater than or equal to the low limit of 69% | Yes | Deviation
2108 | 2 gowning qualifications generated failing trainer results | craig.stevenson@sanofi.com | 76NYA, Room 355 | 2023-08-16 | 2023-08-23 | It is expected that the trainer generate 0CFU/plate and generate passing results
2107 | Pompe - RF1NAD002 - 8NYA - No MCC Buffer Prime Recipe or Instructions for Ascending Column Flow | michael.knapp@sanofi.com | 8NYA 1403 | 2023-08-07 | 2023-08-23 | Column priming should occur on the inlet of the column. For Pompe columns, this would be the bottom, rather than the top.
2106 | Pompe - RF1NAD002 - 8NYA - SUB X-2240 Perfusion destination set to Waste following restart | michael.knapp@sanofi.com | 8NYA 1403 | 2023-08-23 | 2023-08-23 | Pathway remains open to the correct location: Harvest SUM or HBOD filter
2105 | Last count on zero filter was not zero before critical operations | megan.botti-marino@sanofi.com | 45NYA | 2023-08-08 | 2023-08-23 | Per RDSOP-003130, step 5.1.7, "The final reading of the zero count filter check must be within acceptance criteria prior to commencing critical operations."
2104 | 8NYA Pompe - Bioburden samples for UFDF-A operations expired! | stanley.durand@sanofi.com | 8NYA rm 1210 | 2023-08-19 | 2023-08-22 | Bioburden samples have a 24 hour expiry per FBL-EWI-000009 and should be put on test within 24 hours from time samples were taken.
2103 | FBL-EWI-000009 has links for sampling plan as approved for Fz and Cz | daniel.bati@sanofi.com | 8NYA | 2023-08-22 | 2023-08-22 | To open an effective sampling plan
2102 | 805010 - Delta V recipe stuck in loop | leo.gillespie@sanofi.com | 1350 - Solution prep | 2023-08-23 | 2023-08-23 | Batch should proceed after acknowleding prompt
2101 | Controls and Standards Inventory Binder Cannot Be Located | nancy.watkins@sanofi.com | QC Chemistry, 45NYA | NA | 2023-08-09 | A physical count of each assay standard, control or critical material currently in inventory will be performed, at minimum, quarterly and documented on FBL-FRM-000424. All records of FBL-FRM-000424 will be kept with QC per FBL-SOP-000835..
2100 | Gemba-8NYA-Fz-Downstream Nano- Nanofilter was autoclaved and scanned into record but operation can no longer be supported | aidan.mula@sanofi.com | 8NYA | 2023-08-22 | 2023-08-22 | Determination at Gemba
2099 | Gemba-8NYA-Pompe PPQ 1 ANX- Issues closing previous record causing impact to setup record on next run | aidan.mula@sanofi.com | 8NYA RM1503 | 2023-08-22 | 2023-08-22 | Determination at Gemba
2098 | Gemba-8NYA-Pompe PPQ 1 UF-A Downstream- Production record launched incorrect CIP recipe | aidan.mula@sanofi.com | 8NYA RM1503 | 2023-08-22 | 2023-08-22 | Record should run correct recipe
2097 | HA Engineering Lot - DBYENG018; Samples dropped off after hours expired before getting on test | Charlie.Houle@sanofi.com | 76 NYA, Room 364 (HA Fermentation); | 2023-08-21 | 2023-08-22 | It is expected that EM plates are dropped off/submitted within 24 hours so they can be properly incubated by governing procedure.
2096 | 45 NYA MAHU 1 unexpectedly shutdown at approximately 0926 21AUG23 | Alex.Biskaduros@Sanofi.com | 45 NYA rooms 174, 177 186C and additional rooms parts washer interstitial space. | 2023-08-21 | 2023-08-22 | MAHU 1 should not have shutdown unexpectedly
2095 | 76NYA Room 107-QCBP-Weight Set F-21403 50G weight found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-107 | 2023-08-04 | 2023-08-22 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2094 | 75665 - SUM bag leakage identified during initial WFI | Evans.Kamau@sanofi.com | 8NYA Solution prep RM1350 | 2023-08-21 | 2023-08-21 | SUM is not expected to leak
2093 | GMP WFI loop leak on room 166 at 45 nya MSAT | jesus.acosta2@sanofi.com | 45 NYA, ROOM 166 MSAT | 2023-08-21 | 2023-08-21 | no leak expected
2092 | No- Product - No Lots - 68NYA - QC Sampler was able to review their own collection batch in OneLIMs (Batch: 202302746-1) | craig.stevenson@sanofi.com | 51NYA, Multiple Rooms | 2023-07-27 | 2023-08-21 | It is expected that the sampling Analyst is unable to review their own data in OneLIMS.
2091 | Gemba-8NYA RM1503-Pompe PPQ 1-Downstream- 805003 buffer did not have second set of POU samples in LIMS, submitted under MTR | aidan.mula@sanofi.com | 8NYA RM1503 | 2023-08-20 | 2023-08-21 | Determination at Gemba
2090 | 8NYA Fz UF/DF DGHUP012 Load pump alarm | marc.latta@sanofi.com | 1323 | 2023-08-21 | 2023-08-21 | Recipe should run to completion without any alarm.
2089 | Thyrogen/Buffer Prep - Expired Carboy was used when filling DGFBP020 (P/N 1713) | Jeremy.Wooden@Sanofi.com | 80 NYA room 102 | 2023-08-18 | 2023-08-21 | It is expected that all Carboys filled would be within their sterile hold.
2088 | 8NYA Cz UF/DF DGHUP012 Load pump alarm | marc.latta@sanofi.com | 1323 | 2023-08-21 | 2023-08-21 | Load should proceed without load pump alarm
2087 | 8NYA Pompe - Leak observed during Copper pooling. | stanley.durand@sanofi.com | 8NYA room 1503 | 2023-08-20 | 2023-08-20 | SUM bag should fully functional and not damaged. |  | Deviation
2086 | 8NYA-Fz-X2210 SUB TCU not controlling properly | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-08-20 | 2023-08-20 | It is expected that the TCU control the temperature accurately and within processing ranges
2085 | Cerezyme - DGHBN011 - 8NYA - SUB X2230 DO and pH Loss | michael.knapp@sanofi.com | 8NYA 1403 | 2023-08-20 | 2023-08-20 | System is able to maintain DO and pH.
2084 | 8NYA-NEOGAA - UFDF-A Periodate buffer was warm during Oxidation | kenny.leland@sanofi.com | 8 NYA room 1503 | 2023-08-19 | 2023-08-20 | Buffer should be between 6- 10 C
2083 | 8NYA GEMBA: SUB X2230 on hold for out of range DO | lolav.ahmed@sanofi.com | 8 NYA | 2023-08-20 | 2023-08-20 | FBL-CST-000029
2082 | LAL Sample Leak Buffer Batch # DGFB1108 PN# GZ1287-52 | jeff.hill@sanofi.com | 68 NYA QC Micro Receipt Lab | 2023-08-19 | 2023-08-19 | QU-OPE-0182566, FBL-SOP-000080, ALF-SOP-000258 |  | Deviation
2081 | Pompe- UFDF A CIP frozen recipe | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2023-08-20 | 2023-08-20 | Automation intervention will be needed to get recipe restarted. MSAT input will be needed to ensure valves are set to correct setpoints
2080 | Pompe- UFDF-A Stock creation CX Value error | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2023-08-20 | 2023-08-20 | CX value information will need to be gathered and possibly manually entered | Yes | Deviation
2079 | Fabrazyme - DGHCE004 - 8NYA - SUB High Weight and ATF A Stoppage During ATF B Swapout | michael.knapp@sanofi.com | 8NYA 1333 | 2023-08-19 | 2023-08-19 | ATF A does not stop, and SUB perfusion continues uninterrupted during ATF B Changeout.
2078 | -8 NYA GEMBA Request- 805007 unable to enter Glycan load concentration | richard.manasian@sanofi.com | 1364 at 8 NYA | 2023-08-19 | 2023-08-19 | Value would be able to be manually input or value pulled in from related MBR
2077 | 8NYA-Fz-X2200 ATF B Failure | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-08-19 | 2023-08-19 | ATF B is expected to continue to operate normally during processing activities, and remain an integral system
2076 | Pompe ANX Eluate- CX Values failed to propagate through MES | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2023-08-18 | 2023-08-18 | MABC-QU-SOP-0019707 requires a Gemba with QA before manually inputting the CX values gathered from MES
2075 | Pompe ANX- Cloumn/ Skid statuses incorrect | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2023-08-18 | 2023-08-19 | Verify correct recipes were used for Equil and Production despite these incorrect statuses. Determine path forward to correct statuses in MES to align with the current state of the equipment.
2074 | 8NYA-Fz- X2210 Top DO probe installed in port dedicated to ABER, per FBL-EWI-000011 | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-08-18 | 2023-08-18 | It is expected that probes are installed in the order stated in FBL-EWI-000011
2073 | UV Meter unscannable on UFDF-A | Juan.Bonilla@sanofi.com | 8NYA 1503 | 2023-08-18 | 2023-08-18 | scan UV Meter in to UFDF-A record
2072 | 8 NYA GEMBA Request- FIT parameters not available in MES | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-08-18 | 2023-08-18 | Parameters being called for in MBR would be available in delta V
2071 | FRAD23E0610 Cerezyme - Lot DGHDT017 - 8 NYA - QC bioburden found leaking upon testing and compromised | judith.keller@sanofi.com | 68NYA/4700; 8NYA downstream | 2023-08-18 | 2023-08-18 | Sample cannot be compromised
2070 | Gemba - 8NYA - Cz - Leaked observed on ROSS fill bag. | stanley.durand@sanofi.com | 8NYA Cold Room 1220 | 2023-08-17 | 2023-08-18 | It's expected that bags are free of leaks
2069 | 8NYA-Pompe- 2831 SUM bag shows signs of agitator sleeve becoming unseated. Product has been collected in the SUM | john.lytle@sanofi.com | Rm 1343 8NYA | 2023-08-18 | 2023-08-18 | The agitator sleeve is installed properly seated per FBL-EWI-000015
2068 | 8NYA-Pompe-VI skid X2865 Mixer 2 probes need investigation due to erratic readings | john.lytle@sanofi.com | Rm 1343 8NYA | 2023-08-18 | 2023-08-18 | VI probes need to have acceptable response while monitoring the process conditions of the VI skid
2067 | Recipe went into hold due to temperature out of range. Load should be ambient. | kenny.leland@sanofi.com | 8 NYA room 1503 | 2023-08-18 | 2023-08-18 | Load should be at ambient temperature
2066 | MFG at 74 NYA received corrugate shippers p/n GZ7313-01 missing part number and box certificate stamp | alan.rousseau@sanofi.com | room 1017 74 NYA | 2023-08-17 | 2023-08-17 | Expected result is that the materials would have the stamped information as they have had since 2014
2065 | 8NYADOWNSTREAM DGHDT017 Cz shipping barrel pooling | marc.latta@sanofi.com | 1700 | 2023-08-17 | 2023-08-17 | pooling should go smoothly.No leak should not be observed
2064 | Water damage found above ceiling after drip observed in room 45-169 | nicholas.shoenfelt@sanofi.com | 45-169 | 2023-08-17 | 2023-08-17 | It is expected that water will not enter the clean room from outside the clean space.
2063 | 8NYA-GEMBA Pump tubing ruptured during set up. | alex.tello@sanofi.com | 8NYA RM1503 | 2023-08-17 | 2023-08-17 | path forward
2062 | CIP recipe started instead of WFI RInse on D2 | nicole.lufkin@sanofi.com | 51 NYA 140 | 2023-08-17 | 2023-08-17 | MABC-QU-SOP-0056555 |  | Deviation
2061 | 8NYA-Gemba-Downstream-PPQ 1 Copper-Pressure from CNC tote causing higher than set pressure | aidan.mula@sanofi.com | 8NYA RM1503 | 2023-08-17 | 2023-08-17 | Automation intervention to proceed and impact assessment
2060 | Process pH is identified to be out of range during chromatogram review of SP1 lot DGF3B11B. | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 138 | 2023-08-09 | 2023-08-16 | It is expected that the process pH is between 4.8-5.6 per FBL-SOP-000788 |  | Deviation
2059 | Fz Butyl eluate was not chilled before final filtration/transfer. | christopher.arnold@sanofi.com | 8RM1323 | 2023-08-17 | 2023-08-17 | That the eluted material would be chilled to temperature prior to starting the eluate filtration process.
2058 | 68NYA Room 1830-No product impact-Metrology Standard found out of tolerance during an unscheduled calibration. | robert.bauch@sanofi.com | 68-1830 | 2023-08-04 | 2023-08-17 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2057 | Wrong size Bag on BOM | caleb.coppola@sanofi.com | 8NYA 1350 | 2023-08-17 | 2023-08-17 | The BOM would be accurate for materials needed to produce the batch.
2056 | 805010 - 32 liter batch size - 20 liter bag in BOM, record | leo.gillespie@sanofi.com | 1350 - Solution prep | 2023-08-17 | 2023-08-17 | record and BOM should reflect the size of container compatible with the batch size - Hold Gembas with specialists to resolve root cause
2055 | 8NYA - RF1N5D001 805010 Liquid weigh out will not convert from LTR to grams | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-08-16 | 2023-08-16 | Hold Gemba to see path forward.
2054 | 8NYA DOWNSTREAM POMPE COPPER NRSR23SR8159 SAMPLE BAG ISSUE | marc.latta@sanofi.com | 1503 | 2023-08-16 | 2023-08-16 | sample bag should have MES/GMP approve label. Need to find path forward
2053 | GMP refresher training not completed by due date | Jonathan.Lowe@Sanofi.com | 8 NYA | 2023-01-01 | 2023-08-11 | Training was completed by the due date
2052 | 8NYA-Gemba-Pompe PPQ 1-UFDF-B Skid failed visual inspection | aidan.mula@sanofi.com | 8NYA RM1503 | 2023-08-16 | 2023-08-16 | Determination on path forward
2051 | Bioburden received within 10 minutes of sample expiry | sarah.adams@sanofi.com | 68 NYA | 2023-08-16 | 2023-08-16 | Sample to be delivered prior to sample expiry
2050 | SIP timer expired. | Devin.Parks@Sanofi.com | 1403 | 2023-08-16 | 2023-08-16 | SIP should have been started before timer expired
2049 | 8NYA PF8-room 1503-UV easy cal FB-4073 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 8-1503 | 2023-07-07 | 2023-08-16 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2048 | NeoGAA Copper Load Recipe Loop | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2023-08-16 | 2023-08-16 | Update to recipe needed to prevent this loop from occurring for future runs
2047 | 3000L SUM bag leaking | jermaine.kellam@sanofi.com | 8RM 1350 | 2023-08-15 | 2023-08-15 | Bag should not leak at all.
2046 | 8NYA Pompe Navigo Eluate #6 Leaking SU: 114292 | nicholas.candela@sanofi.com | 8 NYA / Rm. 1503 | 2023-08-15 | 2023-08-15 | Bag should not be leaking at time of setup
2045 | Missing Entries for Sampling Start and End Time on Settle Plate Paperwork | krystal.masseri@sanofi.com | 68NYA QCM 4700 | 2023-08-08 | 2023-08-15 | QU-SOP-0057388
2044 | Gemba - 8 NYA - Pompe UFDF-A membrane cleaning incorrect membrane incorrect cleaning status | Kenny.Leland@sanofi.com | 8 NYA 1503 | 2023-08-14 | 2023-08-14 | Membrane should be in correct status post CIP | Yes | Deviation
2043 | 8NYA-Fz-ATF A ruptured diaphragm/ perfusion hold | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-08-15 | 2023-08-15 | The ATF diaphragm should remain integral during operations
2042 | WFI loop 4 failed to complete a sanitization with in the calendar day of 13AUG23 | kenny.leland@sanofi.com | 8 NYA | 2023-08-13 | 2023-08-13 | WFI loop for should have completed a successful sanitization before the end on the calendar day 13AUG23.
2041 | 45NYA Room 1136-QCC45- Balance F-17179 was found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 45-1136 | 2023-08-01 | 2023-08-14 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2040 | Thyrogen sample cracked upon sample transport | michael.montiverdi@sanofi.com | 51 NYA | 2023-08-12 | 2023-08-12 | Integrity of samples should be maintained during the handling of sample logging and transport. |  | Deviation
2039 | 8NYA-Fz- Top media tote drained into bottom media tote due to incorrect setup | john.lytle@sanofi.com | 8NYA CNC space | 2023-08-04 | 2023-08-11 | It is expected that the top and bottom media totes are setup correctly, using pinch clamps and the associated automated valves to determine flowpath
2038 | 8NYA-Pompe-X2240 high pressure event due to vent filter wetted out | john.lytle@sanofi.com | 8nya Rm 1343 | 2023-08-13 | 2023-08-13 | It is expected that the level of foam in the SUB is maintained at suitable levels using regular antifoam dosing and process monitoring | Yes | Deviation
2037 | 8NYA - GEMBA- Sanitization and Flush will be skipped due to recipe running | Alex.tello@sanofi.com | 8NYA RM1503 | 2023-08-13 | 2023-08-13 | Path forward
2036 | Leak Discovered on Column 1 of Fz MCC skid X2510 | Brandon.Steff@sanofi.com | 8RM1333 8NYA | 2023-08-13 | 2023-08-13 | Waste totes are changed out in timely manner after they are filled to avoid having both waste totes filled at the same time.
2035 | Leaking 8NYA Cerezyme sample | stefanie.erickson@sanofi.com | 68NYA, QC Micro | 2023-08-12 | 2023-08-12 | Per QU-SOP-0073109 step 6.1 note viable microorganisms in bioburden samples are sensitive to handling and storage conditions. It is important to control, as much as possible, conditions of storage. It is expected that the bioburden sample will be secured and not leaking/compromised.
2034 | 804997 B13-90mM Na acetate, 10 mM acetic acid, 150mM glycine, pH 5.6--acetic acid not added | leo.gillespie@sanofi.com | 1350 - Solution prep | 2023-08-12 | 2023-08-12 | liquid should have been added prior to acknowledging mix timer
2033 | Small drip observed on media feed line at keleen-pak connector during refeed of D2 Harvest day 11 | adam.hildebrand@sanofi.com | 51 RM-140 | 2023-08-12 | 2023-08-12 | It is expected the kleen-pak connector is leak free during refeed
2032 | Net Weight of TSK Formulated Bulk DGF2B04 exceeded weight limit | Ryan.Geraghty@sanofi.com | 51NYA | 2023-08-11 | 2023-08-11 | Per QU-OPE-0291868 Step 4.7.1.5, the net weight of Formulated Bulk filtered directly into the shipping bag must be </= 38.5kg.
2031 | 8NYA: Leaked bioburden sample ID: 810871 DGHMD045 | bailey.gleason@sanofi.com | QC lab | 2023-08-11 | 2023-08-11 | Hold Gemba to see path forward.
2030 | 8NYA CZ DGHCG011- Sample tube cracked and leaking for sample 810772 CZ2_250_760844_POU760744_ENDO | larissa.fairweather@sanofi.com | 68NYA/4700/QCM | 2023-08-11 | 2023-08-11 | QU-SOP-0070558 Step 6.2.3 - Confirm sample containers are not cracked, leaking.
2029 | Broken Dock Plate at 45NYA hindering buffer movements | joshua.brooks@sanofi.com | 45NYA Loading dock | 2023-08-10 | 2023-08-11 | MABC-QU-SOP-0001975 6.4.2 Material is transported per associated documentation
2028 | 8NYA GEMBA - Butyl did not pass visual inspection | alex.tello@sanofi.com | 8NYA RM1323 | 2023-08-11 | 2023-08-11 | path forward
2027 | Duplicate tracking numbers issued for Athena Master Data Change Requests | timothy.pyne@sanofi.com | 51-107, 68-3350 | 2023-04-20 | 2023-08-11 | It is expected that tracking numbers issued for Athena Master Data Change Requests are sequential and not repeated.
2026 | Thyrogen - No Lots - 51 NYA- Incorrect document uploaded to Veeva | emily.brodsky@sanofi.com | 51 NYA | 2023-08-08 | 2023-08-11 | It is expected that the correct SOP was uploaded to MABC-QU-SOP-0073127 "Room Pressure Differential Monitoring of the Biologics Manufacturing Area at 45".
2025 | 50L bag equilibration steps not witnessed in batch record. | adam.hildebrand@sanofi.com | 51 New York Avenue, RM 152A | 2023-04-17 | 2023-08-10 | It is expected that step is witnessed and signed for at time of operation.
2024 | OneLIMS lot 31546 unable to authorized for lot/product release | craig.stevenson@sanofi.com | TBD | 2023-08-10 | 2023-08-10 | Trying to lot release
2023 | Skid pressure going over the high limit during column prime recipe | zachary.rickard@sanofi.com | 8NYA/1503 | 2023-08-10 | 2023-08-10 | Recipe should run without going over the high pressure
2022 | Action Level alarm during hood ops | david.loik@sanofi.com | 51 NYA | 2023-08-10 | 2023-08-10 | The climet would not alarm
2021 | Climet Alarm Excursion >5.0 | david.loik@sanofi.com | 51 NYA | 2023-08-10 | 2023-08-10 | Climet would be capped after it's turned off
2020 | Missed verifier on shaker table rpm rate | david.loik@sanofi.com | 51 NYA | 2023-08-09 | 2023-08-10 | Step would be verified
2019 | Multiple - 8NYA - Cryostat freezer High High Temp A alarm | michael.knapp@sanofi.com | 8NYA 1403 | 2023-08-10 | 2023-08-10 | Temperature is maintained within control.
2018 | Pompe copper setup record buffer discrepancy | zachary.rickard@sanofi.com | 8NYA CNC | 2023-08-10 | 2023-08-10 | Record and BOM will be the same in what buffer we should use
2017 | 8NYA GEMBA: Cz Vial storage unit blocked in ERP, unable to perform take-out in MES | lolav.ahmed@sanofi.com | 8NYA 1280 | 2023-08-10 | 2023-08-10 | Operator should be able to perform take-out of the vial in MES after one- step transport performed from cryostat dedicated rack (FR5000001-2) to WMS01 for vial retrieval
2016 | DGHCE004, Recipe On Hold Due to Perfusion EM Failure | wilfredo.lopez@sanofi.com | 8NYA Room 1333 | 2023-08-09 | 2023-08-09 | It is expected that the media rate of charge remains unchanged during media tote switchover. The perfusion pumps are expected to remain at a rate below 23-24RPMs independently.
2015 | Fz butyl buffer takeout | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-08-10 | 2023-08-10 | Buffer takeout will occur real time on the record when the CNC breakdown is occurring
2014 | Bioburden and Endotoxin failures for Pompe | craig.stevenson@sanofi.com | 8NYA | 2023-08-07 | 2023-08-09 | No contamination
2013 | 45NYA Room 1136-QCC45- Balance F-14979 was found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 45-1136 | 2023-08-08 | 2023-08-10 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2012 | 51NYA-APT CC Room 140-Conductivity Indicating Transmitter F-5991-013 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 51-140 | 2023-08-09 | 2023-08-10 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
2011 | DGHAF002, Harvest SUM Overflow Pump Waste Filter Failure | wilfredo.lopez@sanofi.com | 8NYA Room 1403 | 2023-08-09 | 2023-08-10 | It is expected that the material being routed to waste goes through the guard filter unobstructed without building pressure in the system.
2010 | The prompt to tare 200L SUM PBM-1520001 did not appear on the Delta-V screen. WFI fill started unprompted while waiting for Automation intervention. | jermaine.kellam@sanofi.com | 8RM 1350 | 2023-08-09 | 2023-08-09 | The prompt should have appeared on screen and allowed us to verify to start the WFI fill.
2009 | 8 NYA GEMBA Request- 804996 is calling for the buffer to be chilled | richard.manasian@sanofi.com | CNC area at 8 NYA | 2023-08-09 | 2023-08-09 | Storage conditions in MBR would match the CST document. Issue was noticed in engineering runs
2008 | 8 NYA GEMBA Request- Tote bag failed to properly deploy | richard.manasian@sanofi.com | CNC area at 8 NYA | 2023-08-09 | 2023-08-09 | Bag would deploy correctly allowing for complete fill
2007 | D1 temperature spike caused a recipe hold requiring automation intervention | david.loik@sanofi.com | 51 NYA | 2023-08-09 | 2023-08-09 | The temperature would maintain within the specifications.
2006 | Incubation Start Time discrepancy for Thyrogen Peptide Map Assay | Kaylyn.Manoogian@sanofi.com | 45NYA, QC Chemistry Lab | 2023-06-29 | 2023-08-09 | Time would've been recorded in military time to avoid any confusion.
2005 | N2 Bulk PO number discrepancy | noah.zahn@sanofi.com | 80NYA | 2023-08-09 | 2023-08-09 | N/A
2004 | 8NYA-Fz-RF503D003 MCC record unable to terminate, needs MES intervention and QA input of missed skid breakdown step | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-07-15 | 2023-08-09 | MES record issues should be resolved with MES team input. MES records should be terminated at the conclusion of production activities.
2003 | Fz Butyl column cycle takeout could not be done | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-08-09 | 2023-08-09 | Column cycle will be taken out automatically at the end of the production recipe.
2002 | SUT failure on harvest SUM PBM-2816001's harvest overflow pump P-2817001 tubing. SUM is connected to SUB X-2240 and MCC skid X-2831. | Brendan.Dickhaut@sanofi.com | 8NYA/Room 1403 | 2023-08-09 | 2023-08-09 | Tubing doesn't fail during normal operations.
2001 | 8NYA- High pressure on X2831 column 3 | Brandon.steff@sanofi.com | 8NYA1403 | 2023-08-08 | 2023-08-08 | Column inlet pressure stays below alarm limit of 20PSI
2000 | Rotameter F-17409 damaged and unable to generate "As Found" calibration results | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2023-08-08 | 2023-08-08 | It is expected that equipment is able to generate "As Found" results during routine calibration.
1999 | 76 NYA PUR DBYHAS007 “Incorrect solution used to adjust the pH". | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2023-08-08 | 2023-08-08 | Per FBS-MBR-000062 step# 7.17.16. It's required to use 1N HCL to adjust the pH.
1998 | GEMBA for Regen 2 buffer not installed correctly in RF639D010 (BHT1582001) | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-08-08 | 2023-08-08 | Buffer installation performed in both shared CNC MCC records so disposal can also be performed in both records when totes are done being used.
1997 | GEMBA for Biolims plan 202302971-1 no longer in biolims | Brandon.Steff@sanofi.com | 8RM1343 | 2023-08-07 | 2023-08-07 | After biolims samples are collected and sent to QC lab, QC is able to receive the samples in biolims.
1996 | Fz Butyl Skid column inlet conductivity and column outlet conductivity reading are out of range | marc.latta@sanofi.com | 1323 | 2023-08-07 | 2023-08-07 | per Del
1995 | Inspection step missing during the installation of Butyl eluate sum | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase1 Room 1323 | 2023-08-07 | 2023-08-07 | DGHBE035 | Yes | Deviation
1994 | Pompe DGHBN002 , Column Pressure issue during load | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2023-08-04 | 2023-08-04 | Skid should not go to hold due to pressurize Column | Yes | Deviation
1993 | pH OOR for TSK DGF3B03A, SP1 DGF4B10A and SP1 DGF4B10C during chromatogram review | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 132 and 138 | 2023-07-15 | 2023-08-07 | It is expected that the process pH be in range per FBL-SOP-000788.
1992 | Downstream parts soaked in upstream soak tank at 51NYA | nicholas.shoenfelt@sanofi.com | 51-122B | 2023-08-06 | 2023-08-07 | SP2 parts are downstream parts, it is expected they be cleaned in a downstream soak tank.
1991 | pH value not recorded DBYFRM034 QU-OPE-0309296 | dianne.moustafa@sanofi.com | 76 NYA Room 364 Fermentation | 2023-08-02 | 2023-08-07 | QU-SOP-0309296 pH should be between (6.95-7.00) and be recorded.
1990 | DGHAF002, Harvest SUM Overflow and Loss of Containement | wilfredo.lopez@sanofi.com | 8NYA Room 1403 | 2023-08-06 | 2023-08-06 | It is expected that the harvest overflow pump turn on when the high weight limit of the vessel is reached. The pump appeared to be on as indicated on the display (reading 100.2rpm) but was not actually running.
1989 | Recipe failed to automatically launch on deltaV | wilfredo.lopez@sanofi.com | 8RM1350 | 2023-08-06 | 2023-08-06 | Recipe is expected to automatically launch once the step is executed in MES. | Yes | No Deviation
1988 | Pompe Media Take Out Issues | robert.morawski@sanofi.com | 8NYA CNC | 2023-08-05 | 2023-08-06 | SCADA PI pull in DGHAF003 gave a correct volume to consume, leaving media volume to be issued to DGHAF002.
1987 | Drip tubes exceeded weight range | Juan.Bonilla@sanofi.com | Column Packing 8NYA Room 1303 | 2023-08-05 | 2023-08-05 | Drip tubes to be within the target range
1986 | Freezer 1421 was OOS in MES - Cleaning requirements don't align with other freezers | Nicholas.grote@sanofi.com | 8 NYA RM 1421 | 2023-08-04 | 2023-08-04 | Freezer in service
1985 | Cerezyme - DGHCG010 - 8NYA - Wave Bioreactor Not Sampled on Day 0 Following Scale-Up | michael.knapp@sanofi.com | 8NYA 1403 | 2023-08-03 | 2023-08-04 | Day 0 sampling is performed on wave reactor following scale up
1984 | Cz Media Take Out Issues | robert.morawski@sanofi.com | 8NYA CNC | 2023-08-04 | 2023-08-04 | DGHBN009 had an accurate SCADA pull, and did not consume the full media tote volume.
1983 | Non Viable data not available for guard filter changeout per 2021FRACC0009 | derek.pandolfo@sanofi.com | Guard filter change out occurs in the LFH in room 1343 at 8NYA. | 2023-06-28 | 2023-08-03 | Per 2021FRACC0009 The EM monitoring (Viable and Non-Viable) will be performed during process and surface sampling will be performed after. The EM data will be reviewed by Manufacturing and Contamination Control.
1982 | 1L hTSH wave cell count exceeded allowable range | adam.rowe@sanofi.com | 51NYA, room 152 | 2023-08-04 | 2023-08-04 | Total cell count would be within allowable range in the MBR |  | Deviation
1981 | Omitted steps in FBL-MBR-000109 during SP2 DGF4B02B process | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 132 | 2023-07-23 | 2023-08-03 | FBL-MBR-000109 requires that the gross weight of the SP2 filtered eluate be recorded per step 4.5.18 and the net weight of the SP2 filtered eluate per step 4.5.19
1980 | Bleed valve closed on active recipe | Ralston.najarro@sanofi.com | CNC AREA, 8NYA1403 | 2023-08-03 | 2023-08-03 | Recipe for X2230 was ended bleed valve shouldve continued to be open due to X2220 still having an active recipe
1979 | Wrong Fz Butyl empty column cleaning recipe | marc.latta@sanofi.com | 8nya/1303 | 2023-08-03 | 2023-08-03 | Recipe should not ask for priming and maintenance. | Yes | No Deviation
1978 | Process pH is identified to be out of range during chromatogram review from PI | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 131 | 2023-07-20 | 2023-08-03 | It is expected that the process pH is between 7.1-7,9 per Fbl-sop-000788 | Yes | Deviation
1977 | HA Lot - DBYFRM034; Chiller pump status reading off | Charlie.Houle@sanofi.com | 76 NYA; Room 364 (HA Fermentation) | 2023-08-01 | 2023-08-01 | The chiller pump status indicator on the PLC should be read "On"
1976 | Fz butyl column CV results out of range | marc.latta@sanofi.com | 8nya/1303 | 2023-08-03 | 2023-08-03 | SWAB 4 RESULT MUST BE BELOW 1.18 ppm | Yes | No Deviation
1975 | HA Lot - DBYFRM034; pH Probe Malfunction led to Re-SIP'ing V-3020 | Charlie.Houle@sanofi.com | 76 NYA; Room 364 (HA Fermentation) | 2023-08-01 | 2023-08-01 | Per MABC-QU-SOP-0020197 step 6.1 the housing with electrode is placed in 37C pH buffer and have reading of 6.96 - 7.00.
1974 | Thyrogen/Buffer Prep - Operator signed for performed by and witnessed by at the same step. | Jeremy.Wooden@Sanofi.com | 80 NYA 101/102 | 2023-08-01 | 2023-08-02 | Per step 6.11.1 of MABC-QU-SOP-0074684 (Good Documentation Practices) Personnel will not review, verify or witness a step they have performed.
1973 | Leak occur at VI manifold. | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2023-08-02 | 2023-08-02 | Manifold should not be leak
1972 | HA Lot - DBYHAH002 - Missing Witness Step | charlie.houle@sanofi.com | 76 NYA; Room 3101 (HA Pre-Purificaiton) | 2023-01-27 | 2023-08-02 | Per FBL-SOP-000942, witnessed by steps must be signed at time of execution.
1971 | PN692302 ( NANOFILTER ASCEPTIC TUBING) not on BOM for Resin Discard but is needed for the process | marc.latta@sanofi.com | 8nya/1303 | 2023-08-02 | 2023-08-02 | BOM and FBL-EWI-000016 should be align
1970 | GEMBA for X2250 high pressure event and dropping capacitance | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-08-01 | 2023-08-01 | Bioreactor stays under 0.5 PSI
1969 | Pompe - RF1NAD003 - 8NYA - SUB DO Probe Drift | michael.knapp@sanofi.com | 8NYA 1403 | 2023-08-01 | 2023-08-01 | Probe readings are accurate and have redundancy.
1968 | R2A media used for dynamic operations, 76NYA fermentation, instead of TSA media | megan.botti-marino@sanofi.com | 76 NYA | 2023-07-11 | 2023-08-01 | TSA media is to be used per FBL-SOP-000755
1967 | Cz Thaw/Pool ambient timer exceeded | zachary.rickard@sanofi.com | 8NYA/1703 | 2023-07-29 | 2023-07-29 | Record will be executed at time of performing steps
1966 | Two (2) GMP Preventative Maintenance (PM) Work Orders - 2854789 and 2854790 - Not Completed by Due Date of 31July2023 | tanni.arif@sanofi.com | Multiple (76-347 and 45 rooftop) | 2023-08-01 | 2023-08-01 | PMs are supposed to be completed/extended on-time
1965 | 51NYA-APT DS-Conductivity Recorder FB-0036 overdue for calibration | robert.bauch@sanofi.com | 51 | 2023-08-01 | 2023-08-01 | BQ-GOP-000046 Calibration Program Version 6 Section 5.7.5.5 Instrument should have been received into Metrology before it's due date
1964 | A gemba is needed for: "76 NYA PUR DBYHAH002 Missing Trainer countersignature."". | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2023-02-03 | 2023-07-31 | FBS-MBR-000093: Each individual performing a task as part of their training process should have a countersignature by a trainer per FBS-SOP-0000942
1963 | 8NYA- GEMBA Request- 805011 record is calling for 1299.694kg of 834732 (L-Histidine Monohydrochloride) should be grams | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-07-31 | 2023-07-31 | Amount that is weighed out should be similar to what was done during shakdown runs and MES records reflect
1962 | Bioburden Pompe samples 749485 and 759882 were both collected forsampling point D0 of lot: DGHAF002 | roodline.bertrand@sanofi.com | 8NYA | 2023-06-07 | 2023-07-30 | TBD
1961 | Can not locate Climet FB-1073 | adam.rowe@sanofi.com | 51NYA, room 140 | 2023-07-31 | 2023-07-31 | Equipment should be on-site and delivered to metrology
1960 | Cap not placed on Purge Valve port, and Purge Valve not closed completely. Leak observed during media transfer DBYFRM033, QU-OPE-0309296 | dianne.moustafa@sanofi.com | 76 NYA Room 364 Fermentation | 2023-07-31 | 2023-07-31 | QU-SOP-0052455 step 6.4.2 close purge valve, step 6.48 cap valve port. QU-OPE-0309296 step 7.6.14 isolate fermentor.
1959 | O2 Bulk OOS reuslt | noah.zahn@sanofi.com | 80NYA | 2023-07-31 | 2023-07-31 | NBO-SPC-000241 step 6.1
1958 | HA - No Lots - 76 NYA - An approved version of FBL-SOP-000941 was used instead of an effective version | craig.stevenson@sanofi.com | 76NYA, Room 355 | 2023-07-19 | 2023-07-31 | It is expexted that effective SOPs are sued to perform the EM assay
1957 | Bioreactor hold due to SUB pressure and ph failure: base buffer pump speed controller interlock. | Brendan.Dickhaut@sanofi.com | 8NYA, Room 1333 | 2023-07-30 | 2023-07-30 | MSAT contacted. trouble shot per FBL-EWI-000004 and adjusted gasses per FBL-EWI-000011. Recipe resumed once PVs were within thier SP's |  | No Deviation
1956 | CZ thaw record not recognizing the correct freeze bags | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1212 | 2023-07-29 | 2023-07-29 | its expected that all bags are assigned a SU that is recognized by the record
1955 | DO probe wire broke/severed on 1L wave bag | nicole.lufkin@sanofi.com | 51 NYA 152A | 2023-07-28 | 2023-07-29 | the DO probe should not be cut, alarm for DO at 1930 on the 28th when noone from manufacturing was in lab
1954 | Steps not completed in record | Ralston.najarro@sanofi.com | CNC AREA | 2023-07-28 | 2023-07-28 | Steps should've been signed for at time of execution
1953 | 76 NYA PUR DBYHAH002 “Missing performed by and Trainer countersignature." | munish.kumar2@sanofi.com | 76 NYA/Room#3102/3103/Purification | 2023-01-31 | 2023-07-28 | FBS-MBR-000062: Steps are supposed to be performed with performed by signature and there should not be any empty spaces left per FBS-SOP-000942. FBS-MBR-000093: Each individual performing a task as part of their training process should have a countersignature by a trainer per FBS-MBR-0000942
1952 | Action level excursion for nonviable sample during setup of Dynamic EM TSH | Adriana.Alban@sanofi.com | 51-152A | 2023-04-10 | 2023-07-28 | Per MABC-QU-SOP-0030277 step 6.4.9 it is expected that comments need to be added into collection notes in onelims in the event of an excursion. |  | No Deviation
1951 | X2250 Vent filter follow up | Devin.Parks@Sanofi.com | 1403 | 2023-07-28 | 2023-07-28 | N/A
1950 | SUT Bag Defect | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1703 | 2023-07-28 | 2023-07-28 | Return the bag to the vendor and replace for a new bag |  | Deviation
1949 | 8NYA EM Collection batch 202302736-1 expired samples | cleiton.lopes@sanofi.com | 8NYA and QC Micro | 2023-07-27 | 2023-07-27 | All samples to be picked and delivered by MLC during the cycle and EM to be incubated within 24hr from sampling | Yes | Deviation
1948 | Missing expiry dates for opened bottles of trypan blue and crystal violet at 51NYA | adam.rowe@sanofi.com | 51NYA, room 140 | 2023-07-27 | 2023-07-27 | Expiry of both materials should be in our cell counting document.
1947 | 8NYA-Pompe-Re-GEMBA automation updates to sparge lines/ Train 3 & Train 4/ Upstream phase 2 Rm 1343 | john.lytle@sanofi.com | 8NYA, Rm1343 | 2023-07-27 | 2023-07-27 | automation update complete
1946 | DGHAF002 and DGHAF003 Recipe Holds Due to Perfusion Failure Events | wilfredo.lopez@sanofi.com | 8NYA Room 1403 | 2023-07-27 | 2023-07-27 | It is expected that the necessary buffer/media sources are installed and ready for operation.
1945 | DGHAF003 SUB 2250 Recipe Hold Due to a High Pressure Event | wilfredo.lopez@sanofi.com | 8NYA Room 1403 | 2023-07-27 | 2023-07-27 | It was expected that the operator installs a manual pinch clamp on the upstream section of the attached condensate collection bag tubing and remove the lower clamp to allow the liquid to drain in to the bag. The procedure would not require sealing the vent line pathway.
1944 | Pompe Media Take Out Issues | robert.morawski@sanofi.com | 8NYA CNC | 2023-07-27 | 2023-07-27 | SCADA numbers produce the correct value for the media tote take out. Storage unit of buffer still exists in the tote.
1943 | 76NYA Room 347-no product impact-Facilities-Conductivity Transmitter found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-347 | 2023-07-26 | 2023-07-27 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
1942 | BIRF form was not filled out when putting media into warm room. | adam.rowe@sanofi.com | 51NYA, room 140 | 2023-07-23 | 2023-07-26 | When media is moved into the warm room, the date and time need to be recorded.
1941 | Hole found in Sterile HA Powder sample pouch Lot DBYHAS005 P/N 5615-30 (Stability Study Sample) | Akintunde.Akinbuli@sanofi.com | 11 Forbes Rd, NBO, Room 1316 Stability Area | 2023-07-26 | 2023-07-26 | It is expected that container closure of stability samples is not compromised | Yes | Deviation
1940 | ASM Butyl Column Pack Record | zachary.rickard@sanofi.com | 8NYA/1303 | 2023-07-26 | 2023-07-26 | Damaged spool piece would have been replaced during vendor PM and column packing operations would continue as planned.
1939 | GEMBA: 8NYA 23SR8123 and 23SR8124 executed in duplicate and incorrectly | judith.keller@sanofi.com | 68NYA/4700; 8NYA upstream | 2023-07-22 | 2023-07-25 | It is expected that NRSRs are performed according to NRSR instructions and executed once per the NRSR.
1938 | 8NYA GEMBA: Incorrect POU sampled Cz for 250mL media equil. | lolav.ahmed@sanofi.com | 8 NYA 1343 | 2023-07-26 | 2023-07-26 | N/A | Yes | Deviation
1937 | 76 NYA PUR BBYHAS008 "FBS-MBR-000056 Missing witness step 7.17.5 and wrong filter lot number on integrity test attachment" | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2021-12-10 | 2023-07-25 | Per the FBS-SOP-000924, step # 6.1.1 and 6.1.2 requires that 1) there should not be any empty spaces within the documents (batch records) as well as the signatures are required to be entered at the time the actions are performed; FBS-MBR-000056, step# 7.4.6 requires that a pre-use integrity test is performed on the filters used during the manufacturing process.
1936 | Ineffective CAPA - Effectiveness Review FRAD22R0110 | gregory.klotz@sanofi.com | 51 NYA | 2023-07-26 | 2023-07-26 | Expected that the ER would be effective. See FRAD22R0110 for details of ineffective CAPA
1935 | 8NYA-GEMBA Request- 705661 leak during transfer RF545D074 DGHBA074 | richard.manasian@sanofi.com | CNC area at 8 NYA Fill Station | 2023-07-26 | 2023-07-26 | Transfer line would not be leaking
1934 | Missing NV Data for Dynamic Sampling | katherine.milliken@sanofi.com | 1343 8 NYA | 2023-06-17 | 2023-07-24 | It is expected that the data would be parsed at the time of the dynamic sampling
1933 | Sporkzlenz Expirations not assigned according to procedure | Edward.kozatek@sanofi.com | 51NYA, Rm 140 | 2023-07-20 | 2023-07-20 | The expiries of prepared wipes should align with MABC-QU-SOP-0073029 |  | No Deviation
1932 | 8NYA-Fz-MCC X2510 Outlet Conductivity failing during CIP operations/ post NaOH WFI rinse | john.lytle@sanofi.com | 8NYA, Rm 1333 Upstream Phase 1 | 2023-07-26 | 2023-07-26 | conductivity result of ,/= 2.1 uS/cm during WFI rinse
1931 | DBYFRM031 did not reach target viscosity FBS-MBR-000058 | dianne.moustafa@sanofi.com | 76 NYA Room 364 | 2023-07-25 | 2023-07-25 | Viscosity target value is ≥1000 cps within maximum usage of 45.1 kg of NaOh
1930 | incorrect time entry recorded and witnessed | Juan.Ramos-Torres@sanofi.com | 76 NYA Room 3102 | 2023-07-14 | 2023-07-26 | witness should have ensured all data is complete and accurate.
1929 | X2250 ATFs stopped running after valve failure. | Devin.Parks@Sanofi.com | 1403 | 2023-07-25 | 2023-07-25 | ATFs should have been cycling normally. |  | Deviation
1928 | GEMBA: 8NYA 760741 Media Formulation/point of batch sample resampled without documenting | judith.keller@sanofi.com | 68NYA/4700; 8NYA solution prep/ 8 NYA laydown area | 2023-07-23 | 2023-07-25 | Resamples need to be documented in onelims and MES.
1927 | Spill event due to pressure Transducer rupture on X2250 eluate SUM disassembly | Devin.Parks@Sanofi.com | 1403 | 2023-07-25 | 2023-07-25 | Pressure transducer should not rupture. | Yes | No Deviation
1926 | X2250 vent filters wetted out | Devin.Parks@Sanofi.com | 1403 | 2023-07-25 | 2023-07-25 | pH range is 6.8-7.0
1925 | 68NYA Room 4700 Dept QCM68-TOC's FB-2424, FB-2426 and FB-2427 were found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 68-4700 | 2023-07-18 | 2023-07-25 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
1924 | 51 NYA rhTSH Purification-atypical A280 results for Q lot DGF2B04A fraction 7 | jacob.kallon@sanofi.com | 45 NYA QC Chem Lab | 2023-07-24 | 2023-07-24 | It is expected that A280 results for fraction 7 be in typical reportable range
1923 | GEMBA for leak on column 1 colder steam valve during buffer prime | Yvesaintoine@sanofi.com | 8RM1333 8NYA | 2023-07-24 | 2023-07-24 | Sample bags are clamped off prior to starting buffer prime so they do not fill with liquid. Column lines remain unobstructed to allow flow of buffer through lines.
1922 | 8NYA-Cerezyme-SUM bag leaking during initial WFI addition | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-07-24 | 2023-07-24 | Determination at Gemba
1921 | Total SIP time not met DBYFRM031 FBS-MBR-000058 step 7.6.11.10/7.6.12 | dianne.moustafa@sanofi.com | 76 NYA Room 364 | 2023-07-24 | 2023-07-24 | SIP must be for a total of a minimum of 30 minutes.
1920 | Bubble trap low level alarm on MCC skid X2731 (Cz) | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-07-24 | 2023-07-24 | Bubble trap remains full for duration of the campaign.
1919 | X2510 bubble trap air filter wetted out | wilfredo.lopez@sanofi.com | 8 NYA 1333 | 2023-07-23 | 2023-07-23 | Expect air filter to get wet
1918 | 8NYA-Fz-Steam trap block valve found to be leaking during SIP operations | john.lytle@sanofi.com | Rm 1333, 8NYA | 2023-07-22 | 2023-07-22 | integral MCC skid, valves not leaking
1917 | Pompe - RF1NAD003 - 8NYA - pH Probe Disabled Following Erratic Readings | michael.knapp@sanofi.com | 8NYA 1403 | 2023-07-20 | 2023-07-20 | Probe readings are accurate. | Yes | No Deviation
1916 | 2 Individuals incomplete on GMP training | gregory.klotz@sanofi.com | 51 and 8 NYA | 2023-06-30 | TBD | Expect that all personnel performing GMO tasks will be trained on the task (specifically current version of SOP, as applicable) |  | Deviation
1915 | Pest Observation in Room 51NYA-140 | craig.stevenson@sanofi.com | 51 NYA Room 140 | 2023-07-20 | 2023-07-20 | It is expected that the controled clean rooms be pest free
1914 | Mold Recovery - P/N 768234 lot DBYSEP034 | paul.consolmagno@sanofi.com | Seprafilm Manufacturing? | 2023-06-21 | 2023-07-19 | N/A
1913 | 76 NYA PUR DBYHAS005 “Discrepancy in product loss during transfer from Tank V-3531 to Tank V-3532". | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2023-07-19 | 2023-07-19 | Per FBS-MBR-000056 step# 7.8.27 the expected transfer amount should be less than or equal to 2175 kg.
1912 | Leaking Tube Seal Downstream of Amber Clamp | thomas.loik@sanofi.com | 8RM1350 | 2023-07-18 | 2023-07-18 | If the Amber clamp is confirmed to have been present prior to the leak starting. The contents of the tote would have remained isolated from the leak at all times. The tote would usable.
1911 | 8NYA- Cerezyme CM equil and sanitization buffer not taken out from the record | stanley.durand@sanofi.com | 8NYA Room 1503 | 2023-06-03 | 2023-07-18 | It's expected for buffer calcs and take out to be performed during Q membrane operations per FBL-EWI-000034
1910 | 76 NYA PUR DBYHAS005 “Tank V-3521 Nitrogen pressure was out of specification ". | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2023-07-14 | 2023-07-18 | Per FBS-MBR-000062 step# 7.7.14 Tank V-3521 nitrogen pressure should be between 5-10 psi and the actual pressure measured and entered on the batch record was 10.64psi.
1909 | 76 NYA PUR DBYHAS005 “Tank V-3521 Nitrogen pressure was out of specification ". | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2023-07-14 | 2023-07-18 | Per FBS-MBR-000062 step# 7.7.14 Tank V-3521 nitrogen pressure should be between 5-10 psi and the actual pressure measured and entered on the batch record was 10.64psi.
1908 | During rhTSH Blue Column Operation chromatogram review, two UV peaks were observed during regeneration 1. FBL-SOP-000788 states there should be only one UV peak during regeneration 1. | craig.valery@sanofi.com | 51 NYA | 2023-07-18 | 2023-07-18 | FBL-SOP-000788 states there should be only one UV peak during regeneration 1. |  | Deviation
1907 | 8NYA-CC8 room 1333-product impact unknown FB-3995 UV easy cal found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1403 | 2023-06-22 | 2023-07-18 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
1906 | Temperature was not in range on fermentation chart, DBYFRM007, FBS-MBR-000058 | dianne.moustafa@sanofi.com | 76 NYA Room 364 | 2023-01-26 | 2023-07-18 | FBS-MBR-000058 step 7.11.20.1 temperature range 35-39C
1905 | Pompe - RF1NAD002 - 8NYA - DO Probe Drift identified on system | michael.knapp@sanofi.com | 8NYA 1403 | 2023-07-18 | 2023-07-18 | DO probes are both accurate and do not drift
1904 | Pompe - RF1NAD003 - 8NYA - VI Skid X2825 pH Calibration Error | michael.knapp@sanofi.com | 8NYA 1403 | 2023-07-14 | 2023-07-17 | It is expected that pH is accurate, so that viral inactivation is successful.
1903 | GDP inconsistency observed by Sodexo cleaning team during cleaning audit | caleb.coppola@sanofi.com | 8NYA | 2023-07-16 | 2023-07-16 | Per FBL-SOP-000942, steps are being recorded real time by the operator physically performing the steps
1902 | GEMBA: Vendor granted access with generic admin account to Nexus | judith.keller@sanofi.com | 68 NYA 4700 QCM | 2023-07-11 | 2023-07-14 | It is expected vendors are set up with unique user names to allow traceability of log in into system.
1901 | No negative control results for Dynamic EM CGFSB007 | Adriana.Alban@sanofi.com | 51-152A | 2022-06-04 | 2023-07-13 | Per FBL-SOP-000755 6.19.3 sample results should be uploaded to LIMS
1900 | 8NYA Cerezyme Manual entry of buffer weights performed without a second verifier. | stanley.durand@sanofi.com | 8NYA room 1212 | 2023-06-01 | 2023-07-17 | It's expected when manual entries of buffer weights are performed a second verifier is required. | Yes | Deviation
1899 | VI mixer X2825 pH Calibration complete prompt not acknowledged. | Devin.Parks@Sanofi.com | 1403 | 2023-07-14 | 2023-07-15 | Prompt is answered following calibration so MCC skid can continue to go through normal column sequence | Yes | Deviation
1898 | DO probe 2 on Cz SUB X2220 diverging, needs to be disabled | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-07-14 | 2023-07-15 | DO probes stay aligned and control to setpoint
1897 | GEMBA for leak in sampling line on Cz SUB X2220 | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-07-15 | 2023-07-15 | Welds are integral and do not leak
1896 | Process Temperature is identified to be out of range during chromatogram review from PI | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 132 | 2023-07-06 | 2023-07-14 | It is expected that the process temperature is between 6-10C per FBL-MBR-000057.
1895 | Crack observed inside in during ATF setup for wetting | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-07-14 | 2023-07-14 | ATF in undamaged upon opening
1894 | ER FRAD23R0025 deemed to be ineffective per BQ-GOP-000249 causing a new deviation event to be created in accordance with BQ-GOP-000248 | paul.mclaughlinjr@sanofi.com | 8NYA | 2023-07-12 | 2023-07-12 | CAPA would be deemed effective
1893 | 8NYA Upstream GEMBA SUB X-2210 ATF B Rupture DGHCE003 | gary.dodakian3@sanofi.com | 8NYA / 1333 | 2023-07-13 | 2023-07-13 | ATF diaphragm are supposed to stay integral throughout the entire bioreactor run.
1892 | The temperature of resistivity probe located in 51-CLN STM GEN 1 (X-9130) is reading 77°F on the transmitter at all times. | brizaida.medina-vega@sanofi.com | 51NYA Mechanical Room 51-146 | 2023-07-13 | 2023-07-13 | Changing by the time
1891 | FZICB sample submitted to QC Micro for ELISA testing in incorrect container | daniel.whitcroft@sanofi.com | 8 NYA | 2023-07-13 | 2023-07-14 | FBL-RPT-000376, Table 16, Sample container for Host Cell Proteins parameter is defined as polypropylene tube
1890 | Incorrect filter part number entered into filter tester DBYFRM006 FBS-MBR-000058 | dianne.moustafa@sanofi.com | 76 NYA Room 364 Fermentation | 2023-01-24 | 2023-07-14 | FBS-MBR-000058 Step 7.6.31 correct part number should be entered into filter tester
1889 | Fabrazyme - DGHCE004 - 8NYA - Aseptiquik membrane tear during vent filter connection to SUB | michael.knapp@sanofi.com | 8NYA 1333 | 2023-07-13 | 2023-07-13 | Aseptiquiks work
1888 | 8NYA-Fz-Probe installation not documented at time of execution | john.lytle@sanofi.com | 8NYA Rm 13333 | 2023-05-10 | 2023-07-13 | probe installation should have been captured at the time of execution
1887 | CTMS EM Plates QC#:23MS01426 bag dropped and 3 plates chip/cracked after receipt. | karen.ramos2@sanofi.com | 68 NYA QC Sample Receipt Rm 4720 | 2023-07-13 | 2023-07-13 | Plates to be received without incident (drops)
1886 | Logbook missing during Q2 logbook audit - Thyrogen Equipment Prep | nicholas.shoenfelt@sanofi.com | 51NYA | 2023-06-01 | 2023-07-13 | It is expected that all logbooks are accounted for in the quarterly audit.
1885 | Pompe - RF1NAD002 - 8NYA - Investigate VI skid for Pump Tubing Wear and Tear | michael.knapp@sanofi.com | 8NYA 1403 | 2023-07-13 | 2023-07-13 | It is expected that the tubing is integral and that there is no noticeable wear and tear on the pump tubing specifically. |  | No Deviation
1884 | Manifold use on 705677 buffer | zachary.rickard@sanofi.com | 8NYA/1210 | 2023-06-10 | 2023-07-12 | Operation should be performed based off of EWI instructions
1883 | 8NYA GEMBA - Equipment not available/no instructions on ewi | Alex.tello@sanofi.com | 8NYA RM1350 | 2023-07-12 | 2023-07-12 | path forward
1882 | Documentation Correction needed for Late Entry | sarah.adams@sanofi.com | 68 NYA | 2023-07-08 | 2023-07-12 | Documentation to occur contemporaneously
1881 | Cerezyme - DGHBN006 - 8NYA - First Harvest Day Sampling Not Performed | michael.knapp@sanofi.com | 8NYA 1403 | 2023-05-07 | 2023-07-12 | First Harvest Day samples are collected per FBL-EWI-000009 on the first day of harvest collection.
1880 | Leaking Bioburden Sample Delivered to QC Sample Management | sarah.adams@sanofi.com | 68 NYA | 2023-07-11 | 2023-07-11 | Sample to be sealed closed correctly
1879 | Fz SP Eluate SUM weight discrepancy | nathaniel.barnes@sanofi.com | 8NYA 1323 | 2023-07-11 | 2023-07-11 | Determine product impact and path forward for Zinc processing
1878 | MCC skid went into hold due to clogged HIBOD waste filter | Brendan.Dickhaut@sanofi.com | 8NYA/Room 1403 | 2023-07-11 | 2023-07-11 | Expected results is that if a filter needs to be replaced, there should be a section in the MES record to document the changeout.
1877 | Chart Recorder Chart missing data DBYFRM029 FBS-MBR-000058 Monitoring Fermentor | dianne.moustafa@sanofi.com | 76 NYA Room 364 | 2023-07-10 | 2023-07-10 | FBS-SOP-000491 chart recorder should record continuously.
1876 | SP Elution transfer delta v. time expired | Christian.Venegas@Sanofi.com | 8NYA Downstream Room 1323 | 2023-07-11 | 2023-07-11 | To ensure that the transfer timer was completed before proceeding with Filtration | Yes | 556 Event
1875 | Thyrogen Drug Product Lot CW0831: An adverse trend was detected for Reconstitution Time at 12 Month Stability Time Point | sylvia.gonzalez@sanofi.com | 45 NYA, NBO | 2023-07-10 | 2023-07-10 | No adverse trend is expected
1874 | 68NYA-Metrology-no product impact-Metrology Standard S-1359-001 no as found data taken standard lost by shipper | robert.bauch@sanofi.com | 68-1830 | 2023-07-10 | 2023-07-11 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
1873 | Hexylene Glycol Components Missed in C&Q Protocols | stephanie.aceto@sanofi.com | 8NYA Tank Farm | 2023-07-11 | 2023-07-11 | All components listed within the Component Impact Assessment and shown on the systems P&ID's would be commissioned & qualified within either the systems respective IV or IQ protocols. | Yes | Deviation
1872 | 8NYA- GEMBA Request- Supplier Deviation for defective 3000L SUM bag | william.lee@sanofi.com | 8NYA/1350 | 2023-07-07 | 2023-07-07 | SUM Bag should have DETIKER PG160 CLAMP attached to bottom outlet
1871 | Rounding Error for Sample Low Load TSH Purity SDS PAGE Assay | john.kapaniris@sanofi.com | 45NYA QC Chemistry | 2023-06-20 | 2023-07-06 | sample load calculation should equal 109.9 Henry Wu recorded 110.0 | No | No Deviation
1870 | MFG identified a partial bucket of L-Glutamine (Material # 726192). The NBO warehouse label is for batch 22GN02346 the HU label is for batch 22GN02345 | michael.montiverdi@sanofi.com | 8NYA | 2023-07-07 | 2023-07-07 | The correct HU labels are applied to the correct material batch number.
1869 | GEMBA for Pompe PA and CO2 sparge path routed to mocrosparger instead of drilled hold sparger (DHS) | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-06-01 | 2023-07-10 | PA and CO2 are routed through DHS in deltaV, not micro sparge | Yes | No Deviation
1868 | Water observed dripping from ceiling tile in soak tank room 51-122B | nicholas.shoenfelt@sanofi.com | 51-122B | 2023-07-10 | 2023-07-10 | It is expected that area clearance (no leak) will be maintained in the manufacturing area.
1867 | 8NYA - Incorrect material added to SUM for 705676 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-07-10 | 2023-07-10 | Scratch batch, reorder material, re batch
1866 | A leak was found at the top of the product SUM bag during Fz Butyl Eluate transfer | christopher.arnold@sanofi.com | 8RM1323 | 2023-07-09 | 2023-07-09 | That there will be no leaking upon starting the transfer.
1865 | Bag leak | maria.palmar@sanofi.com | 8 New York Ave | 2023-07-09 | 2023-07-09 | Drain WFI, use a new bag and start from inflation with a new kit
1864 | 8NYA- GEMBA-Missing sample for DGHBZ015 | richard.manasian@sanofi.com | 8NYA room 1350 | 2023-06-10 | 2023-07-09 | Samples would have been received in BioLims |  | Deviation
1863 | Leaking connection on E3 bioreactor during refeed | david.loik@sanofi.com | 51 NYA | 2023-07-09 | 2023-07-09 | The media refeed would complete without leaking.
1862 | 8NYA-GEMBA- Leak on SUM bag and pump in alarm | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-07-09 | 2023-07-09 | SUM bag lines would not leak
1861 | Pompe - RF1NAD003 - 8NYA - Eluate SUM collected Additional Eluate | michael.knapp@sanofi.com | 8NYA 1403 | 2023-07-09 | 2023-07-09 | Eluate SUM for Pompe switches to next SUM after weight reading is >190kgs. SUM valves pinch off tubing and ensure proper eluate collection.
1860 | 8NYA-CZ2-CZ Recipe did not allow bag to reach set volume. | Juan.Bonilla@sanofi.com | 8NYA 1703 | 2023-07-08 | 2023-07-08 | The recipe was supposed to look at the sum weight instead of the bag weight.
1859 | Pompe - RF1NAD003 - 8NYA - High Pressure Event on Bioreactor | michael.knapp@sanofi.com | 8NYA 1403 | 2023-07-08 | 2023-07-08 | System is controlled without experiencing high pressure events during processing.
1858 | 8NYA-Gemba Request- Buffer 705663 Leak during QS | alejandro.martinez2@sanofi.com | 8NYA Solution Prep, 1350 | 2023-07-08 | 2023-07-08 | Recirculation would stay connected during QS
1857 | 8NYA GEMBA: Leak identified on CZ MCC skid X2560 buffer pump tube connections | Lolav.ahmed@sanofi.com | 8NYA 1403 | 2023-07-08 | 2023-07-08 | Upon completion of pre-CIP setup, pressure test, and SIP the skid remains integral
1856 | 8NYA Cz pooling, exceeded at ambient temperature timer | zachary.rickard@sanofi.com | 8NYA/1703 | 2023-07-07 | 2023-07-07 | Product would be finished pooling within the 180 minutes
1855 | GEMBA 8NYA - BOM item scanned but unable to be used. | alex.tello@sanofi.com | 8 NYA Room 1350 | 2023-07-05 | 2023-07-07 | Path forward
1854 | Pompe - RF1NAD003 - 8NYA - High Pressure Event on Bioreactor | michael.knapp@sanofi.com | 8NYA 1403 | 2023-07-07 | 2023-07-07 | System is controlled without experiencing high pressure events during processing.
1853 | Undocumented training during UF processing | craig.valery@sanofi.com | 51NYA | 2023-07-05 | 2023-07-07 | It is expected that documentation of training be performed through out entire training.
1852 | incorrect values on flow meter totalizers | Jeffrey.boldizar@sanofi.com | 8 NYA 1403 | 2023-07-07 | 2023-07-07 | Totalizer to maintain flow volume
1851 | Pompe - RF1NAD003 - 8NYA - Leak Identified at Tube Seal Post-VI Skid | michael.knapp@sanofi.com | 8NYA 1403 | 2023-07-07 | 2023-07-07 | There are no leaks associated with tube sealing. |  | Deviation
1850 | No data for temptale KJG6N07RB0 - Cerezyme shipment DGHDS006 | margaret.hentz@sanofi.com | Sanofi Waterford - IDA Industrial Park, Old Kilmeaden Road, Ireland | 2023-07-05 | 2023-07-05 | Per BQ-FRM-000692 'C-Safe Acutemp RKN Receipt and Shipping Checklists' Version 3.0, all 4 temptales required to record data |  | Deviation
1849 | Cerezyme - DGHBN010 - 8NYA - Bleed Control did not adjust properly upon starting Capacitance Control | michael.knapp@sanofi.com | 8NYA 1403 | 2023-07-07 | 2023-07-07 | Bleed control switches per recipe without intervention |  | No Deviation
1848 | 76NYA-Room 347-Facilities- Conductivity Meter F-13587-002 on the MECO pure steam gen 2 FEED INLET CONDUCT TEMP CELL B found out of tolerance during a scheduled Calibrations | robert.bauch@sanofi.com | 76-347 | 2023-07-06 | 2023-07-07 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
1847 | 76 NYA PUR DBYHAH003 “Solution temperature is out of specification ". | munish.kumar2@sanofi.com | 76 NYA/Room#3101/Initial Purification | 2023-06-28 | 2023-07-06 | Temperature must be between 20-30C.
1846 | Eluate SUM for RF1NAD003 did not deploy correctly, resulting in high bag pressure | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-07-06 | 2023-07-06 | Bag deploys correctly, is able to achieve target weight of 190Kg minimum for switchover
1845 | Pompe SUB X2250 calibrated to non-representative BGA sample | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-07-06 | 2023-07-06 | BGA measurement does not exceed alarm limits of recipe
1844 | Gemba – 8NYA – Tube Welder failed during welding process | william.lee@sanofi.com | 8NYA CNC area | 2023-07-05 | 2023-07-05 | Tube Welder should have stayed powered on during process.
1843 | 8NYA-Cerezyme- Filter wetting recipe launched manually from DeltaV instead of EBR | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-07-05 | 2023-07-05 | Determination at Gemba
1842 | GEMBA 51NYA: Thyrogen TSK: Load exceeds maximum TSK product to be loaded | william.bradley@sanofi.com | 51NYA RM132 | 2023-07-05 | 2023-07-05 | Product loading should be between the dedicated range.
1841 | TSK Eluate Pool Sampled in Incorrect Container | cindy.guerra@sanofi.com | 51NYA - 132 | 2023-07-04 | 2023-07-05 | Eluate Pool samples would be collected and submitted to QC in the correct polystyrene container |  | No Deviation
1840 | DGHAD006, 760720 Samples taken during formulation incorrectly labeled as point-of-use samples for shared buffer | wilfredo.lopez@sanofi.com | 8NYA Room 1350 | 2023-04-25 | 2023-07-03 | It is expected that all samples are taken and labeled correctly per the sampling plan.
1839 | DGHAD003, 760720 Buffer Lot Not Found in BioLIMS | wilfredo.lopez@sanofi.com | 8NYA CNC space | 2023-07-03 | 2023-07-03 | It is expected that the material lot is int he system so that labels can be printed and required sampling can be performed.
1838 | BOM item not used during batching of 760741 Cz Media | christopher.arnold@sanofi.com | 8NYA 8RM1350 | 2023-07-04 | 2023-07-04 | That all parts listed in the BOM will be used per the guiding EWI to perform the tasks required.
1837 | Leaking 8NYA Pompe Media Sample (Part: 805013) | krystal.masseri@sanofi.com | 68 NYA QCM | 2023-07-04 | 2023-07-04 | It is expected that bioburden samples will be secure and will not be leaking/compromised
1836 | 8NYA- 8NYA- CZ2 DS Sample 763797 spike CV failure for endotoxin assay acceptance criteria | Rebecca.Schmidt@sanofi.com | QC Micro Nexus bench | 2023-06-23 | 2023-07-04 | FBL-SOP-000960 6.4.2, 6.7.3 | Yes | Deviation
1835 | SUM bag leak | maria.palmar@sanofi.com | 8 New York Ave | 2023-07-03 | 2023-07-03 | Drain WFI left on the bag, replace the whole kit with a new one, and remake the buffer.
1834 | Fabrazyme - DGHSD004 -8NYA - OneLims Sampling Error of TSA and Rodac Plates | michael.knapp@sanofi.com | 8NYA 1343 | 2023-07-03 | 2023-07-03 | Correct plates are identified for their activity.
1833 | Unused SUT Assigned to Released MO's | adam.keister@sanofi.com | 8 NYA, 1503 | 2023-07-03 | 2023-07-03 | Path Forward
1832 | Loading material exceed loading capacity of TSK column. | craig.valery@sanofi.com | 51NYA RM132 | 2023-07-03 | 2023-07-03 | It is expected that the loading material be within the column loading capacity
1831 | Cerezyme - DGHBN010 - 8NYA - Media Installation Not Documented In Batch Record | michael.knapp@sanofi.com | 8NYA 1403 | 2023-07-01 | 2023-07-03 | Media installation is captured in both records when performed.
1830 | Sample ID 674113, Fabrazyme Bioburden Vial Thaw receive after it expired. | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2023-07-03 | 2023-07-03 | Bioburden sample submitted within 24 hours
1829 | Cerezyme - DGHBN010 - 8NYA - Bleed Setup Not Performed Prior to Bleed Requirement | michael.knapp@sanofi.com | 8NYA 1403 | 2023-07-03 | 2023-07-03 | Bleed tubing is set up prior to bleed start
1828 | 51NYA TSK Lot DGF4B02B: Incorrect eluate used (DGF2B04 used in error; should've used DGF4B02) | William.Bradley@sanofi.com | 51NYA RM132 | 2023-07-03 | 2023-07-03 | Manufacturing to use the correct lot DGF4B02
1827 | Pompe - RF1NAD002 - 8NYA - Insufficient Sample Quantity Taken for pNP Activity | michael.knapp@sanofi.com | 8NYA 1403 | 2023-07-02 | 2023-07-02 | Correct QTY of sample is taken.
1826 | bug found in 51 NYA room 140 | nicole.lufkin@sanofi.com | 51 NYA 140 | 2023-07-01 | 2023-07-01 | no bug in suite
1825 | Clamp found on media line | dana.hopkins@sanofi.com | 8 NYA CNC space | 2023-07-01 | 2023-07-01 | Lines should be free of clamps |  | Deviation
1824 | SUB X-2250 recipe unable to aquire base pump | dana.hopkins@sanofi.com | 8 NYA room 1403 | 2023-07-01 | 2023-07-01 | Recipe should acquire pumps
1823 | malfunctioning 2 bottle sampler post seed inoc E3 | nicole.lufkin@sanofi.com | 51 NYA 140 | 2023-07-01 | 2023-07-01 | sample valve would open, after inspection by EQP, they determined the stem may have been stuck on the A valve.
1822 | X-2250 N2 valve would not open | dana.hopkins@sanofi.com | 8 NYA room 1403 | 2023-07-01 | 2023-07-01 | Needed N2 per FBL-CST-000030 |  | Deviation
1821 | Pompe Media tote value simulated for switchover due to suspected bad profibus cable. | Devin.Parks@Sanofi.com | CNC Room 1203 | 2023-07-01 | 2023-07-01 | DeltaV should read the media tote weight correctly. |  | Deviation
1820 | 45NYA - Particles discovered falling from ceiling tiles | cindy.guerra@sanofi.com | 45NYA / Rm 171 and 186 | 2023-06-30 | 2023-06-30 | Particles would not be discovered in area
1819 | UPLC Micropipette FB-3494 failed "as found" calibration. Used for Fz UPLC testing | Brandon.Steff@sanofi.com | 8RM1334 8NYA | 2023-06-29 | 2023-06-30 | Micropipette remains within "as found" tolerance when sent out for calibration
1818 | X2220 - ATF A straight line connected to bottom port instead of top port. | Devin.Parks@Sanofi.com | 1403 | 2023-06-30 | 2023-06-30 | Per FBL-EWI-000011 section ATF10 Setup, step 1, Straight line should be connected to top port and filter tube should be connected to bottom port.
1817 | Failed RQ of 70-QC-0765 in 45 NYA | megan.perkins@sanofi.com | 45 NYA / Room 1131 | 2023-06-28 | 2023-06-29 | Per PQ 08-0882, "All individual temperatures for each Temperature Distribution sensor location within usable room/chamber space are within the user specified criteria over the entire monitoring period. Criteria: (2°C to 8°C)." Per PQ 08-0882, "The grand mean of all individual Temperature Distribution sensors located within usable room/chamber space is within the user specified control tolerance from the controller setpoint. Criteria: (3°C to 7°C)." Per PQ 08-0882, "All individual temperature for each Temperature Penetration sensor are within the acceptance criteria over the entire monitoring period. Criteria: (2°C to 8°C)." Per PQ 08-0882, "Difference between Maximum Distribution average and average of the TC adjacent to the EMS sensor Criteria: (±1.5°C)." Per PQ 08-0882, "Temperature Penetration sensors do not rise above the acceptance criteria during the defrost period of the refrigerator. Criteria: (2°C to 8°C)."
1816 | 8NYA - Media expiry conflict | bailey.gleason@sanofi.com | 8NYA | 2023-06-30 | 2023-06-30 | Ideal outcomes are listed in DOC Media Gemba 1
1815 | 8 NYA GEMBA- Determining RoSS PCS Alarms for DGHDT001 DGHDT002 | cory.lefort@sanofi.com | 8 NYA | 2023-06-27 | 2023-06-27 | It is expected that when a PCS alarm is generated in MES it is also generated in ANMS.
1814 | 4 missing logbooks from 74 NYA quarterly audit | vincent.palumbo@sanofi.com | 74 NYA | 2023-06-30 | 2023-06-30 | It is expected that no logbooks would be lost and all of the logbooks would be accounted for. |  | Deviation
1813 | UFDF final yield calc is out of range | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-06-29 | 2023-06-29 | Final yield Calc would be greater than or equal to 85%
1812 | Issues with VI PH Cal record | Jeffrey.boldizar@sanofi.com | 8 NYA 1403 | 2023-06-29 | 2023-06-29 | Execution in record properly
1811 | Two Media lot DGHAW055 and 54 not consumed in SUB record RF639D005 and Rf639D006 | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2023-05-19 | 2023-06-29 | Media batch DGHAW054 and DGHAW055 should take out
1810 | 8NYA-Fz-Eluate SUM found to have shifted installation and will not allow for agitator mixing per process requirements. Additional manipulation may lead to bag integrity issues | john.lytle@sanofi.com | Rm 13333 | 2023-06-29 | 2023-06-29 | manual mixing of the SUM contents is required per FBL-EWI-000009 prior during sampling operations
1809 | Missed HEPA VAC PM's Performed | thomas.loik@sanofi.com | 8RM1350 | 2023-06-01 | 2023-06-28 | Bi-Weekly PM will be scheduled on MFG schedule.
1808 | Expired Conductivity standard used prior to making multiple buffers at 8NYA | richard.manasian@sanofi.com | 8RM1350 | 2023-06-28 | 2023-06-28 | That no expired material will remain inside a workroom. That the weekly/monthly cleanings conducted on 18JUN23 or 25JUN23 would have caught the expired material.
1807 | Fz UFDF ran out of Diafiltration buffer | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-06-28 | 2023-06-28 | Buffer would not run out
1806 | Climet used with incorrect time | valeria.moreno@sanofi.com | 76-355 | 2023-06-27 | 2023-06-27 | Time 0634
1805 | 51 NYA Missing time out entry for movement out of cold room 108C | vincent.palumbo@sanofi.com | 51 NYA | 2023-06-17 | 2023-06-28 | It is expected that the timeout would be recorded on the log sheet at the time of occurrence |  | Deviation
1804 | GEMBA for DGHBN005 first harvest day samples not sampled in biolims | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-04-25 | 2023-06-26 | After harvest is started, bioburden and endotoxin samples taken and sampled in biolims
1803 | Tyvek Holder Assembly PN: GZ2486-01 LN: 23MN00163 Vendor: Oliver Healthcare Packaging Fails for Visual Inspection | fatima.fragaalvarez@sanofi.com | NBO | 2023-04-27 | 2023-04-27 | tyvek holders meet specs in NBO in order to receive the material for production in Framingham
1802 | Tyvek Holder Assembly PN: GZ2486-01 LN: 23MN00163 Vendor: Oliver Healthcare Packaging Fails for Visual Inspection | karen.hempstead@sanofi.com | NBO | 2023-04-27 | 2023-06-27 | tyvek holders meet specs in NBO in order to receive the material for production in Framingham
1801 | Cerezyme Samples incubated for 6 days instead of 3-5 days | Alyssa.Barrett@sanofi.com | 68 NYA/Room 4700 | 2023-05-06 | 2023-06-27 | Per FBL-SOP-000019 Attachment 2, samples should've been tested on liquids form FBL-FRM-000705
1800 | Meeting to Discuss incorrect time recorded on BIRF for Harvest DGF2B0805 | sylvester.kegbeh@sanofi.com | 80NYA ROOMS 139 | 2023-05-17 | 2023-06-26 | It is expected that the correct time be recorded per FBL-SOP-000371.
1799 | GEMBA for Aber probe 2 needs to be disabled on Pompe SUB X2250 | sheila.pandolfo@sanofi.com | 8 NYA Room 1403 | 2023-06-26 | 2023-06-26 | It is expected the Probe 2 reading would not exhibit a sudden drop in the capacitance reading. | Yes | No Deviation
1798 | Missing Falcon Batch Record Issued to FAD21-072 'TFF Skid and Bulk Filtration Batch Record' on 13May21 | damian.barbato@sanofi.com | 76NYA | 2023-03-31 | 2023-03-31 | It is expected that the batch record issued would be returned by Manufacturing and reconciled by QA and the Records Management Group |  | No Deviation
1797 | 51NYA-EP45/51-product impact unknown-Conductivity Meter FB-1343 turned in for cal with incorrect sensor | robert.bauch@sanofi.com | 45-172 and 51-134 | 2023-06-26 | 2023-06-27 | BQ-GOP-000046 section 6 Loop Calibrations
1796 | Fz SP process exceeding the maximum allowed timer of 210 minutes | edgar.vale@sanofi.com | 8NYA, Room 1323 | 2023-06-26 | 2023-06-26 | Process would complete within the allowable 210 minutes
1795 | SIP Failure | jeffrey.mei@sanofi.com | 51 NYA 140 | 2023-06-26 | 2023-06-26 | Per FBL-SOP-000680, the SIP should have continued. The SIP was able to be restarted within the 25.5 hr limit. No product impact expected.
1794 | Samples bags were not entered into the record | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1323 | 2023-06-17 | 2023-06-25 | its expected that all SUT used for sampling will be captured in the record | Yes | No Deviation
1793 | Addition time of Serum not recorded in media record FBL-MBR-000033 | nicholas.shoenfelt@sanofi.com | 51-124 | 2023-06-22 | 2023-06-25 | It is expected that the time the serum is added to the media prep tank the time be recorded in the MBR. |  | Deviation
1792 | 76 NYA PUR DBYHAH003 “Missing witness/verifier steps on batch record FBS-MBR-000062 | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2023-06-26 | 2023-06-26 | Steps supposed to be performed with Witness/verifier per FBS-SOP-000942
1791 | 8NYA- Fz- 705691 Seed media transfer is too slow | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-06-26 | 2023-06-26 | Functionality in recipe to perform manual transfer for seed cart under slower speeds, and increased transfer rate with fixed weight settings for the tote transfers.
1790 | Debris was found in line when attaching column 1 on X2811 for Pompe. | Devin.Parks@Sanofi.com | 1403 | 2023-06-25 | 2023-06-25 | There should be no debris in line | Yes | Deviation
1789 | ATF wye pull tab rip | Jeffrey.boldizar@sanofi.com | 8 NYA 1403 | 2023-06-25 | 2023-06-25 | Expect connection to work properly
1788 | FIT Test did not launch | maria.palmar@sanofi.com | 8 New York Ave | 2023-06-23 | 2023-06-23 | Manually launch FIT test
1787 | Purification tank T-2911 went post-use expired and cleaned without CVs, then used. | alyson.belleville@sanofi.com | 51-139 | 2023-06-16 | 2023-06-23 | It is expected that if the equipment goes post-use it expired, it is cleaned with CVs and provisionally released by QA before forward processing.
1786 | ANX column resin concentration too high | zachary.rickard@sanofi.com | 8NYA/1303 | 2023-06-22 | 2023-06-22 | Slurry concentration would be between 40%-60% |  | Deviation
1785 | 8NYA-Pompe-Process looking for a 1k tote for Base Setup operations, currently utilizing 200L causing logic failure in MES | john.lytle@sanofi.com | Rm 1210\ CNC | 2023-06-23 | 2023-06-23 | Proper tote identification/ installation for base setup
1784 | 8NYA- Cz Nanofilter Failed 3rd FIT | alejandro.martinez2@sanofi.com | 8NYA Solution Prep, 1350 | 2023-06-23 | 2023-06-23 | Automation to update of cz nanofilter FIT recipe
1783 | Particle found in 125ml flask culture | adam.rowe@sanofi.com | 51NYA, room 152A | 2023-06-23 | 2023-06-23 | No particulates should be in the cell culture | Yes | Deviation
1782 | 8NYA-Cedex-MFG ran out of Cedex deproteinizer and obtained some from MSAT. MSAT had the deproteinizer in larger bottles, so the manufacturer’s item number was different than what is listed in FBL-EWI-000026 | john.lytle@sanofi.com | 8NYA-Upstream, multiple | 2023-06-23 | 2023-06-23 | GZ50-583-392 material available
1781 | Insect found in unclassified Gowning Room 204 | nicholas.macdonald@sanofi.com | 76-204 | 2023-06-22 | 2023-06-22 | It is expected that gowning room 76-204 be free from pests
1780 | While batching 760741, during initial WFI, SUM bag leaked | Evans.kamau@sanofi.com | 8NYA1350 | 2023-06-22 | 2023-06-22 | Expected result the sum supposed not leak.
1779 | Lot Disposition Authorization form not completed prior to performing disposition. | timothy.pyne@sanofi.com | 74 NYA, 51 NYA, 8 NYA | 2022-02-02 | 2023-06-22 | It is expected Attachment 1 of FBL-SOP-000567 would be completed prior to performing Quality Dispositions.
1778 | DBYFRM027 Kill tank mixer speed reduced after 53 minutes | nicholas.macdonald@sanofi.com | 76-364 | 2023-06-21 | 2023-06-22 | It is expected that the mixing speed be reduced after approximately 15 minutes.
1777 | Filter not in Contact with media for bioburden sample FZ2_CB_762729_EQUIL | jessica.kychler@sanofi.com | QC Microbiology, 68 NYA floor 4 room 4700 | 2023-06-16 | 2023-06-22 | It is expected that the filter be adhered to the media per step 9.2.8.13 of ALF-MTD-000018
1776 | Pompe VI skid did not elute correctly for X2865 | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-06-22 | 2023-06-22 | Acid is added to eluate, it is mixed, then base is added to neutralize before being sent to eluate SUM
1775 | 68NYA-Metrology-no product impact-Metrology Standard S-784 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 68-1830 | 2023-05-17 | 2023-06-22 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
1774 | 705665, 1.0 M NaOH batch leaking at tube/clamp junction | leo.gillespie@sanofi.com | 1350 - Solution prep | 2023-06-21 | 2023-06-21 | Expected results is that the tubing and clamp do not leak
1773 | 8NYA - Alarms not captured in MES record for RF622D010 | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-06-21 | 2023-06-21 | Hold Gemba to see path forward. |  | No Deviation
1772 | 8NYA-Fabryzyme- 5 foot extension line (714020) not used with exception referencing FRAD23A0065. | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-06-21 | 2023-06-21 | Determination at Gemba
1771 | 8 NYA GEMBA- Cz CM Eluate Collection pH probe not engaged. | cory.lefort@sanofi.com | 8 NYA 1503 | 2023-06-21 | 2023-06-21 | It is expected that the pH probe is engaged in the SUM bag during the pH adjustment.
1770 | rhTSH Form Prod bioburden sample submitted to QC Micro with insufficient volume for testing | amy.roy@sanofi.com | 68NYA/4700/QC Micro | 2023-06-20 | 2023-06-21 | 100mL of rhTSH Form. Prod should be submitted to QC Micro for bioburden testing
1769 | Instructions missing from FBL-EWI-000019 | christopher.brewster@sanofi.com | 8 NYA, Rm. 1350 | 2023-06-20 | 2023-06-20 | The instructions for how to perform these steps should have been included in FBL-EWI-000019.
1768 | Cerezyme - DGHBN009 - 8NYA - New Media Installation Failure | michael.knapp@sanofi.com | 8NYA 1403 | 2023-06-20 | 2023-06-20 | Installation is performed correctly per FBL-EWI-000011.
1767 | Agar plate missing lid | lauren.ruffing@sanofi.com | 68NYA 4700 QC Micro | 2023-06-19 | 2023-06-19 | it is expected that every plate has a lid
1766 | Pompe - RF1NAD003 - 8NYA - DO Setpoint 50% rather than expected 60% | michael.knapp@sanofi.com | 8NYA 1403 | 2023-06-20 | 2023-06-20 | DO setpoint is correct
1765 | DGFB1089 Buffer Manifold Leak | jeff.hill@sanofi.com | 45 NYA | 2023-07-20 | 2023-07-20 | FBL-MBR-000053 Step 4.21, Set up the filter per FBL-SOP-000080 | Yes | Deviation
1764 | Filter stuck to the lid instead of TSA agar for bioburden sample 749995. | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2023-06-15 | 2023-06-20 | For the filter to be placed on the agar for the duration of incubation.
1763 | Agar plate missing lid | lauren.ruffing@sanofi.com | 68NYA 4700 QC Micro | 2023-06-19 | 2023-06-19 | it is expected that every plate has a lid
1762 | 51 NYA Missing time stored entry on BIRF for media movement | vincent.palumbo@sanofi.com | 51 NYA | 2023-03-03 | 2023-06-20 | It is expected that the correct time would be recorded after a time correction
1761 | Pompe - RF1NAD002 - 8NYA - VI Skid pH Probes Not Installed Properly | michael.knapp@sanofi.com | 8NYA 1403 | 2023-06-19 | 2023-06-19 | Probes are installed in correct SUM bags per EWI. |  | Deviation
1760 | Cz ross fill, final bag underweight | zachary.rickard@sanofi.com | 8NYA 1703 | 2023-06-19 | 2023-06-19 | Record will calculate accurately for all bags to be filled within the expected range.
1759 | 8NYA GEMBA - pH was tested under the incorrect standardization. | jack.laplante@sanofi.com | 8NYA RM1350 | 2023-06-19 | 2023-06-19 | Looking for path foward.
1758 | 760741--XVF Media with DTT - Tote SUT leak | leo.gillespie@sanofi.com | CNC Tote filling area | 2023-06-18 | 2023-06-18 | tote bag expected not to leak
1757 | 8NYA-Cerezyme- Filter failed 3 times in post FIT after transfer of 760744 media. | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-06-18 | 2023-06-18 | Determination at Gemba
1756 | LAL sample 750112 (CZ2_MMC_761218_ELUATE_ENDO) was submitted in polypropylene tubes. | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2023-06-16 | 2023-06-18 | Per FBL-SOP-000960 step 6.4.3, "Ensure all samples are received and diluted in polystyrene or depyrogenated glass tubes prior to sample analysis."
1755 | Bag two had leak from sterile connection. | kevin.ornell@sanofi.com | 45 NYA room 190 | 2023-06-18 | 2023-06-18 | Connection would not leak during fill. Per SOP FBL-SOP-000080 step 6.21.1 and 6.21.2 in the event of a connection leak manifold and bags may be disconnected and new manifold bag maybe connected using Kleenpak connectors and fill continued.
1754 | Nano filtration transfer failed | Christian.Venegas@Sanofi.com | 8NYA upstream Room 1503 | 2023-06-17 | 2023-06-17 | To be able to replace SUT's and continue with the transfer
1753 | 8NYA - Pompe PPQ 2-Buffer 805004 RF1N1D002 is calling for 200L bag and adaptor for weigh out | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-06-18 | 2023-06-18 | Update guidance in EWI-19 to allow for this type of weigh out and adjust the record to allow for weigh out flexibility.
1752 | 8NYA Recipe went into hold during Nano filtration | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1503 | 2023-06-17 | 2023-06-17 | Its expected that pressure will reach set target during product filtration
1751 | 8NYA-Pompe PPQ2- Adaptor missing from BOM for 805002 transfer setup. | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-06-17 | 2023-06-17 | Part should exist on the BOM
1750 | Media tote was leaking from seal | dennis.bazinet@sanofi.com | 8NYA CNC area | 2023-06-17 | 2023-06-17 | All sealed lines should have a clamp attached after it is sealed
1749 | A SUM bag began to leak during initial WFI fill | leo.gillespie@sanofi.com | 8RM1350 | 2023-06-16 | 2023-06-16 | Batch would proceed without leaks from the SUM
1748 | CNC Tote Leak at Connection | nathaniel.barnes@sanofi.com | 8NYA 1212 (Phase 2 DS CNC) | 2023-06-16 | 2023-06-16 | Determine if forward processing is ok with the buffer
1747 | 8NYA-Pompe PPQ 2-Probe autoclave cycle aborted with critical alarm "steam pressure pure steam too low" | aidan.mula@sanofi.com | 8NYA RM1373 | 2023-06-16 | 2023-06-16 | Determination at Gemba
1746 | Air handler down in room 152, 51 NYZ | nicole.lufkin@sanofi.com | 51 NYA 152 | 2023-06-16 | 2023-06-16 | air handlers running
1745 | Documentation missing for autoclave run 116845 at 51NYA | nicholas.shoenfelt@sanofi.com | 51-129 | 2023-06-12 | 2023-06-13 | It is expected that the autoclave cycle printout will be uploaded to MES.
1744 | 8NYA-Pompe PPQ 2- Phosphoric Acid weigh out function in MES does not convert liters to grams | aidan.mula@sanofi.com | 8NYA/RM1350 | 2023-06-16 | 2023-06-16 | Record should convert liters to grams and allow operators to proceed without manual conversion.
1743 | Pompe PPQ 1 Grow day 3 daily sampling not performed | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2023-06-13 | 2023-06-15 | Sampling event performed daily for SUB to check Vicell, BGA and Cedex sample per FBL_EWI-000009 and FBL-RPT-001201
1742 | HA Lot - DBYFRM026; error code 'err 001' during SIP setup | Charlie.Houle@sanofi.com | 76 NYA; Room 364 (HA Fermentation) | 2023-06-14 | 2023-06-14 | It is expected that no error code should occur, and the setup should continue to SIP of V-3020.
1741 | Pompe - RF1NAD002 - 8NYA - ATF Guard Filter Pressure Pendotech not reading properly | michael.knapp@sanofi.com | 8NYA 1403 | 2023-06-14 | 2023-06-14 | It is expected that the pressure gauge gives an accurate readout of the pressure in the post-ATF line.
1740 | Pompe - RF1N9D003 - 8NYA - 50L Wave Recipe not ready for inoculation | michael.knapp@sanofi.com | 8NYA 1403 | 2023-06-14 | 2023-06-14 | It is expected that the wave is fully equilibrated in MES and DeltaV prior to inoculation.
1739 | NBO-SOP-000122 omitting Ethanol in V 6.0 | joshua.brooks@sanofi.com | 80NYA | 2023-06-14 | 2023-06-14 | More specific language used in scope of NBO-SOP-000122.
1738 | 45NYA RM.176 TSH - Material Airlock Card Reader malfunction (12B-20) | cindy.guerra@sanofi.com | 45NYA / Rm 176 | 2023-06-14 | 2023-06-14 | Badge read would be functioning properly |  | No Deviation
1737 | Cerezyme - DGHBN009 - 8NYA - Do and Capacitance Probes not enabled at start of campaign | michael.knapp@sanofi.com | 8NYA 1403 | 2023-06-11 | 2023-06-14 | All probes are enabled upon start of bioreactor campaign.
1736 | Climet alarm went off during Fill setup in room 76-354 | valeria.moreno@sanofi.com | 76 NYA- room 354 | 2023-06-13 | 2023-06-13 | FBS-SOP-000607
1735 | FBL-SOP-000743 Equipment and Room Usage Log for the Manufacturing Facility at 74 NYA was obsoleted for the Falcon project (Biologics) but is still needed and in use by 74 NYA Biosurgery | alan.rousseau@sanofi.com | Room 1009 | 2023-06-13 | 2023-06-13 | FBS-SOP-000454 and numerous others. See attached list.
1734 | HA Lot - DBYFRM025; error code 'err 001' during SIP setup | charlie.houle@sanofi.com | 76 NYA; Room 364 (HA Fermentation) | 2023-06-13 | 2023-06-13 | It is expected that no error code should occur, and the setup should continue to SIP of V-3020.
1733 | Frozen Mod Enzymes on Dry Ice were put into Cold Wrong (incorrect storage conditions) | michael.buccini@sanofi.com | 1130 JIT Warehouse 8NYA | 2023-06-06 | 2023-06-12 | Per FBL-SOP-000661 material to be stored in proper storage conditions | Yes | Deviation
1732 | Phenyl Qualified Column does not appear in SAP genealogy report. | timothy.ryley@sanofi.com | 10Cal Rm130 | 2023-05-25 | 2023-06-12 | All materials consumed by Cerezyme 760845 would appear in the SAP report. |  | Deviation
1731 | 8NYA-Cerezyme-706210 Sample not taken in BioLIMS, no manual test request comment made in record | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-06-12 | 2023-06-12 | Determine if sample was taken by verifying credo log records and manual test request archive from QC at Gemba.
1730 | FZ Media Record will not progress during formulation | maria.palmar@sanofi.com | 8 New York Ave. 8RM1350 | 2023-06-11 | 2023-06-11 | The record will move forward when the condition is satisfied.
1729 | 8NYA-Cz-Excursion during sampling operations of post-thaw activities, Critical Surface Monitoring | john.lytle@sanofi.com | Rm 1343 | 2023-06-11 | 2023-06-11 | Per FBL-EWI-000030, EM Sampling takes place prior to conclusion of operations and removal of equipment
1728 | Agitator not working on SP pooling SUM | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-06-11 | 2023-06-11 | Agitator would start once recipe was lauched
1727 | 8NYA-Gemba Part in BOM not needed | Tanner.Raiburn@sanofi.com | 8NYA-1350 | 2023-06-11 | 2023-06-11 | Path forward
1726 | 8NYA-Pompe-Discussion needed for undetermined state of autoclave cycle/ probes for production use | john.lytle@sanofi.com | Upstream CNC | 2023-06-10 | 2023-06-10 | successful completion of probe autoclave cycle
1725 | 8NYA SP Gemba- 807194 (0.5M Phosphoric Acid) Liquid Weighout UoM error | alejandro.martinez2@sanofi.com | 8NYA/8RM1350 | 2023-06-10 | 2023-06-10 | Fix UoM discrepancy
1724 | Aber probe install connection pull tab tear | Jeffrey.boldizar@sanofi.com | 8 NYA 1403 | 2023-06-10 | 2023-06-10 | Expect connection to work properly without tear of pull tab
1723 | 8 NYA SP- 100L SUM Bag leak detected during WFI fill | alejandro.martinez2@sanofi.com | 8NYA/8RM1350 | 2023-06-10 | 2023-06-10 | Replace SUM bag
1722 | 8NYA - Cerezyme - DGHBN005 - Evidence of Leak on Harvest Bag Discovered During Breakdown | michael.knapp@sanofi.com | 8NYA 1403 | 2023-06-09 | 2023-06-10 | Harvest SUM operates for duration of campaign with no breach or perceived breach.
1721 | 8NYA - Pompe - DGHAF002 - Fluctuating %DO Post Inoculation | michael.knapp@sanofi.com | 8NYA 1403 | 2023-06-09 | 2023-06-09 | %DO is able to maintain setpoint.
1720 | 705673-initial pH test result out of spec | leo.gillespie@sanofi.com | Solution Prep | 2023-06-08 | 2023-06-08 | per FBL-EWI-000009,, SD-FRM-FRAD23A0180-MAY23
1719 | 8NYA- GEMBA- Totes not available to transfer into | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-06-09 | 2023-06-09 | Scaled media totes not available to transfer material into
1718 | 76NYA-HAFL-room 355-product impact unknown-Balance F-14980 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-355 | 2023-06-07 | 2023-06-09 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
1717 | Records signed out for Turkiye Inspection not returned by due date | Amy.McIntyre@sanofi.com | 68 New York Avenue, room 1725 | 2023-05-27 | 2023-06-08 | Expected results are (Per FBL-SOP-000210, steps 6.2.7, 6.2.9 and 6.2.10): the records are returned to RM custody by the due date and that RM would inform the requestor that the signed out records need to be returned and the due date for the return. | Yes | Deviation
1716 | 8NYA, CZZ, 3 separate results generated for same POU associated with FRAD23E0409 | craig.stevenson@sanofi.com | 8NYA, tested 68NYA/4700/QC Micro | 2023-06-02 | 2023-06-08 | 1 set of results is generated for each process stage, OneLims investigation to be opened to flag sample, Phenix record to speak to sample impact
1715 | Lost page Manufacturing Walkthrough Checklist | valeria.moreno@sanofi.com | 76 NYA | 2023-06-07 | 2023-06-07 | FBL-SOP-000709 Page 5/5
1714 | CZ AA Oligos Calculation Error for Reagent | Kaylyn.Manoogian@sanofi.com | 68NYA QC Chemistry Lab | 2023-06-06 | 2023-06-08 | Calculation Error for 20mg/mL 2-Picoline-Borane (2-PB). Concentration of reagent would not change despite difference of 1uL of ACN.
1713 | Fire Alarm went off at 76NYA | justin.carpenter@sanofi.com | 76 NYA | 2023-06-08 | 2023-06-08 | Fire Alarm wont go off on a false alarm
1712 | 8NYA - Pompe Buffer record 805002 missing part 808713 in order to transfer. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-06-07 | 2023-06-07 | Hold Gemba to see path forward, order part from MLCs, Transfer.
1711 | Dynamic EM submitted without original Climet printout data | megan.botti-marino@sanofi.com | 8 NYA | 2023-05-21 | 2023-05-31 | Per FBL-SOP-000942, step 6.8.4 the original thermal printout must be kept with the record.
1710 | GEMBA to discuss Harvest volume inadvertently transferred to the waste tank during Clarification DGF2B08A process | Jason.Brodrick@sanofi.com | 51 NYA ROOM 139 | 2023-06-07 | 2023-06-07 | It is expected that the calculated harvest volume be transferred to the T-2911 tank during concentration per FBL-MBR-000035.
1709 | # 80-Coldroom 139 Compressor Change out | justin.carpenter@sanofi.com | 2-8C Cold Room at 139 (INFOR asset # 80-Coldroom 139) | 2021-12-20 | 2023-05-05 | It is expected that for equipment, any change with possible GMP impact (per BQ-GOP-000175) should be driven and monitored using a CCR
1708 | Incorrect FIT recipe loaded for CZ CM | steven.erban2@sanofi.com | 8NYA Downstream | 2023-05-22 | 2023-06-07 | It is expected that MES would auto generate FIT recipe and if not done the correct recipe would be manually created.
1707 | Endotoxin sample tested at higher dilution than the minimum | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2023-06-05 | 2023-06-05 | Per FBL-SOP-000417 the sample is tested at 1:40 dilution
1706 | 8NYA - Pompe Buffer 805004 RF1N1D001 is calling for 200L bag and adaptor. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-06-07 | 2023-06-07 | Hold Gemba to see path forward.
1705 | 8NYA GEMBA - Osmo out of range - POMPE | alex.tello@sanofi.com | 8NYA RM1350 | 2023-06-06 | 2023-06-06 | Sample in range 370-390
1704 | 8NYA-Pompe-DGHAF002-Growth Media Expiry Error | michael.knapp@sanofi.com | 8NYA 1403 | 2023-06-06 | 2023-06-06 | Media expiries are correct |  | Deviation
1703 | Cerezyme & Fabrazyme - Lots TBD - SUM Code Parameter Errors | sheila.pandolfo@sanofi.com | 8 NYA | 2020-06-01 | 2023-06-02 | It is expected that the automation code has the correct values
1702 | Temperature excursion for QCTS freezer 40-QC-0597 | jacqueline.legere@sanofi.com | 68NYA Lab 3700 | 2023-06-05 | 2023-06-05 | Per FBL-SOP-001217, the temperature must be between -50C and -90C.
1701 | 8NYA GEMBA - Incorrect Osmo result | alex.tello@sanofi.com | 8NYA RM1350 | 2023-06-05 | 2023-06-05 | Path forward
1700 | 68NYA Room 4700 Dept QCM68-TOC F-19932 was found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 68-4700 | 2023-06-01 | 2023-06-05 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
1699 | 8NYA - Cerezyme - DGHBN005 - Eluate 6 Collected in two Eluate Bags | michael.knapp@sanofi.com | 8NYA 1403 | 2023-06-04 | 2023-06-05 | SUM eluate bag is able to receive all 21 CZ elutions.
1698 | SUB end of run prompt answered early. | Brendan.Dickhaut@sanofi.com | 8NYA/Room 1403 | 2023-06-05 | 2023-06-05 | FBL-EWI-000011 end of run samples required per sample plan in FBL-RPT-000538
1697 | Evacuation step not passing results on the Ross (post Vital) | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1703 | 2023-06-04 | 2023-06-04 | Be able to continue process
1696 | HBOD waste leak on FZ 3K tote in CNC | Devin.Parks@Sanofi.com | CNC space | 2023-06-04 | 2023-06-04 | 3K HBOD waste tote should not have been leaking.
1695 | 8NYA CZ NANO Recipe did not fill bag 2 | christina.burkhart@sanofi.com | 8NYA, RM 1703 | 2023-06-04 | 2023-06-04 | It's expected that during the post viral shipping bag transfer recipe will fill both bags to target amount
1694 | Leakage on CNC 3K HBOD Waste Tote | marcelo.dossantos2@sanofi.com | 8NYA/1210 | 2023-06-04 | 2023-06-04 | HBOD waste bag should have not leaked.
1693 | MMC Pooling SUM bag leak | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2023-06-04 | 2023-06-04 | Path forward to maintain product integrity
1692 | Gemba – 8NYA – Media component not available in MES | william.lee@sanofi.com | 8NYA/1350 | 2023-06-03 | 2023-06-03 | 828159 media should be reversed in MES.
1691 | Nano filtration, Equilibration Buffer not reaching pressure | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1503 | 2023-06-02 | 2023-06-02 | FBL-EWI-000029
1690 | BHM-1742001, diaphragm on the pressure sensor is damaged. It has several dents in it and its not responding to applied pressure. | Sarah.flattich@sanofi.com | 8 NYA CNC location | 2023-06-02 | 2023-06-02 | Passing calibration.
1689 | Dynamic Climet data not parsing to OneLIMS | adrianna.wooden@sanofi.com | 8 NYA, 51 NYA | 2023-05-28 | 2023-05-28 | Data would parse from climet to onelims per FBL-SOP-000207 or FBL-SOP-001039.
1688 | Incubation information not recorded at time of incubation for HA fill samples. | sarah.adams@sanofi.com | 68 NYA | 2023-05-31 | 2023-06-01 | Incubation time and date recorded
1687 | DBYFRM020 Missing Witness Signatures | nicholas.macdonald@sanofi.com | 76-364 | 2023-06-15 | 2023-06-15 | It is expected that witness steps be signed by both operator and witness at the time of execution.
1686 | The cap for sample ID 708430 was not tight upon sample transport at 8NYA. Sample solution within the saple bag leaked out | michael.montiverdi@sanofi.com | 8NYA | 2023-06-01 | 2023-06-01 | Samples should be secure for sample transport.
1685 | CZ2 CM Eluate pre-filtration Bioburden discovered to be damaged at pick up | stanley.durand@sanofi.com | 8NYA Room 1410 | 2023-06-01 | 2023-06-01 | Sample bags should remain intact and not damaged
1684 | Rapid temp rise for cold room 138 | alyson.belleville@sanofi.com | 51-138 | 2023-06-01 | 2023-06-01 | it is expected that the cold room would remain within spec of 6-10C.
1683 | Cz CM Eluate Filtration filter Connected Backwards | robert.morawski@sanofi.com | 8NYA RM 1503 | 2023-06-01 | 2023-06-01 | Eluate filtration setup executed correctly.
1682 | Clover mite found on secondary packaging material of Seprafilm | alan.rousseau@sanofi.com | room 1017 | 2023-06-01 | 2023-06-01 | FBS-SOP-000432
1681 | 8NYA GEMBA: MCC Skid X-2811 BT-2811001 failed visual inspection during CIP swab sample | lolav.ahmed@sanofi.com | 8 NYA room 1403 | 2023-05-31 | 2023-05-31 | FBL-FRM-000776, FBL-SOP-000397
1680 | During the reaction of Lot DBYCMC090 (part 2139-32), Allen Bradley Panelview 1000 froze causing the shutdown of all linked components. | jan.stepanik@sanofi.com | 76 NYA, room 229 | 2023-05-31 | 2023-05-31 | PLC should not have frozen causing the mixers and pumps to shutoff during the reaction.
1679 | Temperature for resuspension of lof DBYSEP048 failed to maintain temp during degauss. | jan.stepanik@sanofi.com | 76-237 | 2023-05-30 | 2023-05-30 | Temperature should have stayed
1678 | 23 of NAed records are missing from the records management records destruction list | Damian.Barbato@sanofi.com | 80 NYA records management office area, room 120 | 2023-05-17 | 2023-05-30 | It is expected the all the listed NAed records were signed and witnessed to be destroyed by RM and QA
1677 | 8NYA - Issues with PCS Alarms in MES | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-05-18 | 2023-05-29 | ensure data is integral.
1676 | 8NYA - Sample tested under incorrect PH standardization. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-05-26 | 2023-05-30 | Hold Gemba to see path forward.
1675 | 8NYA-Pompe PPQ-805015 Cystine/Tyrosine buffer has unusable part on BOM | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-05-30 | 2023-05-30 | Removal of 714020 from the BOM
1674 | Harvest Vessel lo lo weight alarm keeps occuring | Brendan.Dickhaut@sanofi.com | 8NYA/Room 1403 | 2023-05-29 | 2023-05-29 | Expected results is that the harvest SUM does not fall below the lo lo weight alarm of 20.0kg
1673 | Glacial Acetic Acid GZ2390-01 volume identified to be insufficient during Clarification harvest pool pH and conductivity adjustment for SP1 DGF3B07A chromatogram review from PI | Jason.Brodrick@sanofi.com | 51 NYA ROOM 139 | 2023-05-29 | 2023-05-29 | It is expected that the harvest pool adjusted within 24 hours per FBL-MBR-00064.
1672 | 8NYA Cerezyme Phenyl - Agitator failed to turn on for 15 min mix prior to eluate sampling. | stanley.durand@sanofi.com | 8NYA rm 1503 | 2023-05-28 | 2023-05-28 | It's expected for agitator to turn on at start of mixing recipe
1671 | DGHBN005/761217 - 8 NYA - Incorrect media tote line sealed during media disposal operations | wilfredo.lopez@sanofi.com | 8NYA CNC space | 2023-05-28 | 2023-05-28 | Per FBL-EWI-000011, CNC operations, Disconnect media tote, the tubing that is removed from the automatic pinch valve is to be welded at the 5' extension. that connects the outlet of the tote to the extension tubing leading to the bioreactor. It is expected that the operators seal the correct tube for the media to that is being removed.
1670 | 8NYA-Cz-Shared Media usage discussion for P/N 760720 and 760722 partial solutions, MES stock issues | john.lytle@sanofi.com | Rm 1434 | 2023-05-27 | 2023-05-27 | In order to retain and re-use media under the Shared Media program implementation, Mfg will need the remaining stock available in MES. Current MES logic allows only for emptying of media and removal of remaining stock
1669 | 8NYA- GEMBA- Mod Enzyme record not able to proceed | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-05-27 | 2023-05-27 | Record would accept/reject the scans as per design to prevent miscalculation of activity level
1668 | Bioreactor readings taken outside of the alotted window post refeed | David.Loik@sanofi.com | 51 NYA | 2023-01-29 | 2023-05-26 | Readings would be taken at times instructed
1667 | 8NYA-Cz-MCC X2560 is experiencing consistent holds for Buffer Flow EM. Looking for path forward with Automation/ ENG assistance | john.lytle@sanofi.com | Rm 1434 | 2023-05-26 | 2023-05-26 | Buffer flow controlled by skid Buffer EM should not experience consistent holds for flowrate |  | Deviation
1666 | Recipe when into hold due to filter differential pressure controller low-low- alarm EM-X3090-FLT Failure. | kenny.leland@sanofi.com | 8 NYA/1323 downstream ballroom | 2023-05-26 | 2023-05-26 | Recipe would not go into hold during equilibration. |  | No Deviation
1665 | • After referencing the MCC record for Skid X-2560 RF639D006 this morning, I counted that the buffer pump alarm has now occurred 11 times in the last 3 days. We will need to look into why this keeps occurring. Comment made for restarts overnight. • We performed an in process cal/adjustment on the pH probe for column 2 of X-2543 last night around 2230 due to column timer expiring. Range was looking for around 6.8-7.2 and we were not getting above 6.5. offline reading for the eqwil/wash was 7.02 which was what we adjusted to. The skid ran without further alarms but the meter itself had an alarm/error message on it that read “ch1Error pH offset too high”. Automation and MSAT consulted. MSAT requested that engineering be made aware of this error message this morning to determine if it has impact. Comments added to RF503D003. ITS-INC8899730. | Yves.Antoine@sanofi.com | Room 1333 | 2023-05-24 | 2023-05-24 | FBL-EWI-000032
1664 | MCC Buffer 768506 change out cannot able to execute due to initial set up incorrect. | Bhargav.Patel@sanofi.com | 8 NYA / 1210 | 2023-05-25 | 2023-05-25 | SOP instruction able to follow for FBL-EWI-000032
1663 | Spinners still in FBL-SOP-000692 | nicole.lufkin@sanofi.com | 51 NYA | 2023-05-25 | 2023-05-25 | spinners removed from SOP
1662 | GEMBA: 8NYA Zinc POU buffer sample tested in QCM under aborted MES record | judith.keller@sanofi.com | 68NYA/4700; 8NYA solution prep | 2023-05-20 | 2023-05-24 | Samples are expected to be labeled accurately and represent content in container | Yes | Deviation
1661 | 8 NYA GEMBA- Leak in load filter inlet during fz UF/DF | cory.lefort@sanofi.com | 8 NYA 1323 | 2023-05-25 | 2023-05-25 | It is expected that there is not leak in any of the SUT during the load of fz UF/DF.
1660 | Production recipe launched on day 11 for Fabrazyme run R503D003/Production started on day 11 for Fabrazyme | marcelo.dossantos2@sanofi.com | 8NYA/1333 | 2023-05-24 | 2023-05-24 | Production starts on day 12 for Fabrazyme run
1659 | Seprafilm – Multiple Lots – 76NYA – Gamma Irradiation Dosimeter Placement conditions were not met | fatima.fragaalvarez@sanofi.com | Steris | 2023-05-24 | 2023-05-24 | It is expected that all conditions for the receipt of shipments post-sterilization are met as described in FBS-SOP-000803 Step 6.22 page 9.
1658 | pH meter 2009226 was not re-standardized to the correct parameters for the 2 buffers. | Brendan.Dickhaut@sanofi.com | 8NYA, Room 1350 | 2023-05-24 | 2023-05-24 | Per FBL-EWI-000022 part numbers 760731 and 760729 buffer's pH are measured with a bracketed approach using the range of buffer standards measuring 4-7-10
1657 | Fz Nano Shipping Bags- Recipe failed to end at last bag | nathaniel.barnes@sanofi.com | 8NYA 1313 | 2023-05-24 | 2023-05-24 | Recipe will need to be pushed to completion of the final phase to finish Shipping Bag operations
1656 | HA Lot - DBYHAS002; Live carpenter ant found in 76 NYA Room 3102 (HA Purification) | munish.kumar2@sanofi.com | 76 NYA; Room 3102 (HA Purificaiton) | 2023-05-24 | 2023-05-24 | Refer to FBL-SOP-000512 steps 6.2.2 and 6.3.1.
1655 | blah | kenneth.morgan@sanofi.com | blah | 2023-05-24 | 2023-05-24 | blah
1654 | 8 NYA GEMBA- Cz Freeze bag underfill | cory.lefort@sanofi.com | 8 NYA 1703 | 2023-05-24 | 2023-05-24 | It is expected that during the freeze bag fill the RoSS would fill the freezer bag to a range of 7.2-7.7kgs. However, the SUM ran out of product before the shell could reach that range.
1653 | Failed RQ of Cold Room 88 in 51 NYA | megan.perkins@sanofi.com | 51 NYA / Room 88 | 2023-05-20 | 2023-05-22 | Per PQ 02-2326, Each individual distribution sensor reading (Ti) is within 6°C and 10°C during the entire 48-hour test period.
1652 | 51NYA-Facilities-Room 146-product impact unknown-Temperature indicator F-10964-014 on clean steam generator X-9130 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 51-146 | 2023-05-23 | 2023-05-24 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
1651 | 3000L SUM bag had a leak during initial WFI fill. | jermaine.kellam@sanofi.com | 8RM 1350 | 2023-05-23 | 2023-05-23 | Bag should not leak at all.
1650 | Discussing disabling pH probe 2 on SUB X-2220 | Jeffrey.boldizar@sanofi.com | 8 NYA 1403 | 2023-05-23 | 2023-05-23 | The PH probe to be in range
1649 | Particulate found in SUB sample during once per eluate sampling | Jeffrey.boldizar@sanofi.com | 8 NYA 1403 | 2023-05-22 | 2023-05-22 | No particulates | Yes | Deviation
1648 | Pompe – DGHTY001 – 8 NYA – FBL-RPT-001201 does not align with FBL-SPC-000272 and MES requirements | amanda.sorel@sanofi.com | 8 NYA | 2023-05-23 | 2023-05-23 | FBL-RPT-001201 would only have the pH requirement for this buffer. | Yes | Deviation
1647 | UFDF CIP Record | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-05-23 | 2023-05-23 | Record will terminate correctly
1646 | Time recorded for BSC plating start time is before time recorded for critical hood operations end time. | megan.botti-marino@sanofi.com | 45NYA | 2023-05-05 | 2023-05-23 | Per step 5.7.1 of RDSOP-003130, viable surface monitoring is to take place after the conclusion of BSC activities.
1645 | 8NYA-Pompe-No replacement N2 pads (CI#4944) not available for replacement per FBL-EWI-000030 | john.lytle@sanofi.com | 8NYA/ Rm1303 | 2023-05-23 | 2023-05-23 | N2 pads in stock for use in operations
1644 | DBYHAS002 Processed before release of Action Notice | nicholas.macdonald@sanofi.com | 76-3101, 76-3102 | 2023-05-22 | 2023-05-23 | It is expected that Action Notice release occur prior to forward processing.
1643 | During RQ study minimum temperatures reached 5.17°C, acceptance criteria range is 6°C - 10°C | ahmedgaffar.syed@sanofi.com | 51 NYA/Cold Room 87 | 2023-05-19 | 2023-05-19 | RQ study failed to meet the requirements since 3 (TC 6,7,8) out of 20 thermocouples (TCs) failed to operate within the acceptable range of 6 to 10 degrees Celsius. |  | No Deviation
1642 | E4 hose PM, unable to find replacement log from 2022. Hoses are supposed to be replaced annually, and inspected every 6 months. However, PMs are only listed every 12 months in Infor. | adam.rowe@sanofi.com | 51NYA,rm140 | 2022-04-30 | 2023-05-22 | Per SOP hoses to be replaced annually and inspected every 6 months.
1641 | HA Lot - DBYHAS002; Calculation Error and incorrect WFI addition amount | charlie.houle@sanofi.com | 76 NYA; Room 3101 (HA Pre-Purificaiton) | 2023-05-19 | 2023-05-22 | Per FBS-MBR-000060, v.11.0, step 7.11.6, we are to add the calculated amount of WFI or 1200 kg (whichever is larger). The calculated amount recorded was 1132.0 kg and 1134.0 was recorded as added.
1640 | 8NYA Cerezyme - CM eluate filter clogged. | stanley.durand@sanofi.com | 8NYA ROOM 1503 | 2023-05-22 | 2023-05-22 | Filter is supposed to not clog during eluate filtration | Yes | Deviation
1639 | 8NYA-Cz-MCC X2731 Bubble trap wetted after Buffer tote emptied, Bubble trap prime attempted | john.lytle@sanofi.com | Rm 1434/ 8NYA | 2023-05-22 | 2023-05-22 | Buffer swit chusing the shared media update will change the buffer at 50kg remaining. Buffer 760720 was found empty but was still reading 247kg. The incorrect weight led to the buffer running out
1638 | 2 different Climet 450t flow rates in FBL-SOP-001039 | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2023-05-22 | 2023-05-22 | It is expected that the flow rates align and be 50LPM, and has a +/-5% specification (45.5-52.5LPM)
1637 | 8NYA-Pompe- no HU label available in MES for Flask operations | john.lytle@sanofi.com | 8NYA- Rm1303 | 2023-05-21 | 2023-05-21 | Batch record should include an automated step to create a HU/ label information associated with flask operations
1636 | 8NYA POU samples Signed for during Formulation of Buffer | thomas.loik@sanofi.com | 8RM1350 | 2023-05-19 | 2023-05-21 | Correct samples in BioLims
1635 | 76-AHU 4A Fan failure (Common are) | john.beaule@sanofi.com | 76-AHU 4A Common area in-suite | 2023-05-21 | 2023-05-21 | FBS-SOP-000747, area pressures
1634 | Low recovery of Formulated Bulk CGF4B15 | alyson.belleville@sanofi.com | 51 NYA RM 87 | 2023-05-21 | 2023-05-21 | It is expected that the recovery for the Formulated Bulk operation would be within the range of 93-104%.
1633 | 8NYA Cerezyme - Temp probe on Nano load SUM bag is damaged | stanley.durand@sanofi.com | 8NYA ROOM 1503 | 2023-05-22 | 2023-05-22 | SUM bag should fully functional and not damaged.
1632 | Gemba – 8NYA – Pump head from supplier built incorrectly | william.lee@sanofi.com | 8NYA/1350 | 2023-05-19 | 2023-05-19 | Pump head assembly 713975 should be built correctly
1631 | Gemba – 8NYA – HVAC and EMS alarms | william.lee@sanofi.com | 8NYA/ 1313/ 1323 | 2023-05-21 | 2023-05-21 | HVAC units should be working properly.
1630 | Media POU samples for CZ taken, but not submitted. | Devin.Parks@Sanofi.com | 1403 | 2023-05-17 | 2023-05-19 | Samples should have been submitted right after taking them. | No | No Deviation
1629 | bioreactor X2220 and X2230 went into hold for perfusion failures | brendan.dickhaut@sanofi.com | 8NYA, Ballroom 1403 | 2023-05-19 | 2023-05-19 | Clamps would be removed from media supply line during change out.
1628 | ANMS and MES service interuption | Shiful.Kohli@sanofi.com | site wide (ANMS), 8NYA (MES) | 2023-05-18 | 2023-05-18 | TBD | Yes | Deviation
1627 | Abnormal UV peak discovered during SP1 Wash 1 Chromatography Review | cindy.guerra@sanofi.com | 51NYA | 2023-05-06 | 2023-05-18 | Initial decrease to baseline, then two distinct peaks |  | Deviation
1626 | 8NYA Bioburden sample CV UFDF X307001 CIP leaked | cleiton.lopes@sanofi.com | 68 QC Micro | 2023-05-18 | 2023-05-18 | Leaked and cracked samples will not proceed to testing |  | No Deviation
1625 | Thyrogen - MES does not align with step 6.4.16 of FBL-SOP-000496 for confirmatory reading for 7/10/12.45 standards | Jeremy.Wooden@Sanofi.com | 45 NYA/room 190 | 2023-05-19 | 2023-05-19 | It is expected that MES would align with the SOP for the correct buffer standard to be used.
1624 | Data Integrity - Incorrect instrument recorded in Onelims during routine EM Collection | craig.stevenson@sanofi.com | 8NYA | 2023-05-17 | 2023-05-17 | It is expected that EM is performed with equipment that is active within OneLIMS. | Yes | Deviation
1623 | A SUM bag began to leak during initial WFI fill | christopher.arnold@sanofi.com | 8RM1350 | 2023-05-18 | 2023-05-18 | That all SUM bags will not leak.
1622 | Pompe - RF1N9D002 - 8NYA - Wave Equil Started Prior to Postpone of Thaw | michael.knapp@sanofi.com | 8NYA 1403 | 2023-05-18 | 2023-05-18 | Wave media equil occurs on day prior to use. |  | No Deviation
1621 | LOT: DBYMES007 Labeled with Incorrect Expiration Date | rebecca.Cunningham@sanofi.com | 76NYA, 74NYA | 2023-03-28 | 2023-05-18 | All product is labeled with the correct expiration date. CofA, CofC and labels should have the same expiration date that is assigned to the product per FBS-SOP-000488
1620 | Pompe - RF1N9D002 - 8NYA - Wave Equil Started Prior to Postpone of Thaw | michael.knapp@sanofi.com | 8NYA 1403 | 2023-05-18 | 2023-05-18 | Wave media equil occurs on day prior to use.
1619 | 8NYA-Fz-DeltaV parameters for SUB Perfusion EM "Weight Controller High-High Limit(RPM)" setpoint inadequate to prevent continuous hold conditions after perfusion rate changes | john.lytle@sanofi.com | Rm1333 | 2023-05-18 | 2023-05-18 | Automation has encountered this problem in the past and increased the current setpoint to 65rpm, which will cover the rate changes experienced on the ATF pumps during perfusion
1618 | 8NYA-Cerezyme-MOD Enzyme MES/Athena BOM ZUMERRORS | thomas.loik@sanofi.com | 8RM1350 | 2023-05-01 | 2023-05-01 | Adjust Athena BOM
1617 | 51 NYA rhTSH Purification CV's Failing Conductivity for RQ | jacob.kallon@sanofi.com | 68 NYA, QC Micro Lab | 2023-05-17 | 2023-05-17 | It is expected that WFI rinse sample meets passing criteria per USP <645> |  | No Deviation
1616 | Inoculation timer exceeded on D1 | david.loik@sanofi.com | 51 NYA | 2023-05-17 | 2023-05-17 | Inoculation would be completed within the 7.5 hour allowable timeframe.
1615 | An amber clamp on a buffer tote discharge line was not completely shut, allowing buffer to flow down the line and wet out the aseptic connection seal. | richard.manasian@sanofi.com | 8RM1202, the CNC space outside of Solution Prep | 2023-05-18 | 2023-05-18 | That the setup would lead to a tote that does not leak.
1614 | DGHAW053/760741 Media with DTT - 8NYA CNC - Contaminant observed during point-of-use sampling | wilfredo.lopez@sanofi.com | 8NYA CNC space | 2023-05-17 | 2023-05-17 | It is expected that the material be free of contaminants prior to use as it has been filtered post production.
1613 | Particulates assay did not use proper number of vials for testing | aimee.park@sanofi.com | 45 New York Avenue | 2023-04-17 | 2023-05-17 | The analyst would have used the appropriate number of vials for testing
1612 | 8NYA Cerezyme - CM eluate filter clogged. | stanley.durand@sanofi.com | 8NYA room 1503 | 2023-05-17 | 2023-05-17 | Filter is supposed to not clog during eluate filtration
1611 | 8NYA-GEMBA-Parts missing from BOM 805012 | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-05-17 | 2023-05-17 | All material should be included within the BOM
1610 | ATF tubing connection defect discovered during ATF 10 build | Jeffrey.boldizar@sanofi.com | 8 NYA 1403 | 2023-05-17 | 2023-05-17 | Expect connection to work properly
1609 | One incorrect roll of large illustration labels (808141) was issued along with 3(808148) rolls of small illustration labels to US PK lot DBYSEP024 p/n 787515 | alan.rousseau@sanofi.com | room 1017 | 2023-05-17 | 2023-05-17 | FBS=MBR-000090 FBS-SOP-000453
1608 | 8NYA-GEMBA-Wrong part listed in takeout 805012 | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-05-17 | 2023-05-17 | Correct part number should be listed on takeout step
1607 | No Activity level documented during routine EM on 13Jan2023 at 45NYA CTMS Suite 1 Room 115. | craig.stevenson@sanofi.com | 45NYA Room 115 | 2023-01-13 | 2023-05-17 | It is expected that EM samplers would document the activity level on the associated paperwork.
1606 | 8NYA-Cz-RFID for Train 3 and 4 tote parking spot are not working and Trian 4 valve closing while pulling buffer | john.lytle@sanofi.com | Rm1210 CNC Corridor Phase 2 Upstream | 2023-05-17 | 2023-05-17 | functional RFID and associated buffer valve logic functional
1605 | Filter stuck to the lid instead of R2A for bioburden sample cv 23MS00994 Permeate X3071 RFICBD007 | Stefanie.Erickson@sanofi.com | QC Micro counting | 2023-05-11 | 2023-05-17 | For the filter to be placed on the agar for the duration of incubation
1604 | Non-contemporaneous data entry for Fabrazyme ICB AA Oligos | John.Kapaniris@sanofi.com | QC Chemistry 2nd floor Lab 68NYA | 2023-05-03 | 2023-05-15 | Contemporaneous data entry is required
1603 | 68 NYA QCM GDP Error in Spor-Klenz Lot Documented for F-16088 Hood Clean and Manifold Sanitization on 02FEB2023 and 04FEB2023 | Grayson.guyton@sanofi.com | 68 New York Ave., 4700 , QC Micro | 2023-02-02 | 2023-05-16 | Per FBL-SOP-000594 step 6.5.3 it is expected that documentation is performed properly before cleaning can occur | Yes | Deviation
1602 | 45NYA-QCC45 Room 1136-product impact unknown- F-17365 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 45-1136 | 2023-04-20 | 2023-05-17 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results | Yes | Deviation
1601 | 45NYA-Dept SP45 Room 190-Filter Integrity Tester F-17119 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 45-190 | 2023-03-29 | 2023-05-17 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 unit should have produced passing results
1600 | Phenyl Skid recipe not moving forward | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1503 | 2023-05-16 | 2023-05-16 | to be able to resumed recipe
1599 | Lower sensor on bubble trap is showing liquid/ Bubble trap incorrect reading. Should be reading air. | marcelo.dossantos2@sanofi.com | 8NYA/1333 | 2023-05-16 | 2023-05-16 | Both sensors should be reading air coming off SIP.
1598 | Wave reactor Tare perform inadvertally | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2023-05-16 | 2023-05-16 | Should be tare prior to Media load | Yes | No Deviation
1597 | DBYFRM020: Fermenter Rupture Disc Ruptured | nicholas.macdonald@sanofi.com | 76-364 | 2023-05-15 | 2023-05-15 | It is expected that the pressure remain within the MBR range of 5 - 20psi
1596 | NBO-SPC-000259 obsoleted while associate part (Bioshield) still in use | nicholas.shoenfelt@sanofi.com | 51NYA | 2023-05-10 | 2023-05-10 | It is expected that SPCs will not be obsoleted until after a part has been discontinued from use.
1595 | 8NYA-Pompe PPQ-805015 Cystine/Tyrosine buffer has discrepancy between sample plan and MES sample requirements. | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-05-15 | 2023-05-16 | All documentation should align to the correct sampling plan to avoid confusion during batching.
1594 | TOC sample submitted without Swab - Thyrogen Hose Station RQ | nicholas.shoenfelt@sanofi.com | 68NYA QC Micro Lab | 2023-05-14 | 2023-05-15 | It is expected that the TOC swab is snapped off into the TOC vial after swabbing.
1593 | 8NYA-Cereyzyme-Mod Enzyme cocktail requesting manual inputs for volume calculations for final bottle | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-05-15 | 2023-05-15 | Determination at Gemba | Yes | No Deviation
1592 | Fabrazyme - DGHCE003 -8NYA - Vicell "Write Error" Resulting in Sample Data not Saving on Equipment | michael.knapp@sanofi.com | 8NYA 1333 | 2023-05-14 | 2023-05-14 | Sample data is saved properly and sent to PI.
1591 | Cerezyme - DGHBN004 - 8NYA - CO2 Probe 1 Disabled Due to Reading Loss | michael.knapp@sanofi.com | 8NYA 1403 | 2023-05-12 | 2023-05-13 | Probe readings are accurate and have redundancy.
1590 | Cz Bioburden sample spilled at sample receipt | diane.collins@sanofi.com | 68 NYA ,4720 sample receipt | 2023-05-12 | 2023-05-12 | Expected that bioburden samples are not opened during receipt
1589 | Incorrect Cleaning recipe launched for 8NYA Qual Skid | zachary.rickard@sanofi.com | 8NYA/1303 | 2023-05-12 | 2023-05-12 | Recipe should contain a CV
1588 | Late entry on MBR | alexandre.silva2@sanofi.com | 76 NYA (Purification). | 2023-02-10 | 2023-05-12 | It's expected that Steps performed at time of execution.
1587 | 8NYA- GEMBA- Tote bag Connections for 760725 wetted with residual liquid | richard.manasian@sanofi.com | CNC at 8 NYA | 2023-05-11 | 2023-05-12 | Glycol line would remain connected without breaking off during removal/attachment
1586 | 8NYA-Fz-No instruction for DO Process Calibration using EWI instructions for newer model of M400 Mettler Toledo Transmitters | john.lytle@sanofi.com | Rm1333 | 2023-05-12 | 2023-05-12 | EWI instructions needed to cover operation of the 2nd model of M400 available for use in Mfg
1585 | Fabrazyme - 8NYA - Nano Guard Filter Changeout SD instructions need clarification following Biolims implementation | michael.knapp@sanofi.com | 8NYA 1343 | 2023-05-12 | 2023-05-12 | Instructions align with current EM Procedures.
1584 | Incorrect connection on ATF B during ATF Setup | Yves.Antoine@sanofi.com | Room 1333 | 2023-05-11 | 2023-05-11 | FBL-EWI-000032
1583 | 51 NYA Missing time out entry for buffer movement out of cold room 108C | vincent.palumbo@sanofi.com | 51 NYA | 2023-05-03 | 2023-05-11 | It is expected that the time would be recorded on the log sheet
1582 | Suspected Failure of Agitator Steam Seal on Bioreactor D1 | james.letourneux@sanofi.com | 51 NYA RM 140 | 2023-05-11 | 2023-05-11 | The agitator mechanical seal should not allow condensate or steam to pass through it into the vessel. The vessel interior should be dry, or only have residual moisture following WFI soak and rinse, at this phase of reactor preparation (post-PM and agitator clearance check, but prior to reactor SIPs).
1581 | HA - DBYHM003 - 76NYA - Climet settings not correct and climets stopped prior to end of critical operation | craig.stevenson@sanofi.com | 76NYA, Room 355 | 2023-04-04 | 2023-05-11 | It is expected that the climets perform continuous moniitoring during the duration of the HA Fill. It is also expected that the climet be stopped at the end of the Fill and not early.
1580 | Sample 696172-DGHAW045-MEDIA_760741_ENDOTOXIN submitted in polypropylene | larissa.fairweather@sanofi.com | 68NYA/4700/QCM | 2023-05-11 | 2023-05-11 | Per FBL-SOP-000817 Step 6.1.3 Ensure all samples are received and diluted in polystyrene or depyrogenated glass tubes prior to sample analysis.
1579 | Missing time entry for signage posted and work execution of 2023-F-029 | vincent.palumbo@sanofi.com | 51 NYA | 2023-05-10 | 2023-05-11 | It is expected that the date and time would be recorded in these fields at the time of occurrence.
1578 | DO probes rebuilt on D1 Bioreactor but not recorded in prep record | adam.hildebrand@sanofi.com | 51 NYA RM-140 | 2023-05-11 | 2023-05-11 | It is expected that probe rebuild would be recorded contemporaneously.
1577 | Fz X2210 SUB Probes not scanning on MES due to expiry. | paola.tello@sanofi.com | 8NYA RM133 | 2023-05-11 | 2023-05-11 | N/A
1576 | Incorrect Carbohydrates results reported on COA for 8GF1B4V | aimee.park@sanofi.com | 45 New York Avenue | 2019-04-17 | 2023-05-10 | Correct Carbohydrates results would be reported on the COA for Lot: 8GF1B4V
1575 | Shipping barrels have a whole at the base | marc.latta@sanofi.com | 8nya/1703 | 2023-05-05 | 2023-05-05 | barrel should not have a whole |  | Deviation
1574 | Shipping barrel base has a whole | marc.latta@sanofi.com | 8nya/1703 | 2023-05-05 | 2023-05-05 | barrel should not have a whole
1573 | Site 219 in 76 NYA was tested for bioburden as WFI instead of RO/DI Purified | alexander.decoste@sanofi.com | 68 NYA, Room 4790 (bioburden room), QC Micro | 2023-05-04 | 2023-05-10 | Per FBL-SOP-000019 the sample should have been tested with 1 mL diluted with 10mL of phosphate buffer. | No | Deviation
1572 | 8NYA-Cerezyme-RF623D003/DGHAX003-Seed Media (760744) bag was found in 1220 cold room with open clamp to sample bag which was also resting on the floor. | aidan.mula@sanofi.com | 8NYA/RM1220 | 2023-05-10 | 2023-05-10 | Determination at Gemba
1571 | DGHAW045, 760741 XVF MEDIA WITH DTT Formulation Sample Labeling Error | wilfredo.lopez@sanofi.com | 8NYA Room 1350 | 2023-05-09 | 2023-05-10 | Per FBL-RPT-000538, Table 26, it is expected that at the time of batching 760741 Media Buffer, samples are taken for Bioburden and Endotoxin, however the labels printed and applied to the samples incorrectly reflect that the sampling for Point-Of-Use Bioburden and Endotoxin (per Table 8) was performed in the BioLIMS system. POU samples are expected to be taken at the time of use in SUB production.
1570 | 8NYA- GEMBA Request- Missing samples | William.Lee@sanofi.com | 1350 at 8 NYA | 2022-11-24 | 2023-04-28 | Samples would have been delivered to QC micro and recieved
1569 | HA Lot - DBYHAM005 | charlie.houle@sanofi.com | 76 NYA; Room 3102 (HA Purificaiton) | 2023-05-09 | 2023-05-09 | Compression unit should work as intended, compressing the filters.
1568 | LoLo pO2 alarm during harvest of D2 bioreactor | David.Loik@sanofi.com | 51 NYA | 2023-05-09 | 2023-05-09 | Recipe would proceed without alarm.
1567 | Cerezyme – DGHAN004 – 8 NYA – Initial WFI Addition Target Weight Exceeded 1% Tolerance for PN 760731 | amanda.sorel@sanofi.com | 1350 8 NYA | 2023-05-09 | 2023-05-09 | Final volume would be in target following QS | Yes | No Deviation
1566 | First line of BIRF for SP1 eluate DGF2B04B not filled out for initial storage | alyson.belleville@sanofi.com | 51 NYA RM 138 | 2023-05-01 | 2023-05-09 | It is expected that the information for initial storage of the eluate for storage location, room temperature, and time of storage would be filled out at the time of BIRF creation and eluate storage. | Yes | Deviation
1565 | 8NYA- GEMBA Request- WFI over fill 760731 RF613D004 | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-05-09 | 2023-05-09 | Final volume would be in target following QS
1564 | Missing site ID for one non-viable sample, HA Routine EM 76 NYA, 18APR23 | megan.botti-marino@sanofi.com | 76 NYA | 2023-04-18 | 2023-05-09 | Per FBL-SOP-000207 "Use of the Climet CI-1054", step 6.5.4, "Choose appropriate sample location by selectin [ID] from the main screen."
1563 | 8NYA Fabrazyme - Leak observed at the filter inlet at start of UFDF load. | stanley.durand@sanofi.com | 8 NYA Room 1323 | 2023-05-09 | 2023-05-09 | It's expected during UFDF that SUTs do not fail during operation |  | Deviation
1562 | 8NYA-Fz- Re-GEMBA-Low Viability of Fz SUB-Potential ViCell interference | john.lytle@sanofi.com | 8NYA/ Rm1333 | 2023-01-27 | 2023-01-27 | viability >/80%
1561 | Fz ICB 100 L SUB Daily Viability Excursion for CGHCE006 Re-Gemba | mark.lindblad@sanofi.com | 8 NYA Phase 1 Upstream | 2023-01-27 | 2023-01-27 | Per FBL-CST-000028, it is expected all 100 L SUB daily cell viabilities are greater than or equal to 80 % for all culture days
1560 | Lack of PRS Inventory Management in QC Chemistry Laboratory at 68NYA | Kaylyn.Manoogian@sanofi.com | 45NYA and 68NYA | 2023-05-05 | 2023-05-04 | Per FBL-SOP-000182, "Subtract aliquot of PRS for each assay occasion or use." "Document Inventory adjustments for a Primary Reference Standard in the comments section of FBL-FRM-000424 and FBL-FRM-000395. Notify Quality Assurance per BQ-GOP-000248 if a Quarterly Inventory Count discrepancy cannot be reconciled.
1559 | Cerezyme - DGHBN006 - 8NYA - Aber Probe Reading and Trend Discussion | michael.knapp@sanofi.com | 8NYA 1403 | 2023-05-08 | 2023-05-08 | Probe readings are accurate and have redundancy.
1558 | 8NYA / Cerezyme DGHBN004 Stock # 69456 Inadequate amount sampled | michael.aiello@sanofi.com | 8NYA / 68 QCM | 2023-05-07 | 2023-05-07 | It is expected that amount of samples submitted matches LIMS quantity
1557 | No 84 uS/cm conductivity standard on site | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-05-08 | 2023-05-08 | There should be a stock of the standard on site
1556 | Partner phase hold on MCC skid X-2560 | Brendan.Dickhaut@sanofi.com | 8NYA, Room 1403 | 2023-05-08 | 2023-05-08 | Per FBL-EWI-000032, SUT lines coming from the phnumatic valves BV-605 and BV-606 would be directed to the correct SUMs and SUMS would be positioned in their allocated parking spots.
1555 | rhTSH SP1 Lot DGF2B04C Mix Time (KPP) Discrepancy | Ryan.Geraghty@sanofi.com | 51NYA | 2023-05-06 | 2023-05-08 | It is expected that the start and end time for mixing equals 5 to 10 minutes per FBL-MBR-000064, Step 4.3.11
1554 | rhTSH Clarification Lot DGF2B04C used incorrect ELISA results for calculations | Ryan.Geraghty@sanofi.com | 51NYA | 2023-05-04 | 2023-05-07 | It is expected the ELISA results used for calculations are the correct results for each harvest day.
1553 | Upon inspection of 3000L SUM bag it was found to not have proper connecting tubing for the inflation process | william.lee@sanofi.com | 8RM1350 | 2023-05-07 | 2023-05-07 | That the SUM bag will have all parts necessary to perform installation and inflation tasks.
1552 | HU for Material GZ6138-02 (lot: 22GN00387) not in Athena for consumption | nicholas.shoenfelt@sanofi.com | 51-124 | 2023-05-06 | 2023-05-06 | It is expected that the HU would be available to be consumed to the process order.
1551 | 8 NYA GEMBA- Fz Zinc elution collection line inadvertently sealed prior to elution being done | cory.lefort@sanofi.com | 8 NYA 1323 | 2023-05-06 | 2023-05-06 | Eluate collection to fully collect before line being sealed. | No | Deviation
1550 | GEMBA for leak from pump insert during ATF wetting | Brandon.Steff@sanofi.com | 8RM1403 | 2023-05-06 | 2023-05-06 | Pump inserts are used to prevent pump tubing from tearing
1549 | 8 NYA GEMBA- Cz RoSS Freeze bag underfill | cory.lefort@sanofi.com | 8 NYA 1703 | 2023-05-05 | 2023-05-05 | It was expected that the freeze bag would fill to a volume of around 7.7 kgs | Yes | Deviation
1548 | Media EM alarm for bioreactor X-2220 received via log-mate. Media was not being fed into bioreactors 2220/2230. | Brendan.Dickhaut@sanofi.com | 8NYA CNC logistics corridor 1210 and ballroom 1403 | 2023-05-05 | 2023-05-05 | Per FBL-EWI-000011 media changeout, operator is to ensure tote is reading in delta V. Tote was not receiving signal in delta V due to Profibus cable not being properly attached.
1547 | 80 NYA freezer keys logbook used with old revsion | vincent.palumbo@sanofi.com | 51 NYA | 2023-05-04 | 2023-05-05 | It is expected that the most current revision be used.
1546 | 3 HU labels were incorrectly applied at 8NYA for material 692169 | michael.montiverdi@sanofi.com | JIT 8NYA | 2023-05-09 | 2023-05-09 | The correct HU labels are applied to the correct material batch number.
1545 | 80NYA Cold Room 139 Compressor replacement | David.zhong@sanofi.com | 80NYA Cold Room 139 | 2021-12-20 | 2023-05-04 | FBL-SOP-000910 section 6.5.3 and per FBL-SOP-001460 section 6.14.1. |  | Deviation
1544 | DGHAW040 760741 XVF Media Formulation Sample Labeling Error | wilfredo.lopez@sanofi.com | 8NYA Room 1350 | 2023-05-04 | 2023-05-04 | Per FBL-RPT-000538, Table 26, it is expected that at the time of batching 760741 Media Buffer, samples are taken for Bioburden and Endotoxin, however the labels printed and applied to the samples incorrectly reflect that the sampling for Point-Of-Use Bioburden and Endotoxin (per Table 8) was performed in the BioLIMS system. POU samples are expected to be taken at the time of use in SUB production.
1543 | Percent Recovery of the rhTSH Formulation FB CGF1B14 is identified to be out of range during record review. | sylvester.kegbeh@sanofi.com | 51 NYA ROOMS 88 and 87 | 2023-05-04 | 2023-05-04 | It is expected that Percent Recovery be between 93%-104% per FBL-MBR-00114 step 4.6.3
1542 | 8 NYA Cerezyme Nano Formulation Nano load SUM ran dry during buffer chase. | stanley.durand@sanofi.com | 8NYA ROOM 1503 | 2023-05-04 | 2023-05-04 | Pump alarm should be cleared prior to running.
1541 | FBL-MTD-000039 Phosphate determination performed on Lambda 25 in error | john.kapaniris@sanofi.com | QC Chemistry 45NYA | 2023-05-03 | 2023-05-03 | Per CAPA FRAD20A0061, All assays for PHosphate determination need to be performed on the Lambda 35 Spectrophotometers
1540 | Hyaluronic Acid - DBYHAM004 - 76NYA - Multiple errors in HA Fill packet | craig.stevenson@sanofi.com | 76NYA, Room 355 | 2023-05-03 | 2023-05-04 | It is expected that the EM sampler follows FBS-SOP-000607
1539 | Cerezyme - DGHBN005 - 8NYA - Capacitance Probe Divergence Leading to Potential Overbleeding | michael.knapp@sanofi.com | 8NYA 1403 | 2023-05-04 | 2023-05-04 | Probes remain consistent with each other and do not deviate or result in inaccurate bleed rates. |  | No Deviation
1538 | 8NYA-PF8-Room 1503-Delta V weight indicator FB-4312-007 had no as found data taken before replacement of transmitter | robert.bauch@sanofi.com | 8-1503 | 2023-05-03 | 2023-05-03 | BQ-GOP-000046 Calibration Program Version 6 Section 5.8.3 as found data should have been taken.
1537 | 8NYA GEMBA - SUM bag has incorrect recirculation connection | alex.tello@sanofi.com | 8NYA RM1350 | 2023-05-04 | 2023-05-04 | SUM bag has male connection should be female.
1536 | Cerezyme -DGHBN005-8NYA-Leak observed on media tote during installation | michael.knapp@sanofi.com | 8NYA 1403 | 2023-05-03 | 2023-05-03 | No evidence of potential leaks on media totes. |  | 556 Event
1535 | ER FRAD21R0012 deemed to be ineffective per BQ-GOP-000249 causing a new deviation event to be created in accordance with BQ-GOP-000248 | paul.mclaughlinjr@sanofi.com | 68 NYA | 2023-05-03 | 2023-05-03 | Not-Effective; BQ-GOP-000249 step 5.25.6
1534 | Non-viable data for 51NYA did not parse correctly in biolims | megan.botti-marino@sanofi.com | 51 NYA | 2023-04-11 | 2023-05-03 | It is expected that data parse correctly in biolims
1533 | Pressure gauge F-7611 in use while de-activated in Gage InSite | jacob.kallon@sanofi.com | 51 NYA room 139 | 2023-04-27 | 2023-04-27 | FBL-SOP-000943
1532 | Incorrect Media quantity distribution between records causing tote to have a remainder not charged to each SFO | Jeffrey.boldizar@sanofi.com | 8 NYA 1403 | 2023-05-02 | 2023-05-03 | expected that the record would distribute medias full quantity to each SFO
1531 | pH adjust not performed during media prep at 51NYA | nicholas.shoenfelt@sanofi.com | 51-124 | 2023-05-03 | 2023-05-03 | It is expected that a pH adjustment will be made if the initial pH is out of range.
1530 | 8NYA Cz Post viral SUM weight issue | zachary.rickard@sanofi.com | 8NYA/1703 | 2023-05-03 | 2023-05-03 | SUM will be at a steady weight and not changing while it is not being touched
1529 | Two QC Micro Balances (FB-3513 and F-21321) failed annual calibration | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2023-05-03 | 2023-05-03 | It is expected that balances generate passing As Found data.
1528 | Cerezyme - DGHBN006 -8NYA - Bleed Pump Mod Bad Alarms | michael.knapp@sanofi.com | 8NYA 1403 | 2023-05-03 | 2023-05-03 | No Mod Bad alarms during operations.
1527 | Approved version of FBL-FRM-000797 used for 8 NYA dynamic EM and surface monitoring | megan.botti-marino@sanofi.com | 8 NYA, room 1343 | 2023-03-26 | 2023-05-03 | It is expected that effective versions of documents be used to execute procedures.
1526 | 51NYA 2-8 C Refrigerator Asset #2050814 in Room 140 major repair performed without CCR | David.zhong@sanofi.com | 51NYA, Room 140, Asset# 2050814 | 2023-03-17 | 2023-05-02 | FBL-SOP-000910, When changes affect: • The validated state of a system or asset A Change Control or other Quality Change Management system may be required; FBL-SOP-001460, Major changes may require initiation or revision of engineering and validation lifecycle elements. |  | Deviation
1525 | 45 NYA - Thyrogen - doors between 189B and MSAT pilot lab blew open. | Jeremy.Wooden@Sanofi.com | 45 NYA, room 189 B | 2023-05-02 | 2023-05-02 | That the door would remain closed. | Yes | No Deviation
1524 | CTMS 45NYA WO 2859801 and WO 2859798 Referenced the Incorrect Deviation # in INFOR and Did Not Have A Deviation Opened for Late PM last Nov2022 | tanni93@bu.edu | 45-113 and 45-114 | 2022-12-01 | 2023-05-02 | FBL-SOP-000081 6.4.9
1523 | SoloVPE not parsing data to OneLIMS | marygrace.holmes@sanofi.com | 68NYA, QC Chemistry | 2023-05-02 | 2023-05-02 | SoloVPE data would would parse to correct location
1522 | TSA Single Wrap Media (Lot 23GN00639) received into lab with open packaging | amanda.lavalley@sanofi.com | QC Micro Lab 68 NYA | 2023-04-27 | 2023-05-01 | Single wrap TSA should have fully closed packaging
1521 | Change Control 2021FRACC0014 Cold Rooms 1220, 1131 Routine Use Works Open while Cold Rooms 1220, 1131 in GMP Use | stephanie.aceto@sanofi.com | 8 NYA Rooms 1220 and 1131 | 2023-05-02 | 2023-05-02 | Release for use form would have been completed as described per works 18 and 19. |  | No Deviation
1520 | HA Lot - DBYHAM004 - Mixing speed of HA-CTAB solution out of specification | Charlie.Houle@sanofi.com | 76 NYA; Room 3101 (HA Pre-Purificaiton) | 2023-04-20 | 2023-05-02 | Per FBS-MBR-000060, v.10.0, step 7.9.6 requires a target mixing speed of 120.0 rpm with a range of 70.0 - 120.0
1519 | 8 NYA GEMBA- FZ SP Leak observed in the Load SUM during product transfer | cory.lefort@sanofi.com | 8 NYA 1323 | 2023-05-02 | 2023-05-02 | It is expected that the Fz SP product would transfer into the Load SUM bag without any leaks.
1518 | Thyrogen - Recirculation end time recorded as prior to start time for Clarification DGF3B03C | ryan.geraghty@sanofi.com | 51NYA | 2023-04-17 | 2023-05-02 | It is expected that the correct times are recorded and total time is within parameters per FBL-MBR-000035 | Yes | Deviation
1517 | Expired reagent used in execution of FBL-MTD-000074 | nancy.watkins@sanofi.com | 68NYA, QC Chemistry Lab | 2022-10-21 | 2023-05-01 | FBL-MTD-000074 step 6.1.6 Label and store at room temperature for up to 1 month
1516 | Approved version of FBL-SOP-000755 used twice for 76NYA, once for dynamic EM, once for an aseptic qualification | megan.botti-marino@sanofi.com | 76 NYA, Hood F-0704 | 2023-05-04 | 2023-05-01 | It is expected that effective versions of documents be used to execute procedures.
1515 | 68NYA-QCM-Product impact unknown-TOC FB-2426 found out of tolerance for conductivity during a scheduled calibration | robert.bauch@sanofi.com | 68-4700 | 2023-04-26 | 2023-05-01 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results | Yes | Deviation
1514 | D2 Inoculation Sampling Impossible under AN for Corridor 83 Ceiling Breach at 51NYA | kevin.devanna@sanofi.com | 51 NYA Rm 140 | 2023-05-01 | 2023-05-01 | Ceiling tiles should not be opened without proper planning. Bioreactor should be sampled for BGA readings, and Mycoplasma allocated in the BSC, within 15 minutes of inoculation completion.
1513 | LAL samples for 51CV-0014 and 51CV-0016 were submitted in polypropylene tubes. | roodline.bertrand@sanofi.com | 51 NYA | 2023-04-30 | 2023-05-01 | Per FBL-SOP-0000960 step 6.4.3 "Ensure all samples are received and diluted in polystyrene or depyrogenated glass tubes prior to sample analysis."
1512 | Item missing from Cz Nanofiltration BOM (691804, 200L Waste bag) | jarad.kukla@sanofi.com | 8NYA / 1503 | 2023-04-29 | 2023-04-29 | It is expected that the 691804 is included on the BOM as is it needed to collect equilibration waste during the nanofiltration operations.
1511 | 51NYA-APT DS-Room 131-Pressure Gauge FB-1207 is overdue for calibration | robert.bauch@sanofi.com | 51-131 | 2023-05-01 | 2023-05-01 | BQ-GOP-000046 Calibration Program Version 6 Section 5.7.5.5, Portable GMP Critical instruments shall be received into Metrology before its due date.
1510 | Weight Set FB-4589 Brought Down to Metrology Past due date | kyle.lebherz@sanofi.com | QC Micro | 2023-05-01 | 2023-05-01 | For instrument to be brought down before calibration due date
1509 | Cz Shipping Bag transfer recipe won't complete | robert.morawski@sanofi.com | 8NYA RM1703 | 2023-04-30 | 2023-04-30 | Recipe looks to fill bag off of remaining SUM bag volume and not shipping bag volume. |  | Deviation
1508 | Cerezyme - DGHBN006 - 8NYA - Antifoam pump tubing leak during antifoam prime | michael.knapp@sanofi.com | 8NYA 1403 | 2023-04-29 | 2023-04-29 | No leaks on SUT
1507 | 8NYA GEMBA- Crack on multiple air filters P/N 717325 | Alex.tello@sanofi.com | 8NYA RM1350 | 2023-04-30 | 2023-04-30 | SUM bag should be completely inflated.
1506 | Cerezyme - DGHBN006 - 8NYA - Media Valve Not Opening during Perfusion Start | michael.knapp@sanofi.com | 8NYA 1403 | 2023-04-29 | 2023-04-29 | Correct media valve opens and perfusion is initiated
1505 | Cerezyme - DGHBN006 - 8NYA - High levels of foam observed prior to perfusion start | michael.knapp@sanofi.com | 8NYA 1403 | 2023-04-29 | 2023-04-29 | SUB system does not generate foam which could potentially wet vent filters prior to perfusion start and routine antifoam dosing.
1504 | F-16088 BSC usage log not completed for assay occasion ALF018RM-000000006 | alexander.decoste@sanofi.com | 68 NYA, Room 4790 (bioburden room), QC Micro | 2023-04-28 | 2023-04-28 | FBL-SOP-000594 V.8.0 Step 6.5.3 Document pre- and post-use cleaning/disinfection on FBL-LBK-000202. Perform in order of operations. and Step 6.3.4 Ensure magnehelic gauge is within the acceptable range for the BSC and document on FBL-LBK-000202.
1503 | 76 NYA PUR DBYHAM004 “Chart recorder was set for 8hrs instead of 24hrs ". | munish.kumar2@sanofi.com | 76 NYA/Room#3102/Purification | 2023-04-28 | 2023-04-28 | FBL-SOP-000458 Step# 6.8.20
1502 | FBL-MBR-0000048 obsoleted without appropriate quality driver | gregory.klotz@sanofi.com | 51 NYA | TBD | 2023-04-28 | Expected that MBR obsoletion would be performed as part of appropriate quality driver (CCR) | Yes | Deviation
1501 | EMS alarm forcing staff to leave during hood process | Jeffrey.boldizar@sanofi.com | 8NYA 1343 | 2023-04-28 | 2023-04-28 | room remains integral during operation
1500 | hole discovered in wall of 8NYA cnc area room 1212 apparently caused by fork truck impact | daniel.sciotti@sanofi.com | 8NYA room 1212 | 2023-04-28 | 2023-04-28 | it is expected that any impact to the wall would be reported to HSE as an event
1499 | 8NYA_Fabrazyme Nano Guard Filter Swap SUTs not charged against process order in MES or Athena. | stanley.durand@sanofi.com | 8NYA Room 1323 | 2023-04-07 | 2023-04-25 | SUT should be issued against process order in MES or Athena.
1498 | 8NYA Upstream FZ End of Run NRSR (23SR8050) samples for DGHCE002 | nicholas.grote@sanofi.com | 8 NYA 1333 | 2023-04-24 | 2023-04-27 | Per the NRSR 6 samples should have been taken.
1497 | 51NYA Daily Critical Utility water site 51-DI2-5 not collected but to Action Notice sign not being removed. | craig.stevenson@sanofi.com | 51 NYA Room 146 | 2023-04-26 | 2023-04-28 | It is expected that all A/N stops signs are removed from entry-ways to allow access after work has been completed
1496 | Unable to Enter Data for Operator that Performed Activity without LIMS account/Left the Company | jessica.kychler@sanofi.com | 68 NYA 4th floor 4700 QC Micro | 2023-04-07 | 2023-04-14 | It is expected that operators performing sampling for LIMS data will have an account set up in LIMS
1495 | PCO2 probe F008-AI-2220028 is out of range/reading erratically. | Brendan.Dickhaut@sanofi.com | 8NYA, Room 1403 | 2023-04-28 | 2023-04-28 | both probes are within acceptable range. option to perform in-line probe adjustment was not in the MES record.
1494 | Filter integrity test performed by unqualified operator at 8NYA | edgar.vale@sanofi.com | 8NYA, Room 1290 | 2023-04-14 | 2023-04-27 | It's expected that the operators performing tasks are qualified on the associated OJTs
1493 | 45NYA Seiver Unit (Main DI #1) does not seem to be functioning properly. | craig.stevenson@sanofi.com | 45NYA | 2023-01-01 | 2023-04-27 | It is expected that the Seiver unit collect TOC and Conducitivity.
1492 | Obsolete version of New Equipment Form submitted for cold room onboarding. | tanni.arif@sanofi.com | 76 NYA. Cold rooms 421A and 421B. Asset numbers 5001455 and 5001456 respectively. | 2022-09-29 | 2023-04-27 | Good Document Practices
1491 | Incorrect Date recorded for RI index result | cindy.guerra@sanofi.com | 80NYA/102 | 2023-04-22 | 2023-04-26 | RI sample would have the correct witness date |  | Deviation
1490 | Missing page from Record Review Tracking Log | rebecca.Cunningham@sanofi.com | 76 NYA | 2023-04-06 | 2023-04-07 | All records be signed out per FBS-SOP-000602 and the Record Review Tracking Log be kept and maintained.
1489 | 8NYA-Fabryzyme- Initial WFI addition exceeded tolerance by 2% for 705675 buffer | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-04-27 | 2023-04-27 | Determination at Gemba |  | Deviation
1488 | OOS Refractometer reading. Reference Temperature cannot be Changed to Perform Sampling. | thomas.loik@sanofi.com | 8RM1350 | 2023-04-25 | 2023-04-25 | 8.95-12.04
1487 | 8NYA room 1503 is weekly expired | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA Room 1503 | 2023-04-26 | 2023-04-26 | its expected that during weekly / monthly cleanings the correct room will be scanned into the record
1486 | Cerezyme -8NYA - SUB Shared Media Take-Out QTY Error | michael.knapp@sanofi.com | 8NYA 1403 | 2023-04-23 | 2023-04-25 | Media is issued against BOM during disposal activities.
1485 | E3 into hold for HiHi steam pressure | adam.rowe@sanofi.com | 51NYA,rm140 | 2023-04-25 | 2023-04-25 | No hihi pressure alarm, or hold | Yes | No Deviation
1484 | 8NYA Cerezyme CM incorrect prompt selected during eluate pH adjust. | stanley.durand@sanofi.com | 8NYA room 1503 | 2023-04-24 | 2023-04-24 | It's expected for the correct prompt to be selected for pH adjust | Yes | Deviation
1483 | Inexecutable Biosurgery SOP | Rebecca.Cunningham@sanofi.com | N/A | 2022-08-17 | 2023-04-24 | SOP should be executable
1482 | 74NYA Room PU1-24-product impact unknown-FB-3659-014 Retentate flow indicator on TFF skid X-3211 found out of tolerance during an unscheduled calibration. | robert.bauch@sanofi.com | 74-PU1-24 | 2023-04-24 | 2023-04-24 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1481 | Missed entry on Portable Media Storage and Usage Log while transporting from 51NYA to 80NYA | stephen.roberts2@sanofi.com | 51NYA Dock | 2023-04-24 | 2023-04-24 | It is expected that the media barrel is moved onto cold truck and all entries are filled in on the Portable Media Storage and Usage Log.. |  | Deviation
1480 | Cerezyme - 8NYA - Aber Probe Deviation | michael.knapp@sanofi.com | 8NYA 1403 | 2023-04-24 | 2023-04-24 | Aber Probes do not deviate or require disabling. | Yes | No Deviation
1479 | Cerezyme - 8NYA - SUB media loss resulting in SUB perfusion and weight alarms | michael.knapp@sanofi.com | 8NYA 1403 | 2023-04-24 | 2023-04-24 | Media is installed and media switchover does not impact SUb perfusion or SUB weight. | Yes | Deviation
1478 | 76NYA-Room 107- QCBP-Product impact unknown-F-9055 Thermometer found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-107 | 2023-04-21 | 2023-04-24 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1477 | Nosebleed in controlled area during media filtration in room 51-124 | nicholas.shoenfelt@sanofi.com | 51-124 Media Batching Room | 2023-04-23 | 2023-04-23 | It is expected that blood is not in the controlled manufacturing areas. |  | Deviation
1476 | MES Record out of Order | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2023-04-22 | 2023-04-22 | A path forward needs to be determined to allow us to continue executing in the record and forward process and meet schedule needs. A more permanent fix of the record order steps will also be needed
1475 | After adding ammsulf there is a 15-minute mix time. But the agitator did not come on | kelvin.glean@sanofi.com | 8 NYA | 2023-04-21 | 2023-04-21 | FBL-EWI-000033 | No | Deviation
1474 | Leak occurred during load / concentration 1 | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1323 | 2023-04-21 | 2023-04-21 | its expected that load will finish without leaks
1473 | 8NYA FZ UFDF Leak occurred during load / concentration 1 | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1323 | 2023-04-21 | 2023-04-21 | Its expected that SUT comes from the vendor assembled appropriately. Filter p/n 690766
1472 | Retain for rhTSH Harvest Fluid (Harvest 11), L/N: DGF1B05 is contaminated | craig.stevenson@sanofi.com | 51 NYA Room 140 | 2023-04-19 | 2023-04-21 | It is expected that the sample be free of foreign matter.
1471 | 8 NYA GEMBA- Fabrazyme UF/DF leak during load concentration 1 | cory.lefort@sanofi.com | 8 NYA 1323 | 2023-04-21 | 2023-04-21 | It is expected that the UF/DF load would run without any leaks coming from the SUT.
1470 | GEMBA for Cz SUB X2230 bleed pressure modbad | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-04-21 | 2023-04-21 | Bleed pump does not have MODBAD alarms
1469 | GEMBA- 8NYA- SUM disconnected following SUM set up | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-04-20 | 2023-04-20 | SUM would remain connected during operation
1468 | CZ Sub 2220x, vent filter connected to wrong WYE on Sub bag. | joseph.gonzalez@sanofi.com | 1403 | 2023-04-20 | 2023-04-20 | SUT remain, integral for production, installation of new vent filters
1467 | 8NYA-CC8 room 1403-product impact unknown-FB-4221-047, FB-4221-048 on MCC Skid X2831 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 8-1403 | 2023-04-20 | 2023-04-20 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1466 | During milling of Lot CBYCMC067 a 1.5 inch gasket was accidentally milled. | jan.stepanik@Sanofi.com | 76 NYA , room 224. | 2023-04-20 | 2023-04-20 | No gasket should have been inside the mill.
1465 | Media was not released before wave scaleup was started | adam.rowe@sanofi.com | 51NYA, room 152A | 2023-04-19 | 2023-04-19 | The media MBR should have been reviewed and product released. The media batch was performed on 08Apr23, the media was used on 19Apr23.
1464 | Gemba – 8NYA – Leak noticed on tote bag during transfer | richard.manasian@sanofi.com | CNC area at 8 NYA | 2023-04-19 | 2023-04-19 | Tote bag would not leak
1463 | CO2 control accidently turned off on Wave-C, after scaleup | adam.rowe@sanofi.com | 51NYA, room 152A | 2023-04-19 | 2023-04-19 | Operator should have clicked the the correct button.
1462 | 8NYA GEMBA-Mod Enzyme Thaw timer exceeded. | Alex.tello@sanofi.com | 8 NYA Room 1220 | 2023-04-19 | 2023-04-19 | path forward
1461 | No reportable results generated for Bioburden CZ Potential Contamination NRSR DGHBN001 23MS00719 | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2023-04-06 | 2023-04-18 | That the analyst document the positive controls in order to verify anaerobic conditions. | Yes | Deviation
1460 | No Reportable result for RODAC B for site 45-112-59, 45 CTMS EM 31MAR23 | megan.botti-marino@sanofi.com | 68 NYA QC microbiology lab | 2023-04-07 | 2023-04-19 | That results be recorded on their associated paperwork
1459 | Approved version of FBL-SOP-000941 used for Routine EM monitoring of the HA fill areas on 29MAR23 | megan.botti-marino@sanofi.com | 76NYA, rooms 355, 354, 356, 358, 352A, 357, 359 | 2023-03-29 | 2023-04-19 | That effective versions of documents be used when executing procedures. | Yes | Deviation
1458 | Cz Zinc Buffer out of Spec | zachary.rickard@sanofi.com | 8NYA 1503 | 2023-04-18 | 2023-04-18 | buffer would be in range during conductivity check
1457 | 8NYA GEMBA: Leak discovered on MCC Skid X-2811 buffer pump inlet of pump P-28110002 at 3/4" x 1/2" reducer Tri-clamp connection | lolav.ahmed@sanofi.com | 8 NYA, room 1403 | 2023-04-19 | 2023-04-19 | n/a
1456 | 3k SUM bag began leaking prior to buffer transfer | richard.manasian@sanofi.com | 8RM1350 | 2023-04-18 | 2023-04-18 | That the bag will not leak.
1455 | Cz Zinc Buffer out of Spec | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2023-04-18 | 2023-04-18 | Forward processing decision to complete Cz zinc production
1454 | Incorrect Orientation for Aseptiquick X connectors on Manifold (762423) | jarad.kukla@sanofi.com | 8NYA / 1212 | 2023-04-18 | 2023-04-18 | The manifold SUT should have been built with the Aseptiquik X connectors in the correct orientation (per NBO-RPT-000417).
1453 | 8NYA-PF8-Room 1364, unknown product impact-Filter Integrity Tester FB-3039 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1364 | 2023-04-14 | 2023-04-18 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1452 | 8NYA-Cerezyme-Tubing ripped during transfer on 768506 Hexylene Glycol buffer | aidan.mula@sanofi.com | 8NYA CNC space | 2023-04-18 | 2023-04-18 | Determination at Gemba |  | No Deviation
1451 | TSH Gemba for DGFSB007 1L Wave Day 3 incorrect gas concentration was selected based on pH | brian.mcgee@sanofi.com | 51NYA/Rm 152A | 2023-04-16 | 2023-04-17 | It is expected for the gas to be changed to 5% CO2 when the pH result is below 6.950. In this case, it wsa 6.915
1450 | CZ zinc WFi is out of temp during dilution | Pavel.dobrovolskiy-wynes@sanofi.com | 1503 | 2023-04-18 | 2023-04-18 | its expected that wfi will remain within temp
1449 | Cz SUB X2250 did not reach capacitance for variable rate bleed | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-04-18 | 2023-04-18 | Capacitance increases until 145pF, triggering variable bleed
1448 | 68NYA-Metrology-Metrology Standard S-1695-010 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2023-04-06 | 2023-04-18 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1447 | Fabrazyme Sample Tested with Incorrect Rinse Fluid | krystal.masseri@sanofi.com | 68NYA QCM 4700 | 2023-04-11 | 2023-04-17 | Per Attachment 2 of FBL-SOP-000019, the sample must be rinsed with 100mL of Fluid A.
1446 | ABYHAM001-A EPIV Stability Revised Results for Initial, 3M, and 6M Timepoints | kristina.erickson@sanofi.com | QCBP 76 NYA RM 107 | Feb2020-Sep2020 | 2023-04-10 | EPIV calculations require the purity result run instead of the control value result which was used inadvertently to process per FBS-MTD-000053 step 6.11.1. | Yes | Deviation
1445 | collection batch ID and autoclave run steps not included in various sections of CZ seed culture MES record | nicole.lufkin@sanofi.com | 8 NYA | 2023-04-17 | 2023-04-17 | steps should be at every LFH set up step |  | No Deviation
1444 | ph/do concerns with TSH bioreactor | tracy.chaves@sanofi.com | 51 NYA room 140 | 2023-04-16 | 2023-04-16 | per FBL-MBR-000091 ph is expected to be between 6.4 - 7.1. 17April23 ph was observed to be at 5.25 |  | Deviation
1443 | ABYHAM001-A EPIV Stability Incorrect Purity Values for Initial, 3M, and 6M Timepoints | kristina.erickson@sanofi.com | QCBP 76 NYA RM 107 | Feb2020-Sep2020 | 2023-03-01 | Correct purity result is 99.6% instead of the 98.6% result used found in the initial Glucuronic assay for ABYHAM001.
1442 | HA Media Simulation - DBYSM2001 - Leak observed in 10L Sample Bags | charlie.houle@sanofi.com | 68 NYA QC sample submission / 76NYA Bulk HA Purification | 2023-04-15 | 2023-04-15 | 10L bags are to be submitted
1441 | GEMBA -8NYA-PH out of range for 705668 RF562D004 DGHBH004 | Richard.Manasian@sanofi.com | 8NYA RM1350 | 2023-04-16 | 2023-04-16 | PH result to be in range between 7.25 - 7.54
1440 | 8NYA-Fz-Buffer information for P/N 705676 not captured at time of execution . | john.lytle@sanofi.com | Rm1333 | 2023-02-09 | 2023-04-15 | it is expected that Mfg operations are documented extemporaneously and capture the required data according to the process requirements
1439 | Leak observed on top of Blue column during buffer cycling after pack. | alyson.belleville@sanofi.com | 51-138 | 2023-04-15 | 2023-04-15 | It is expected that the column would not leak during operation.
1438 | 8NYA- Cz-EM Excursions during flask scale up operations | john.lytle@sanofi.com | 8NYA Upstream Seed Prep | 2023-04-14 | 2023-04-14 | </= 2 EM samples @5micron
1437 | 8NYA-Fz- SUT intall step skipped during HBOD Waste tote setup | john.lytle@sanofi.com | Rm1333 | 2023-02-07 | 2023-04-12 | SUT required for setup operations are utilized according to the record
1436 | Cz-8NYA-SUB HBOD Waste Pinch Valve Leakby | michael.knapp@sanofi.com | 8NYA 1403 | 2023-04-14 | 2023-04-14 | Automated pinch valve for HBOD waste is able to stop all fluid from passing by.
1435 | Gemba – 8NYA- HU missing in comment required per SD-FRA-2022FRACC0074-OCT22 | richard.manasian@sanofi.com | 1350 at 8 NYA | 2022-12-14 | 2023-04-14 | HU would be recorded in record
1434 | CZ 1L Transfer Cap Installation and the need of a VS-CS | shawn.nguyen@sanofi.com | 1343 | 2023-04-06 | 2023-04-14 | Clarity or improves on EWI-000030 EM setup and FBL-SOP-000755 step 6.12 |  | No Deviation
1433 | Missed signature step on Attachment 1 of FBL-SOP-000371 for Fz Bulk Intermediate Reconciliation Form | michael.daly@sanofi.com | 8 NYA cold room 1220 | 2023-03-27 | 2023-04-14 | It is expected the form is filled out completely | Yes | Deviation
1432 | DO high high alram on bioreactor X2230 | Bhargav.Patel@sanofi.com | 8NYA/Room 1403 | 2023-04-13 | 2023-04-13 | Expected results is that the DO stays within accepable range without causing alarm/hold state. | Yes | Deviation
1431 | Insufficient Recorded Time Elapsed Between End of Refeed and Parameter Readings | Brian.McGee@Sanofi.com | 51 NYA RM 140 | 2023-02-07 | 2023-04-13 | Parameter readings must be taken no sooner than 15 minutes after completion of media refeed, and no more than 30 minutes after. | Yes | Deviation
1430 | 51 NYA Thyrogen purification, packed blue column bed height does not meet specified range per FBL-MBR-000029 | Luis.Dominguez@sanofi.com | 51NYA/139 | 2023-04-13 | 2023-04-13 | it is expected that the final packed bed height be 10.0 cm to11.0 cm | Yes | No Deviation
1429 | CZ MMC production recipe could not acquire eluate SUM for "tare" command | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA Room 1503 | 2023-04-13 | 2023-04-13 | its expected that all equipment will be available to be acquired by the recipe
1428 | Recipe went into hold during elution due to SUM hi hi temp | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA Room 1503 | 2023-04-13 | 2023-04-13 | its expected that during elution the SUM temp control will be set to hi hi of 24
1427 | 8NYA-Fabryzyme- pH was re-tested in record after initial failed result. | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-04-13 | 2023-04-13 | Determination at Gemba
1426 | Incorrect Incubator Listed on Settle Plate Paperwork | jessica.kychler@sanofi.com | QC Micro 68 New York Avenue, floor 4, room 4700 | 2023-04-08 | 2023-04-12 | Correct 30-35C incubator should be listed
1425 | 8 NYA GEMBA- Fabrazyme UF/DF skid set-up step missed | cory.lefort@sanofi.com | 8 NYA 1323 | 2023-04-10 | 2023-04-12 | Skid setup would be signed for prior to starting the UF/DF skid cleaning.
1424 | GZ5025-59 Refractive Index sample out of spec | cindy.guerra@sanofi.com | 45NYA Rm. 190 | 2023-04-12 | 2023-04-12 | RI Sample would be within spec
1423 | 76NYA-room 360 HAFL- F-13565 balance found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-360 | 2023-04-12 | 2023-04-12 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1422 | missing check box in FBS-MBR-000083 page 9 for lot DBYSEP006 | fatima.fragaalvarez@sanofi.com | 76NYA room 231 | 2023-01-25 | 2023-04-11 | Document equipment calibration information, ensuring all equipment calibration will remain effective for the duration of this operation.
1421 | Training - Missed count of 1 CFU on 2200035192, Viable Air, Site 226, Room 364, 76 NYA Zone 2, 18Jun2022, Action Level | paul.consolmagno@sanofi.com | 68NYA, QCM | 2022-06-25 | 2023-04-12 | Training - Count of 1 CFU would be submitted for identification per FBL-SOP-XXXXX as the sample is an action level excursion per FBL-SOP-000941
1420 | 76NYA-HAFM room 364-F-4274-017 Temperature Indicator found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 76-364 | 2023-04-12 | 2023-04-12 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1419 | 76NYA-HAFM room 364-F-8090 mixer RPM found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 76-364 | 2023-04-12 | 2023-04-12 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1418 | Change Control 2021FRACC0035 Routine Use Closed without all Pre-Routine Use Works Closed | stephanie.aceto@sanofi.com | 8NYA | 2023-04-03 | 2023-04-12 | All pre-routine works would have been closed prior to the routine use work being closed
1417 | Empty column storage not documented after cleaning | Luis.Dominguez@sanofi.com | 51NYA/122B | 2023-04-09 | 2023-04-12 | FBL-SOP-000907 step 11.8.3.58, column must be stored in 0.5M NaOH | Yes | Deviation
1416 | GEMBA: 8NYA DGHCG004/005 Seed train samples incorrectly sampled in onelims/potential missing bioburden | judith.keller@sanofi.com | 68 NYA / 4700/ QCM/ 8NYA / MFG-CZ- upstream | 2023-04-10 | 2023-04-11 | It is expected that samples are labeled based on their content. It is expected that sampling times in biolims accurately reflect real time sampling time from MES.
1415 | 7.5 hour timer from 24L wave sample to inoc of E4 exceeded | David.Loik@outlook.com | 51 NYA | 2023-04-11 | 2023-04-11 | Inoculation would be completed prior to timer exiry. | Yes | Deviation
1414 | Autoclave probe load was completed but had an Alarm during cycle. | Alex.tello@sanofi.com | 8NYA 1373 | 2023-04-11 | 2023-04-11 | Path forward
1413 | Manual entry missing second verifier in EBR | michael.daly@sanofi.com | 8 NYA 1323 | 2023-03-10 | 2023-04-11 | It is expected manual entries have a second verifier entry in the EBR
1412 | 8NYA-Nilfisk GM80 HEPA Vac removed from Service | thomas.loik@sanofi.com | 8RM1350 | 2023-04-10 | 2023-04-11 | Repair HEPA Vac
1411 | 51NYA D2 bioreactor viral retain sample rejected for leaking | David.Loik@sanofi.com | 51 NYA | 2023-04-10 | 2023-04-10 | Conical would not be leaking. | Yes | Deviation
1410 | CZ-8NYA-MCC Skid 2811 Leak Identified during Buffer Prime | michael.knapp@sanofi.com | 8NYA 1403 | 2023-04-10 | 2023-04-10 | There are no leaks. SIP integrity test captures all poor connections so that the system is integral prior to SIP start.
1409 | Part 714049 not listed as BOM for 761217 Cz record | Bhargav.patel@sanofi.com | 8 NYA / 1403 | 2023-04-10 | 2023-04-10 | Part should be listed as BOM for 761217
1408 | Valves not opening for Buffer Prime | Jeffrey.boldizar@sanofi.com | 8NYA CNC | 2023-04-10 | 2023-04-10 | Get valves to open for recipe
1407 | Climet Alert and Action excursions during Grade A Set-up Procedure | jonathan.lucas@sanofi.com | 51 NYA / RM 140 | 2023-04-10 | 2023-04-10 | FBL-SOP-000755, step 6.7.4. And Table 3. | Yes | No Deviation
1406 | No as found data for Chloride Meter F-3394 | michael.garcia@sanofi.com | 45NYA Room 1130 | 2022-11-06 | 2023-03-20 | Per BQ-GOP-000158 step 5.3.1.2, ensure as-found data can be collected from the UUT (unit under test) | Yes | Deviation
1405 | 76 23MS00740 DBYPM5002 PM5 Media found turbid upon read | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2023-04-08 | 2023-04-09 | Per FBL-SOP-001293, samples are expected to be non-turbid to pass sterility verification
1404 | FRAM_ACTIVITY_CZ-TEC ONELIMS assay template waterbath temperature limits are set to incorrect significant figure | aimee.park@sanofi.com | 68NYA, QC Chemistry 2nd floor | 2023-04-06 | 2023-04-10 | The limits should have been set to 37.0+/- 0.3C, not 37C per Step 6.2.2 of FBL-MTD-000087
1403 | 8NYA-CZ-SUB X2230 Pressure Zeroing Error | michael.knapp@sanofi.com | 8NYA 1403 | 2023-04-10 | 2023-04-10 | The pressure is zeroed within the recipe prior to start of inflation. |  | No Deviation
1402 | 8NYA-Cerezyme-FIT failed three times for filter on 768506 buffer | aidan.mula@sanofi.com | 8RM1290 | 2023-04-09 | 2023-04-09 | Determination at Gemba
1401 | Flake found in Media in 500ml shaker flask at 125ml Volume | adam.hildebrand@sanofi.com | 51-152A | 2023-04-09 | 2023-04-09 | It is expected that process medium does not contain flakes
1400 | 8 NYA GEMBA- Prompt answered prior to probes standardization during Butyl | cory.lefort@sanofi.com | 8 NYA 1323 | 2023-04-08 | 2023-04-08 | Probes are to be standardized and then installed on skid before prompt is answered.
1399 | Thyrogen Media bag #3 for 20L (639945) was overfilled while bag was filling. | kevin.ornell@sanofi.com | 51 NYA Room 124 | 2023-04-08 | 2023-04-08 | Bag 3 clamp would stay secured. Bag 3 and 4 would be filled to required spec.
1398 | 8 NYA GEMBA- Leak on Aquesyn valve during SP skid setup. | cory.lefort@sanofi.com | 8 NYA 1323 | 2023-04-08 | 2023-04-08 | Recipe was put into hold, management was notified, and a GEMBA was scheduled.
1397 | Harvest goods receipt performed multiple times resulting in incporrect harvest day sequential numbering. | adam.hildebrand@sanofi.com | 51-140 | 2023-04-05 | 2023-04-05 | It is expected Goods receipt would be performed once per harvest day
1396 | 8NYA-Fabryzyme-Solution Prep- 30" filter failed FIT due to pressure leak | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-04-07 | 2023-04-07 | At this material to existing NBOD23E0058.
1395 | 8 NYA GEMBA- Recipe failed to load for Fz UF/DF CV | cory.lefort@sanofi.com | 8 NYA 1323 | 2023-04-07 | 2023-04-07 | It is expected that the record loads the correct recipe to delta-v.
1394 | Nitrate sample for WFI site 144 in 76 NYA from 04APR23 can not be found | alexander.decoste@sanofi.com | 68 NYA, Room 4700, QC Micro | 2023-04-04 | 2023-04-06 | Sample can not be located and no valid results will be reported | No | Deviation
1393 | 8NYA-Fabryzyme- Filter made loud noise during FIT and recipe aborted, failing to pressurize. | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-04-07 | 2023-04-07 | Determination at Gemba
1392 | Bioburden sample 647152 CZ2_10L_760844_POU760744 DGHCG003 expired at 1028 and will not have reportable results | alexander.decoste@sanofi.com | 68 NYA, Room 4790 (bioburden room), QC Micro | 2023-04-06 | 2023-04-06 | Sample should have been tested before 1028 on 06APR23
1391 | PH Adjustment Error Buffer DGFB2039 | jeff.hill@sanofi.com | 45 NYA, RM190 | 2023-04-06 | 2023-04-06 | Correct volume of sodium hydroxide will be added for PH titration.
1390 | 51NYA-APT SS Corridor 127-Differential Pressure Indicator F-16113-310 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 51-127 | 2023-04-05 | 2023-04-06 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1389 | 8NYA - Cerezyme - pH probe calibration performed on SUB with incorrect value | michael.knapp@sanofi.com | 8NYA 1403 | 2023-03-09 | 2023-04-02 | pH calibration is performed with the adjusted BGA value.
1388 | iLims number canceled by QCM accidently for Sodium Hydroxide Solution buffer 1.0M @ 45NYA (iLims #: 230007327) | craig.stevenson@sanofi.com | 68NYA Room 4700 | 2023-04-04 | 2023-04-04 | It is expected that QC does not cancel iLims numbers associated with product/buffers that require QC testing.
1387 | Questionable/unverifiable GDP entry on FBS-SOP-000737 | craig.stevenson@sanofi.com | 76NYA, Room 355 | 2023-04-04 | 2023-04-05 | It is expected that FBL-SOP-000942 V2.0 step 6.1.1, step 6.1.2, step 6.1.3, step 6.1.4 for GDP.
1386 | Incorrect sample volume taken during Fabrazyme butyl end of hold sampling | michael.daly@sanofi.com | 8 NYA 1333 | 2023-04-03 | 2023-04-04 | It is expected the correct volume is sampled
1385 | 8NYA-Fabryzyme- Follow-up Gemba of FRAD23M00095 Event | aidan.mula@sanofi.com | 8RM1350 | 2023-03-24 | 2023-03-24 | Follow up to determine root cause
1384 | 68NYA-Room 4700 Dept QCM- Glass thermometers F-16400 and F-16479 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 68-4700 | 2023-04-03 | 2023-04-04 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1383 | Cold room 51-138 in QCA | jacob.kallon@sanofi.com | 51 NYA Room 138 | 2023-04-04 | 2023-04-04 | It is expected that the cold room temperature be within range (6-10 degrees Celsius)
1382 | 8 NYA GEMBA- Incorrect recipe loaded from the record for Qual Skid Cleaning | cory.lefort@sanofi.com | 8 New York Ave / 1303 | 2023-04-03 | 2023-04-03 | The expected results would be that the record would load a qual skid cleaning with a CV recipe not qual. skid cleaning without CV. | Yes | Deviation
1381 | BRX PM Procedure Step Signed but Not Performed | james.letourneux@sanofi.com | 51 NYA RM 140 | 2023-04-02 | 2023-04-03 | Per FBL-SOP-000819, Attachment 3, "O-Ring Replacement" section, Steps 2-3, operator must replace o-rings for "pH Probe Interior/ Exterior" and "pH Backup Probe Interior/Exterior."
1380 | 8NYA Fz x2200 MICRO/MACRO sparge lines are swapped | paola.tello@sanofi.com | 1333 | 2023-03-07 | 2023-04-03 | MICRO line connected to the front line. Macro line in the back line
1379 | 76NYA-HAFM-Biosurgery-Temperature Gauge F-18693 found out of tolerance during an unscheduled calibration for not working | robert.bauch@sanofi.com | 76-Varies | 2023-03-31 | 2023-04-03 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1378 | 8NYA CC8 Room 1403-Skid BR-2240001, Temperature indicating controllers FB-3988-029 and FB-3899-032 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 8-1403 | 2023-03-31 | 2023-04-03 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1377 | Thyrogen Drug Product Lot 9W1648 36M Stability Time Point: Isoaspartic Acid Testing Performed, but not required per FBL-PTM-000049 | yekaterina.zavgorodniy2@sanofi.com | 45 QCC | 2022-08-13 | 2023-04-03 | It is expected that Isoaspartic acid is not tested at 36M for lot 9W1648.
1376 | 8NYA CZ DGHBN001 CB-1 Activity samples pending receipt in OneLIMS | cleiton.lopes@sanofi.com | 8NYA | 2023-02-07 | 2023-04-03 | MFG to sample per sampling plan and submit to drop off station for MLC pickup and delivery to QC.
1375 | Quarterly Trend Reports Not Completed by Due Date Per FBL-SOP-001222 | Brendan.Scanlon@sanofi.com | 68 NYA | 2023-04-01 | 2023-04-01 | Reports are completed by end of quarter per FBL-SOP-001222
1374 | 8 NYA AN 2023-I-016 stop signs removed prior to QA Release and 2023-I-017 released by QA with already removed stop signs | molly.pomeroy@sanofi.com | Room 1212 (CNC North), Room 1422 (CNC Vestibule), Room 1411 (MAL)) | 2023-03-31 | 2023-04-03 | It is expected that the AN stop signs are removed following the release of the AN by QA.
1373 | Gemba – 8NYA – Leak noticed on SUT during Fz Harvest Sampling | Yves.Antoine@sanofi.com | Room 1333 | 2023-04-03 | 2023-04-03 | FBL-EWI-000009, FBL-EWI-000032, FBL-EWI-000015, FBL-EWI-000031 | Yes | Deviation
1372 | Storage unit unable to be scanned due to a unit of measure error. | ricardo.rodriguez@sanofi.com | 8RM1350 | 2023-04-02 | 2023-04-02 | All storage unit labels will scan.
1371 | 8NYA - Cerezyme - Approved Document FBL-SOP-000755 Used for Dynamic Monitoring | michael.knapp@sanofi.com | 8NYA 1343 | 2023-03-31 | 2023-04-02 | The Effective Document is used for operations.
1370 | GEMBA: 45 NYA - Mixing time did not reach min. | William.Bradley@sanofi.com | 45NYA | 2023-03-26 | 2023-04-02 | Mix is to be within the acceptable range of 15 minutes to 240 minutes according to the record.
1369 | Gemba – 8NYA – WFI overshoot during QS on SUM 1505 | richard.manasian@genzyme.com | 1350 at 8 NYA | 2023-04-02 | 2023-04-02 | WFI would fill to target per receipe
1368 | Batch overfilled during QS WFI fill. | maria.palmar@sanofi.com | 8 New York Ave, 8RM1350 | 2023-04-01 | 2023-04-01 | The bag should not overfill
1367 | Filter failed FIT “error pressure not obtainable.” | alex.tello@sanofi.com | 8 NYA Room 1350 | 2023-04-01 | 2023-04-01 | Wetted with WFI following FBL-EWI-000021
1366 | A FIT test recipe launches but reaches a "pressure not obtainable" | christopher.arnold@sanofi.com | 8RM1350 | 2023-04-01 | 2023-04-01 | That the FIT would launch and test to either a pass or failure.
1365 | Initial WFI fill exceeded target by more than 1%. | julie.auguste@sanofi.com | 8RM1350 | 2023-03-31 | 2023-03-31 | Initial WFI will be within target specifications.
1364 | Gemba – 8NYA – Expiration incorrect for buffer | william.lee@sanofi.com | 8NYA/1350 | 2023-03-30 | 2023-03-30 | The buffer expiry should be 60 days | No | Deviation
1363 | Incorrect number of cycles generated during Cz zinc repack | michael.daly@sanofi.com | 8 nya 1303 | 2023-02-14 | 2023-03-30 | It is expect the correct number of cycles is taken out and generated during a repack. For a Zinc cerezyme repack the correct number of cycles to be taken out during the unpack is all current remaining cycles. Then when the repack is complete all current remaining cycles are created after the qualification passes
1362 | Travel stop improperly set on 2-bottle sampler resulting in unusable sample. | james.letourneux@sanofi.com | 51 NYA Rm 140 | 2023-03-30 | 2023-03-30 | Travel stops on valves should be set properly to allow the actuator pin to fully compress the valve's diaphragm to a closed state, fully obstructing flow. |  | No Deviation
1361 | RF505C006, CGHCM006 failed TOC swab result failing CV (22MS02072) | daniel.bati@sanofi.com | 8NYA | 2022-08-29 | 2023-03-30 | Scheduled a Gemba | Yes | Deviation
1360 | 45NYA Gemba for missing GCB Logbook FBL-LBK-000027 (2022LBT0060) | brian.mcgee@sanofi.com | 45NYA Room 101, Records management room | 2023-03-30 | 2023-03-30 | It is expected that the logbook remain in the RM room or in the GCB area while it is in use. |  | Deviation
1359 | 68NYA-Room 1740-QCVAL- IRTD F-21406 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1740 | 2023-03-22 | 2023-03-30 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1358 | Big particle observed in Aseptic Challenge sample | Rebecca.Schmidt@sanofi.com | QC Micro | 2023-03-29 | 2023-03-29 | FBL-SOP-000873 6.2.4
1357 | Vial shipment not picked up for shipment | adam.hildebrand@sanofi.com | 80NYA RM166 | 2023-03-29 | 2023-03-29 | It is expected logistics picks up vial shipment to go to loading dock for fedex.
1356 | D1 Inoculation Line cannot be used for D1 Inoculation as required by FBL-SOP-000680, as the line does not reach the required RTD/steam trap. | james.letourneux@sanofi.com | 51 NYA Rm 140 | 2023-03-29 | 2023-03-29 | Equipment that satisfies procedural requirements should be provided.
1355 | Scales standardization ranges changed in iNOVA MES | jacob.kallon@sanofi.com | 51 NYA Room 132/134 | 2023-03-28 | 2023-03-28 | It is expected that floor scale and bench to scale standardizations have the correct ranges in MES
1354 | RQ Failed acceptance criteria of 2-8C, Temperature drops to 0.3C | David.zhong@sanofi.com | 51NYA | 2023-03-28 | 2023-03-28 | RQ - F051-REF-FB-1991-51NYA-CQ-0003. Temperature must be 2-8C |  | No Deviation
1353 | Gemba – 8NYA – Leak noticed on SUT during buffer preparation | william.lee@sanofi.com | 8NYA | 2023-03-28 | 2023-03-28 | SUM bag should not have leaked
1352 | 76NYA-Room 3105 Dept HABP-product impact unknown--no as found data taken on Pall FIT FB-2830 before repair | robert.bauch@sanofi.com | 76-3105 | 2023-03-28 | 2023-03-29 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 As found data should have been collected.
1351 | UFDF Final Yield calculation out of range | michael.daly@sanofi.com | 10 Cal Room 138 | 2023-03-29 | 2023-03-29 | It is expected the final yield calc is greater than or equal to 85%
1350 | 8NYA-CC8 room 1403-product impact unknown-FB-2899-069, FB-2899-070 on MCC Chromatography Skid X-2560001 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1403 | 2023-03-28 | 2023-03-29 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results | No | Deviation
1349 | Sum Bag leak upon starting WFI fill. | ricardo.rodriguez@sanofi.com | 8RM1350 | 2023-03-28 | 2023-03-28 | That the bag would not leak.
1348 | D1 HiHi temperature alarm during aeration | david.loik@sanofi.com | 51 NYA | 2023-03-28 | 2023-03-28 | Temperature would be controlled within range.
1347 | Gemba – 8NYA – Cracked air filter observed during inflation | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-03-28 | 2023-03-28 | SUT should be free from defects | Yes | Deviation
1346 | JIT Warehouse RQ Study Failure | Robert.petersjr@sanofi.com | 8NYA/JIT Warehouse | 2023-03-28 | 2023-03-28 | Temperature must be between 25 - 30C per F008-ED-UR-0111 | Yes | No Deviation
1345 | 8NYA - Cerezyme - SUB Aber Probe Transmitter and DeltaV Scaling Issues | michael.knapp@sanofi.com | 8NYA 1403 | 2023-03-28 | 2023-03-28 | The transmitter and deltaV scaling for Aber probes is aligned and scaled appropriately for Cz. |  | No Deviation
1344 | Conductivity Sample failed and retested without resample - no ability in MES or EWI to instruct path forward if a sample fails | molly.pomeroy@sanofi.com | 8 NYA Solution Prep | 2023-02-08 | 2023-03-28 | FBL-SOP-001320 Step 6.8.4 Review of Exception Log (QA) - Review “system description” field, “user description” field, and “comment” fields for each exception and ensure a clear explanation, resolution and justification are provided
1343 | Dynamic EM for an aseptic qualification performed without completing 10 non-viable samples and no zero filter taken | megan.botti-marino@sanofi.com | 51 NYA | 2023-03-08 | 2023-03-28 | That at least 11 non-viable readings be taken during the critical operation and that a zero filter be performed prior to the operation.
1342 | DBYFRM015 PLC display did not match transmitter or chart recorder | nicholas.macdonald@sanofi.com | 76-364 | 2023-03-27 | 2023-03-27 | It is expected that the pH on PLC display read the same as the pH reading on the transmitter and chart recorder.
1341 | Steps executed without necessary training | charlie.houle@sanofi.com | 76 NYA; Room 3102 (HA Purificaiton) | 2023-03-26 | 2023-03-27 | It is expected those executing and verifying steps within MBRs and logbooks have required training for those steps.
1340 | Fz Nano load SUM ran empty during buffer chase step | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-03-27 | 2023-03-27 | During buffer chase pump should add chase buffer fast enough to keep the sum weight between 5-10kg | Yes | Deviation
1339 | Expired 51NYA Aseptic Challenge samples | cleiton.lopes@sanofi.com | 68 Sample Management/QC Micro 4700 | 2023-03-24 | 2023-03-24 | Test Request to be filled out and submitted properly and EM to be incubated within 24hr of sampling.
1338 | 8NYA Fz Zinc Chromatography - ANMS alarm no populated in CIP record. | stanley.durand@sanofi.com | 8NYA Room 1323 | 2023-03-20 | 2023-03-27 | Quality Critical alarms should be reflected in the batch record
1337 | 8NYA Cerezyme CM column failed CV | stanley.durand@sanofi.com | 8 NYA room 1303 | 2023-03-24 | 2023-03-27 | TOC results came back 5.02ppm for swab 2, which falls outside of our reporting range of 0.76 to 2.4ppm per FBL-SOP-000484.
1336 | 8 NYA - Wave BR-2162001 Load Cells found out of specification during Vendor PM | nicholas.grote@sanofi.com | 8 NYA Room 1403 | 2023-03-16 | 2023-03-16 | It is expected that the 1.0 kg and 12.kg test points for the load cells are within the acceptable ranges. | No | No Deviation
1335 | 8NYA GEMBA: Cz Wave Bag Sampling step not executed in the record | lolav.ahmed@sanofi.com | 8NYA room 1403 | 2023-03-26 | 2023-03-27 | sampling step requires execution on the day of sample performed
1334 | Temperature control of Microcarriers and PBS solution failed in D1 Bioreactor | adam.hildebrand@sanofi.com | 51-140 | 2023-03-26 | 2023-03-26 | Detailed instruction to run the Bioreactor Temp circuit manually | Yes | Deviation
1333 | 8 NYA GEMBA- Burst product inlet line for Fz UF/DF | cory.lefort@sanofi.com | 8 NYA 1323 | 2023-03-26 | 2023-03-26 | Expected results are that the product inlet line would maintain integrity and would flow to filter without issue.
1332 | Condo probe pulled out of the skid to perform re-calibration | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1323 | 2023-03-25 | 2023-03-25 | its expected that all instruments will remain in the skid post system cleaning
1331 | Supervisor not trained in the latest Batch record review SOP | valeria.moreno@sanofi.com | 76 NYA | 2023-03-23 | 2023-03-23 | FBS-SOP-000466(en-EN)15.0
1330 | Maintenance PHT and SIP not documented in MBR but was performed according to DELTA-V recipe completion. | adam.hildebrand@sanofi.com | 51-140 | 2023-03-23 | 2023-03-25 | It is expected that data is recorded contemporaneously when Maintenance PHT and SIP | No | Deviation
1329 | 8NYA Room 1343 Incubators not sending data to EMS | Brandon.Steff@sanofi.com | 8NYA 8RM1343 INC-2100-002, INC-2100-003 | 2021-07-27 | 2023-03-24 | PI records incubator data
1328 | 51NYA TSH Overdue Manual Cell count requalification | Adriana.Alban@sanofi.com | 51NYA | 2023-03-24 | 2023-03-24 | Operators are requalifying on manual cell counts every 730 days
1327 | Butyl MCC Load adjust Product sensor leak during Transfer | christian.venegas@sanofi.com | 8NYA upstream Room 1333 | 2023-03-24 | 2023-03-24 | To be able to replace sensor and continue with the transfer
1326 | 8 NYA GEMBA- Failed visual and CV swab for CM Column | cory.lefort@sanofi.com | 8 NYA Column Packing room | 2023-03-22 | 2023-03-22 | There should be no discoloring on the distribution plate in the CM column at anytime.
1325 | No product-no lots-68NYA: F-16089 BSC usage log not completed for water bioburden assay occasion | larissa.fairweather@sanofi.com | 68NYA/4700/QCM | 2023-03-22 | 2023-03-23 | FBL-SOP-000594 V.8.0 Step 6.5.3 Document pre- and post-use cleaning/disinfection on FBL-LBK-000202. Perform in order of operations. and Step 6.3.4 Ensure magnehelic gauge is within the acceptable range for the BSC and document on FBL-LBK-000202. | Yes | Deviation
1324 | Storage units not scanning into to EBR due to unit of measure error. | jermaine.kellam@sanofi.com | 8RM 1350 | 2023-03-23 | 2023-03-23 | All storage units are expected to scan into the EBR.
1323 | Step performed without verification | valeria.moreno@sanofi.com | 76-355 | 2023-03-21 | 2023-03-23 | FBS-MBR-000076
1322 | 2023-I-014 AN Work Completed without Area Clearance Section being completed | molly.pomeroy@sanofi.com | 8 NYA Room 1403 | 2023-03-19 | 2023-03-22 | It is expected that area clearance is performed Doer of Work verifies with Area Owner(s) that work may commence and verifies and that Area Owner (s) have signed the Area Clearance per FBL-SOP-000336 step 6.4.5 prior to the completion of the work
1321 | Low cell count on final harvest day of bioreactor E3 | david.loik@sanofi.com | 51 NYA | 2023-03-23 | 2023-03-23 | No deviations from specifications in FBL-CST-00000007
1320 | FBS Historical data from OSI PI Historian not available for Gemco Dryer in room 76-229 | Shiful.Kohli@sanofi.com | FBS 76-229 | TBD | 2023-03-23 | TBD
1319 | 20R APS Run 6 EM sample by date and times occur after EM packet submitted and received. Attachment 2 of 21 is missing from EM packet. | katherine.ditommaso@sanofi.com | 76-406 | 2023-03-22 | 2023-03-22 | Dates and times align with activities performed contemporaneously.
1318 | LIMS admin not trained on FBL-SOP-001471 (BioLIMS master data management) | nicholas.lapointe@sanofi.com | N/A | 2023-02-16 | 2023-03-23 | Admin should have been trained on procedure prior to performing change management activities for OneLIMS.
1317 | 68NYA-Room 1830-Metrology-No product impact-Humidity Standard FBS-1568 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2023-03-17 | 2023-03-23 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1316 | 8NYA Upstream WFI Batch Records with Incorrect/Missing Takeouts | michael.knapp@sanofi.com | 8NYA 1333 | 2022-09-05 | 2023-03-19 | The correct bag is used for the associated record and the takeout is performed.
1315 | 8NYA Fabrazyme SUT recorded into record missing handling unit | stanley.durand@sanofi.com | 8NYA Room 1323 | 2023-03-06 | 2023-03-22 | It's expected that for all manual entry into the record of non BOM items Part numbers, Batch, HUs and expiry should be entered
1314 | Multiple GDP errors on HA MEDIA Fill packet from 21Mar2023 | craig.stevenson@sanofi.com | 76NYA, Room 355 | 2023-03-21 | 2023-03-22 | All data to be collected per FBS-SOP-000607
1313 | Incorrect choice made on Fz UFDF cleaning record step | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-03-21 | 2023-03-21 | Operator should have answered "yes" causing the recipe to run a CV |  | No Deviation
1312 | 8NYA-Room 1323 Dept PF8-Delta V weight indicator FB-4235-007 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 8-1323 | 2023-03-21 | 2023-03-22 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1311 | Thyrogen - DGF3B03 - 51 NYA - Missed witness step during production | conner.booker@sanofi.com | 51 NYA, Room 140 | 2023-03-11 | 2023-03-11 | The witness step was to be signed by the person who witnessed the full operation of the procedure per the SOP.
1310 | QCM Analyst inadvertently removed results from Bioburden sample. | brandon.lebherz@sanofi.com | 68 NYA, QC Micro Lab, | 2023-03-11 | 2023-03-12 | Is it expected that analysts should not be able to remove results from tests during routine operations. There is currently not a procedure that outlines the steps to remove results. | Yes | Deviation
1309 | No reportable result for non-viable site 210, 76 Biosurgery for week of 19FEB23 | megan.botti-marino@sanofi.com | 76 NYA, room 355 | 2023-02-23 | 2023-03-21 | Per FBL-SOP-000941, site 210 is to be sampled weekly
1308 | Harvest bag replacement on X2200/ Harvest Collection Bag Replacement on X2200 | marcelo.dossantos2@sanofi.com | 1330 | 2023-03-20 | 2023-03-20 | Harvest material should not be allowed to expire.
1307 | 76NYA-Room 3102- HABP-product impact unknown- Pressure Gauge F-18956 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 76-3102 | 2023-03-20 | 2023-03-21 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1306 | Remote mixing on UF\DF Load sum not starting | Christian.Venegas@Sanofi.com | 8NYA upstream Room 1323 | 2023-03-20 | 2023-03-20 | to fix the issue and remote mixing to be resumed |  | No Deviation
1305 | Analyst used an incorrect rinse volume while testing 23MS00557 | Mitchell.Wilkins@Sanofi.com | 4700 | 2023-03-12 | 2023-03-20 | The rinse volume should read 100mL
1304 | GEMBA: 68NYA-QCM (8NYA Skid qual, CU samples) M9 run parameters with incorrect # rejects per FBL-SOP-000595 for TOC/COND | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2023-03-09 | 2023-03-19 | Per FBL-SOP-000595 6.4.7
1303 | Reject Key Usage Log not completed for 80NYA Reject Cage Keys | Paul.McFarland@sanofi.com | 68NYA Records Mgmt Room | 2023-03-15 | 2023-03-15 | Attachment 3 of FBL-SOP-000900 would be completed each time the reject keys were accessed by QA.
1302 | No Negative Control Submitted for EM paperwork 76NYA CTM | kathryn.lentoni@sanofi.com | 76NYA | 2023-03-13 | 2023-03-20 | Negative control should've been submitted
1301 | flakes found in multiple BI's during final read for 23MS00549 | Stefanie.Erickson@sanofi.com | QC Micro counting bench | 2023-03-10 | 2023-03-18 | for the BI's to be non-turbid
1300 | 76NYA-Room 3104 Biosurgery-product impact unknown- F-21186-011 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 76-3104 | 2023-01-11 | 2023-03-20 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1299 | 51NYA-APT CC rooms 148 and 153-product impact unknown-FB-0411 and FB-0416 Terminal HEPA filters found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 51-148 and 51-153 | 2023-02-16 | 2023-03-20 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1298 | Buffer tote found to have a leak when tote was to be moved. | maria.palmar@sanofi.com | 8RM 1202 | 2023-03-19 | 2023-03-19 | That the tote bag will not leak.
1297 | FZ MCC 2510 UV prompt sensor baseline troubleshoot during production start up | shawn.nguyen@sanofi.com | 8NYA Room 1333 | 2023-03-19 | 2023-03-19 | UV baseline to be zero'd and recipe progressed
1296 | 8NYA GEMBA: Launched CIP w/CV recipe failed to progress on DeltaV for MCC Skid X2560 | Lolav.ahmed@sanofi.com | 8NYA room 1403 | 2023-03-19 | 2023-03-19 | n/a
1295 | Fz-The MCC X-2543 column was loaded outside the range of 1-5% Delta UV which is a nKPP in the control strategy | john.lytle@sanofi.com | Rm1333 | 2023-03-18 | 2023-03-18 | It is expected that equipment profibus connections are in an adequate state to allow for proper execution of Delta V automation controls. It is also expected that Fz MCC column load be maintained between 1-5% breakthrough
1294 | Part 692169 failed bag inspection | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-03-19 | 2023-03-19 | Bag would have one airline and one recirculation line
1293 | FZ Once every Adhere eluate #4 Harvest Vessel Retain sample labeled with two different inventory numbers. Missing separate Harvest Vessel sample | shawn.nguyen@sanofi.com | 8NYA Room 1333 | 2023-03-17 | 2023-03-17 | Per FBL-RPT-000376, Table 11, two Process Impurities Retains were to be taken.
1292 | Cerezyme Phenyl Equil BIOB sample unable to be received by QC due to sample expiry | stanley.durand@sanofi.com | 8NYA rm 1503 | 2022-12-16 | 2022-12-17 | Per FBL-EWI-000009 Cz Phenyl_Equil_BIOB sample should be received and tested by QC within the 24 hour expiry.
1291 | 760744-RF623D002 part 692321 can not be used and is in BOM | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-03-18 | 2023-03-18 | All materials in the BOM should be able to be used.
1290 | Cell count incorrect and missing verification | adam.rowe@sanofi.com | 51NYA,rm140 | 2023-02-07 | 2023-03-17 | Cell count should be accurate, and verified
1289 | Filter Integrity Test (FIT)multiple results of PRESSURE NOT OBTAINABLE | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-03-17 | 2023-03-17 | Filter would have a passing result
1288 | Non-contemporaneous data entry for Fabrazyme ICB RP Purity batch | John.Kapaniris@sanofi.com | 68NYA | 2023-01-09 | 2023-03-16 | Data needs to be entered at time of assay occurrence | Yes | Deviation
1287 | Missed inspection of (3) pest control devices during February service inspections | andrew.doheny@sanofi.com | 51 NYA Utility Room 113 and 8 NYA Interstitial Space East | 2023-02-08 | 2023-03-15 | In accordance with the site pest control procedure FBL-SOP-000512, Pest Control, V9, step 6.2.2, the vendor is to inspect all rodent traps, insect lights and monitors.
1286 | 8NYA GEMBA: SUB X2220 high SUB pressure | Lolav.Ahmed@sanofi.com | 8NYA room 1403 | 2023-03-17 | 2023-03-17 | FBL-CST-000029
1285 | EM setpoints for temperature did not save on butyl adjust SUM and agitator did not work properly once recipe was launched. | Brendan.Dickhaut@sanofi.com | 8NYA/ Room 1333 | 2023-03-16 | 2023-03-16 | Expected results is that recipe set points and equipment function properly under recipe
1284 | Step was not executed prior to launching butyl production recipe step scadDA101 | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1323 | 2023-03-16 | 2023-03-16 | its expected that operator will fulfill every step in the record prior to starting production recipe.
1283 | Cz samples were sampled in Biolims but never delivered/received by QC | Brandon.Steff@sanofi.com | 8RM1403 | 2023-02-07 | 2023-03-10 | Samples taken delivered/received by QC, corresponding results entered into biolims
1282 | 68 NYA/8NYA- NRSR 23SR8024 taken in incorrect container | judith.keller@sanofi.com | 68 NYA / 4700/ QCM/ 8NYA / MFG-FZ/ downstream | 2023-03-16 | 2023-03-16 | Samples should be sampled per NRSR
1281 | No Products - No Lots - MassBio Campus - No Means to Document an Event-Based Audit Trail Review | shana.kelley@sanofi.com | QA Digital / 68 NYA | 2023-03-16 | 2023-03-16 | Requiement to perfom an Event-Based audit trail review is documented in Step 5.1.3 of BQ-GOP-000300 and section 7 of BQ-FRM-000899: Audit Trail Review Assessment Form
1280 | 8 NYA EM Paperwork missing verifiable source for late entry for initial incubation of samples | megan.botti-marino@sanofi.com | 68 NYA | 2023-02-12 | 2023-03-15 | It is expected that incubation of samples be documented at the time of occurrence and that late entries include verifiable sources.
1279 | SUM Failure - leak | leo.gillespie@sanofi.com | 1350 - Solution perp | 2023-03-15 | 2023-03-15 | Batch put on hold as per FBL-EWI-000039
1278 | CM Column Resin Expired | zachary.rickard@sanofi.com | 8NYA/1303 | 2023-03-14 | 2023-03-14 | Column resin would not be expired | Yes | Deviation
1277 | FZ MCC 2510 Column 2, bleed bag to column connect assembly found melted on 1 unused port of the 3 available ports | shawn.nguyen@sanofi.com | 8NYA Room 1333 | 2023-03-14 | 2023-03-14 | SUT remain integrals for production
1276 | Fz-MCC X-2210 bubble trap wetted out during Buffer Prime operations | john.lytle@sanofi.com | Rm1333 | 2023-03-14 | 2023-03-14 | bubble trap functionality is dependent on calibrated sensors to check for liquid/ air levels.
1275 | Fz-MCC X2543 held during Buffer Prime for SUM Agitator position | john.lytle@sanofi.com | Rm1333 | 2023-03-13 | 2023-03-13 | MCC should not be looking for SUM 1 agitator position to ensure completion of Buffer Prime operations
1274 | Unlabeled sample received with ILIMS test request for DGF1B01B ILIMS 2300005599 | katherine.ditommaso@sanofi.com | 68 NYA, 4th floor sample receipt | 2023-03-13 | 2023-03-13 | FBL-SOP-000964 section 3.0. Ensure samples are labeled per requirements.
1273 | Analyst did not record monthly water bath clean on attachment 3 "Monthly Maintenance Checklist for Water Baths Log sheet" for FB-2323 during both December and January. | Mitchell.Wilkins@Sanofi.com | 4700 | 2023-03-13 | 2023-03-13 | For there to be a filled out form. | Yes | Deviation
1272 | 8NYA GEMBA: FZ MCC Buffer Alignment Error | Lolav.Ahmed@sanofi.com | 8NYA room 1333 | 2023-03-13 | 2023-03-13 | FBL-EWI-000032 REF 1 table for Fz buffer information and MCC inlet valves | Yes | Deviation
1271 | insufficient volume during media load for 12L wave | lindsey.hymer@sanofi.com | 152A | 2023-03-12 | 2023-03-12 | per FBL-MBR-000108 weight should reach between 11.00kg and 11.22kg
1270 | Incorrect buffer connected to cnc manifold port 6 (pompe copper shakedown run) | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA cnc Suite 2 | 2023-03-11 | 2023-03-12 | Its expected that correct buffer will be connected to the correct tote
1269 | E3 Agitator Turned on Prematurely During Harvest | james.letourneux@sanofi.com | 51 NYA Rm 140 | 2023-03-11 | 2023-03-11 | Per FBL-SOP-000680, Step 6.19.8, "Once impeller hub is covered with media, start agitation." Agitator should not have been started until media refeed reached sufficient volume in the bioreactor to cover the impeller hub.
1268 | Leak found in ceiling near previously remediated leak | adam.hildebrand@sanofi.com | 51-140 | 2023-03-10 | 2023-03-10 | No leaks detected
1267 | FBS Alarm Responders Not Trained on FBL-OJT-000697 | rebecca.Cunningham@sanofi.com | 76 NYA | 2022-08-02 | 2023-03-10 | All alarm responders trained on OJTs for tasks they are performing.
1266 | Rust observed on top of BSC in media prep room 51-124 | nicholas.shoenfelt@sanofi.com | 51-124 | 2023-03-10 | 2023-03-10 | It is expected that the clean rooms and process equipment within them be free of rust.
1265 | BQ-TRS-000355 GMP Concepts course- iLearn content and approved storyboard are not the same version | patricia.flowerselovitz@sanofi.com | Mass BioCampus | 2020-08-07 | 2023-03-09 | The expected outcome is the GEODE+ approved item is the same version as the Articulate file in our content repository.
1264 | Incorrect buffer volume transferred GZ1289-56 DGFB2049 | cindy.guerra@sanofi.com | 45NYA Rm. 190 | 2023-03-07 | 2023-03-07 | Buffer volumes should align with transfer volume | Yes | No Deviation
1263 | Discrepancy with Butyl MCC eluate weight | Christian.Venegas@Sanofi.com | 8NYA upstream Room 1333 | 2023-03-10 | 2023-03-10 | to have the correct weight for MCC Eluate
1262 | GEMBA: 68 NYA - QCM LAL plate reader missing data file | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2023-03-08 | 2023-03-08 | Instruments should be compliant and data files should not delete
1261 | FBS Alarm Responders not trained on the effective version of FBL-SOP-000316 | rebecca.Cunningham@sanofi.com | 76NYA, 74NYA | 2023-01-17 | 2023-03-09 | All alarm responders are trained on the effective version of FBL-SOP-000316
1260 | Missed Bench cleaning entries in QC Chemistry Labs | John.Kapaniris@sanofi.com | 68NYA QC Chemistry ICB Lab/45NYA QC Chemistry Lab | 2022-05-29 | 2023-03-09 | Initial/Date indicating completion of Bench Cleaning in Logbook FBL-LBK-000181 | Yes | No Deviation
1259 | Missing Verification (Initial/Date) on FBL-PTM-000332, Table A2.2 and A2.3 | sheila.pandolfo@sanofi.com | 8 NYA Room 1403 | 2022-11-28 | 2022-11-28 | It is expected the verifier would have initial and dated the first row of both tables.
1258 | Qual Skid Post Pack CV | zachary.rickard@sanofi.com | 8NYA/1303 | 2023-03-09 | 2023-03-09 | Record will have correct recipe to be launched for post use skid CV.
1257 | No reportable result for bi-weekly EM Non-Viable site CC2-80 | Craig.Stevenson@sanofi.com | 74NYA, CC2-01 | 2022-12-01 | 2023-03-09 | Per FBL-SOP-001039 V5.0 Step 6.5.6, the flow rate should be 50LPM, ±5% (47.5-52.5LPM).
1256 | 8 NYA Trend of Multiple Data Integrity Deviations | valerie.staras@sanofi.com | 8 NYA | 2023-03-09 | 2023-03-09 | N/A
1255 | End time on Credo logbook 2023QCC0022 was not entered on 04Mar2023 | michael.montiverdi@sanofi.com | Samples were delivered to 68 NYA Sample receipt room | 2023-03-04 | 2023-03-08 | End time to be entered. |  | No Deviation
1254 | 68NYA Room 4700 Dept QCM68-TOC's F-19931 and F-19933 were found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 68-4700 | 2023-03-02 | 2023-03-09 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results | Yes | Deviation
1253 | Multiple systems put into GMP use for Cerezyme without all CCR works closed for equipment in common with commercial products | debra.finucane@sanofi.com | 8NYA | 2022-11-01 | 2023-03-09 | All CCR Works to be completed before system released per BQ-GOP-000281 | No | Deviation
1252 | Cz Zinc Column Resin Prep agitated during mixing time at incorrect speed. | robert.morawski@sanofi.com | 8NYA RM1303 | 2023-03-08 | 2023-03-10 | Agitation was run at 237 RPM during mixing time, and not 125 RPMs.
1251 | 76NYA-Room 360 Dept HAFL-Balance F-5674 not in it's approved location and moved without as found data being taken | robert.bauch@sanofi.com | 76-360 | 2023-03-08 | 2023-03-09 | STD-000236 Section 3.8 as founds are to be taken before balance is moved and BQ-GOP-000077 section 5.5.1.4 a deviation is generated for equipment not in their approved location.
1250 | Not enough 200L barrels to completed buffer batch. | maria.palmar@sanofi.com | 8 New York Ave | 2023-03-08 | 2023-03-08 | Prior to staring buffer area clearance should be performed to insure all material for batch are available.
1249 | Cz Zinc Column Stock Creation did not execute | robert.morawski@sanofi.com | 8NYA RM1303 | 2023-03-08 | 2023-03-08 | Stock is successfully created.
1248 | XVF Media with DTT part 760741, Lot DGHAW018 Manual Test request missing Time Submitted | katherine.ditommaso@sanofi.com | 8 New York Ave Solution Prep | 2023-03-07 | 2023-03-08 | Per ALF-SOP-000258 and ALF-FRM-000005 Time Submitted must be recorded on Manual Test Request ed | Yes | No Deviation
1247 | Sartorius 2000L Bag Leak | stanley.durand@sanofi.com | 8NYA ROOM 1503 | 2023-03-07 | 2023-03-07 | Bags are expected to not leak during usage.
1246 | 74NYA room CC!-18-Dept APT CC-F-13857-002 Delta V temperature indicator for -20 C freezer F-13857 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 74-CC1-18 | 2023-03-08 | 2023-03-08 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1245 | ASM Free Thiol Documentation Error for Sample Concentrations | John.Kapaniris@sanofi.com | 68NYA QC Chemistry ICB Lab | 2023-01-21 | 2023-03-08 | Sample Concentrations should be entered correctly for the sample.
1244 | 18 GMP Preventative Maintenance (PM) Work Orders Not Completed by Due Date of 28Feb2023 | Tanni.Arif@sanofi.com | 51NYA, 8NYA, 76NYA | 2023-03-01 | 2023-03-07 | All PMs completed by due date, end of month | Yes | Deviation
1243 | Momentary Power Interruption Occurred on 07Mar2023 | tanni.arif@sanofi.com | Line 1, which feeds 8, 45, 68, 74, 80 | 2023-03-07 | 2023-03-07 | It is expected that electrical power interruption would not impact operations at Framingham Campus. | Yes | Deviation
1242 | SUM Failure | leo.gillespie@sanofi.com | 1350 | 1350-solution prep--mixer PBM-1505001 | 6:50 | following FBL-EWI-000039 SUT fail instructions
1241 | use of uncontrolled paperwork used to document media fill PQ execution and associated environmental monitoring | robert.lavalleejr@sanofi.com | 76 NYA HA purification | 2022-10-15 | 2023-03-02 | batch records for GMP activity are required to be issued as lot-specific documents. other GMP activities (ie; environmental monitoring) are required to be documented on approved procedures
1240 | Admin Audit trail review not performed per schedule | coreen.morin@sanofi.com | N/A | 2023-03-06 | 2023-03-06 | Improved clarity in surrounding execution of audit trail review. |  | Deviation
1239 | 3 EM packets from 45 CTMS do not have verifiable sources for initial incubation of samples | megan.botti-marino@sanofi.com | 4700 NYA, QC Microbiology lab | 2022-12-03 | 2023-03-07 | It is expected that the initial incubation be documented at the time of incubation.
1238 | No Product - No Lot - 68NYA - Incorrect Magnehelic Gauge Certificate Value on BSC F-16080 | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2022-09-01 | 2023-03-06 | Per FBL-SOP-000594 V8.0 step 6.3.4, the magnehelic gauge is to be within the acceptable range for the BSC and document on FBL-LBK-000202. | Yes | Deviation
1237 | Fz Butyl Load not forward processed within 24 hours | nathaniel.barnes@sanofi.com | 8NYA 1323 | 2023-03-06 | 2023-03-06 | Forward processing decision for Butyl Load
1236 | HA Lot - DBYFRM011; Steps not executed. | charlie.houle@sanofi.com | 76 NYA; Room 364 (HA Fermentation) | 2023-03-01 | 2023-03-06 | Expected results would be to have the steps executed during processing.
1235 | Unacceptable Zero filter obtained prior to routine EM | Craig.Stevenson@sanofi.com | 76 New York Ave | 2023-02-02 | 2023-03-05 | Per FBL-SOP-001039 V5.0 Step 6.4.8, the fifth Zero Count Filter Check reading for channels (0.5µm and 5.0µm) should be 0 when sampling a Grade A environment.
1234 | No Product - Not Lots - 76NYA - Missing Zero Filter | Craig.Stevenson@sanofi.com | 76 New York Ave, Room 3104 | 2023-02-17 | 2023-03-05 | Per FBL-SOP-001039 V5.0 step 6.4.1, a zero filter check should be performed prior to the day’s sampling event at the building being monitored.
1233 | 76NYA-Room 347-Facilities TOC analyzer on R&D loop F-11534 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-347 | 2023-02-13 | 2023-03-06 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1232 | 68NYA-Metrology-no product impact-Metrology Standard S-1174-001 and S-1174-004 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2023-02-23 | 2023-03-06 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1231 | HiHi SUB Temperature Alarm/Hold on X2200/TIC-026 (F008-TIC-2200026) | dana.hopkins@sanofi.com | 1333 | 2023-03-05 | 2023-03-05 | SUB temperature should stay between 36-38 degrees C.
1230 | Fz Butyl Set up Recipe needs to be corrected | robert.morawski@sanofi.com | 8NYA RM 1323 | 2023-03-05 | 2023-03-05 | The record worked correctly, but through process knowledge it was realized there is a gap in the process during PMs. A cappa will be needed to ensure this doesn't not happen again.
1229 | Conductivity meter incorrectly standardized resulting in EBR issues | christopher.arnold@sanofi.com | 8RM1350 | 2023-03-04 | 2023-03-04 | Per FBL-EWI-000022, it is expected that the meter will be set to two decimal places to allow the EBR to function correctly.
1228 | 8NYA-Fz-X-2200 not able to execute EWI instructions for pH probe configuration | john.lytle@sanofi.com | Rm1333 | 2023-03-04 | 2023-03-04 | EWI instructions accurate to task being performed, references to menu and configuration settings should match what is seen at the local faceplate
1227 | Portable media storage and usage form missing storage information | adam.rowe@sanofi.com | 51NYA, room 147 | 2023-01-29 | 2023-03-04 | Step should be signed
1226 | 8NYA-Fz-X-2200 not able to execute EWI instructions for DO probe configuration | john.lytle@sanofi.com | Rm1333 | 2023-03-04 | 2023-03-04 | EWI instructions accurate to task being performed, references to menu and configuration settings should match what is seen at the local faceplate
1225 | Annual Compressed Air site 74-CA-04 not sampled in 2022 | Craig.Stevenson@sanofi.com | 76 NYA / Room 347 | 2022-12-31 | 2023-03-03 | Per FBL-SOP-000366 V13.0 Attachment 1 (page 11 of 17), 74-CA-04 should be sampled annually.
1224 | Vitry user i0229454 (Amhed El Hadjam) received Framingham sample 480829. | brandon.lebherz@sanofi.com | 68 NYA, QC Micro Lab, | 2023-03-01 | 2023-03-02 | Is is expected that users from other sites are not able to receive Framingham data.
1223 | HA lot DBYFRM012 offline pH out of range step 7.11.3.3. FBS-MBR-000058 | dianne.moustafa@sanofi.com | 76 NYA/364 | 2023-03-03 | 2023-03-03 | Offline pH be with range of 0.3 of PLC pH
1222 | Omnix Vision Inspection System PC Workstation not working | alan.rousseau@sanofi.com | Room 1009 74 NYA | 2023-03-03 | 2023-03-03 | FBS-SOP-000509
1221 | Did not complete signature in ilearn for revision for FBL-SOP-000595 before performing testing | vanessa.jaskula@sanofi.com | 68NYA QC Micro | 2023-02-25 | 2023-03-02 | SOP revision in ilearn needs to be signed prior to assay
1220 | QCM Incubator FB-2054(75-QC-0206) setpoint value set at 29.2C. Expected setpoint value is 29.0C | brandon.lebherz@sanofi.com | 68 NYA, QC Micro Lab, Room 4760 | 2023-03-01 | 2023-03-01 | Per Study# 21-0404 (PQ) / F068-ED-UR-0008 User Requirement (UR) set point was qualified to be 29.0C
1219 | 45NYA - 172 POU Cooler not functional | cindy.guerra@sanofi.com | 45NYA / Rm 172 | 2023-03-02 | 2023-03-02 | POU Cooler would be operable for MFG use
1218 | Suspected roof leak in RM 140 at 51 NYA. | adam.hildebrand@sanofi.com | RM 140 at 51 NYA | 2023-03-02 | 2023-03-02 | It is expected that the ceiling tiles and roof are integral.
1217 | Guidance missing for labeling and storage of WFI spray bottles used for testing and KCL used for pH probe storage. 8NYA/Non-product specific | aidan.mula@sanofi.com | 8NYA | 2023-03-02 | 2023-03-02 | Analysis of storage, labeling, and expiration of these materials and guidance in SOP/EWI. Further determination at Gemba. | Yes | No Deviation
1216 | TSK CGF1B14A fraction 6 PI/ rhTSH seems to be atypical from TSK fractions pooling Table | sylvester.kegbeh@sanofi.com | 51 NYA ROOMS 132 | 2023-02-28 | 2023-03-01 | FBL-MBR-000057, FBL-SOP-000057
1215 | ASM DS Stability Activity Samples failed Assay Acceptance Criteria, Substrate was not clear/free of particulates | John.Kapaniris@sanofi.com | 68NYA | 2023-01-20 | 2023-02-28 | Substrate should be clear | Yes | Deviation
1214 | UF\DF CIP with CV recipe issues | Christian.Venegas@Sanofi.com | 8NYA Downstream suite 1 | 2023-03-01 | 2023-03-01 | To be able to sync the status of the correct state for UF\DF membranes
1213 | Buffer tote leaked, pt# 804998 from the pompe MMC line | dennis.bazinet@sanofi.com | Room 1212, pompe MMC line | 2023-03-01 | 2023-03-01 | All tote connections that are seal should also have a clamp on the line
1212 | CIP Cycle Development of the Zinc Chromatography Column C-3112020 at 8NYA Run 1 Cerezyme, swab 1 from lot RF638D004_003 #1021 had a final reported result of 2.5ppm. Per FBL-SOP-000484, this result falls outside of the acceptable range of 0.76 to 2.4ppm. | paul.consolmagno@sanofi.com | 68 NYA, QCM | 2023-02-24 | 2023-02-24 | TOC swabs would generate results within range
1211 | 8NYA-CC8 room 1333-product impact unknown MCC skid X-2560 (33 items) and MCC Skid X-2543 UV (6 items) overdue for calibration | robert.bauch@sanofi.com | 8-1333 | 2023-03-01 | 2023-03-01 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 units should have been calibrated before their due date.
1210 | 51 NYA- Thyrogen- Bioreactor (CGF4B10) pH probe adjustments made but not required per MBR. | ian.odell@sanofi.com | 51 NYA Room 140 | 2022-08-03 | 2023-02-28 | Per FBL-MBR-00091, step 4.4.10 adjust pH probes on the bioreactor to BGA reading per FBL-SOP-000680 if difference of either probe is >0.05 or as needed.
1209 | DBYFRM011 PM5 Bottle inoculated prior to minimum 14.5 hour equilibration time | nicholas.macdonald@sanofi.com | 76-364 | 2023-02-28 | 2023-02-28 | It is expected that the PM5 equilibrate for the minimum 14.5 hours.
1208 | DBYFRM010 PM5 Bottle OD Exceeded allowable range | nicholas.macdonald@sanofi.com | 76-3104 | 2023-02-27 | 2023-02-27 | It is expected that the OD reading be within the allowable range of 0.200 - 0.400
1207 | FZ 20L Wave CO2 Flow Inconsistency | michael.knapp@sanofi.com | 8NYA 1333 | 2023-02-27 | 2023-02-27 | CO2 flow levels are steady.
1206 | 68NYA-Metrology-no product impact-Metrology Standard FBS-0101-001 had no as found data taken during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2023-02-14 | 2023-02-28 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have had as found data taken
1205 | 68NYA-Metrology-no product impact-Metrology Standard S-1337-005 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2023-02-16 | 2023-02-28 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1204 | Investigate low pressure pump R&D WFI loop/ pump shut down | MACJONES.OSEIKUFFOUR@SANOFI.COM | BUILDING 76NYA ROOM 347 | 2023-02-28 | 2023-02-28 | FBL-SOP-000870
1203 | Fz UF cassette hydraulic pressure exceeded 2000 psig | michael.daly@sanofi.com | 8 NYA Room 1323 | 2023-02-11 | 2023-02-27 | It is expected the hydraulic pressure of the cassettes is between 1600-1800 psig | No | No Deviation
1202 | Cerezyme Zinc Column failed CV | stanley.durand@sanofi.com | 8 NYA room 1303 | 2023-02-17 | 2023-02-27 | It is expected for swab final results to fall within range per FBL-SOP-000484
1201 | MCC Skid X-2543 Status State Errors | michael.knapp@sanofi.com | 8NYA 1333 | 2023-02-04 | 2023-02-26 | System status changes allow for proper scanning and identifying of equipment. | Yes | No Deviation
1200 | ASM DS Stability Activity Samples failed Assay Acceptance Criteria, Samples outside the testing window. | john.kapaniris@sanofi.com | QC Chemistry 68 New York Avenue ICB Lab | 2023-01-13 | 2023-02-27 | Blank wells Must be ≥ -0.005 and ≤ +0.005 at 415 nm.
1199 | Two No Reportable results for Dynamic EM for Batch 202204535-1 | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2023-01-04 | 2023-02-27 | Per FBL-SOP-001496, all results should be counted and entered into LIMS to generate complete results.
1198 | No Product - No Lots - 68NYA - Incorrect EMS Parameters for Incubator FB-2057 (75-QC-0201) | Craig.Stevenson@sanofi.com | 68NYA / Room 4760 / QCM | 2023-02-27 | 2023-02-27 | It is expected that the OQ and EMS documentation align with the PR and UR documents.
1197 | Higfh high alarm limits set to 0.0PSIG on Clean steam pressure controller (PIC-06304) causing recipe to hold at start. Agitator speed controller not responding during recipe at start (SIC07304). | adam.hildebrand@sanofi.com | 51 NYA | 2023-02-27 | 2023-02-26 | It is expected that the recipe would load correct setpoints for the operation at hand.
1196 | Item in EBR BOM not able to be used. | alex.tello@sanofi.com | Room 1350 | 2023-02-27 | 2023-02-27 | Path forward
1195 | Agitator did not turn on during Fz Zinc Eluate Mix Delta V Recipe | stanley.durand@sanofi.com | 8NYA rm 1323 | 2023-02-26 | 2023-02-26 | Agitator is expected to turn on for 15 min mix per FBL-EWI-000033.
1194 | MCC Rinsate Samples X-2510 Product outlet not collected during Post use CIP | Yves.Antoine@sanofi.com | Room 1333 | 2023-02-26 | 2023-02-26 | Proudcut waste samples must be first sample taking from MCC Skid.
1193 | Electrical safety box removed from wall in corridor 83 in building 51 | charles.hart@sanofi.com | 51NYA Corridor 83 | 2023-02-26 | 2023-02-26 | N/A
1192 | high high alarm on X2210 weight controller/High high alarm on X2210 weight Controller | marcelo.dossantos2@sanofi.com | 1330 | 2023-02-24 | 2023-02-24 | HiHi alarm should not occur.
1191 | 51NYA Thyrogen 2300004326 Maxisette Retentate Equilibration CGF2B17A PR-065-17 tested with expired reagent for BIOB | larissa.fairweather@sanofi.com | 68 NYA / 4700/ QCM | 2023-02-23 | 2023-02-23 | ALF-MTD-000018 V.12.0 Step 9.2.2.2. Obtain all applicable materials needed for testing as listed in Section 7 / (Section 7) 7.1. Critical Reagents: Sterile, peptone (fluid A) pH 7.1 ± 0.2, BD Catalog #290652 or equivalent (FRAM_RINSE_FLUID) | Yes | Deviation
1190 | Non-viable data for Dynamic operations missing in Onelims system | Adriana.Alban@sanofi.com | 51NYA/152A | 2022-12-27 | 2023-02-24 | It is expected that the non-viable data is parsed properly into Onelims after sampling event and checked by operator per FBL-SOP-000207.
1189 | OneLIMS LAL sample 560133 CGHHE023 COV_AX_827589_POU810991 was not received and put on test in QCM | paul.consolmagno@sanofi.com | 68 NYA, QCM, Sample receipt | 2022-11-19 | 2023-02-24 | The endotoxin sample would be received and put on test in the QCM laboratory | Yes | Deviation
1188 | Cell Count Data Incorrectly Input into Batch Record | james.letourneux@sanofi.com | 51RM152 | 2022-12-12 | 2023-02-23 | The 275 counted viable cells should have been reported both in the Total Cells field as well as the Viable Cells field. Non-viable cell count should not have been recorded.
1187 | Manual launching of MCC blank run/Blank Run 1 Recipe not launching from MES | marcelo.dossantos2@sanofi.com | 1330 | 2023-02-24 | 2023-02-24 | Recipe launch manually
1186 | Bulk Gas sampling of allston bulk delivery truck | vincent.palumbo@sanofi.com | 80 NYA | 2023-02-23 | 2023-02-23 | per NBO-SOP-000122, step 6.3.1 and 6.3.2 the bill of lading must be verified before sampling and testing occurs
1185 | 8NYA Inoc. Prep EM Plates DGHSD002 expired upon delivery | cleiton.lopes@sanofi.com | 8NYA and 68 Sample Management | 2023-02-23 | 2023-02-23 | EM must be incubated within 24hr of sampling.
1184 | 8NYA-Fz-pressure event on Harvest line, recipe held. Leak identified at Pendotech connections post event | john.lytle@sanofi.com | Rm 13333 | 2023-02-22 | 2023-02-22 | NO pressure event
1183 | 76NYA-Room 3103 Dept HABP-Vacuum Indicator F-6938-027 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-3103 | 2023-02-22 | 2023-02-23 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1182 | BIRF form missing initials and date | adam.rowe@sanofi.com | 51NYA, room 148 | 2023-01-22 | 2023-02-22 | All movements should be signed for |  | Deviation
1181 | 8 NYA Utility Monitoring Inconsistencies | Brendan.Scanlon@sanofi.com | 8 NYA | 2023-02-22 | 2023-02-22 | FBL-FRM-000819 would align with sampling determined by HACCP and FRM document would be accurate.
1180 | Media not placed in cold room. | stephen.roberts2@sanofi.com | 51NYA/108 | 2023-02-21 | 2023-02-22 | It is expected that the media barrel would have been transferred to cold room 108C with the rest of the lot and stored at 2-10C
1179 | 8NYA Anaerobic EOR bioburden sample not tested for routine anaerobic bioburden per FRAD22A0215/2022FRACC0043/FBL-PTM-000293 | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2023-02-18 | 2023-02-21 | Anaerobic bioburden samples are expected to arrive to QC via non-routine MLC sample pick up to allow for testing with communication from MFG. Samples are expected to be received and placed in appropriate testing bin. It is expected that sample receipt verify the labels against the test request upon receipt | Yes | Deviation
1178 | No product - No Lot - 68NYA - No Pass/Fail BSC Gauge reading recorded on 11Feb2023 and 16Feb2023 | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2023-02-11 | 2023-02-21 | Per FBL-SOP-000594 V8.0 step 6.3.4, Ensure magnehelic gauge is within the acceptable range for the BSC and document on FBL-LBK-000202.
1177 | 8 NYA GEMBA- Buffer Stock Incorrectly taken out in MES | cory.lefort@sanofi.com | 8 NYA | 2023-02-21 | 2023-02-21 | FBL-SOP-001320 | Yes
1176 | Handling units for SUT are referenced in two separate batch records. 8NYA/Cerezyme/RM1350 | aidan.mula@sanofi.com | 8NYA, RM1350 | 2023-02-21 | 2023-02-21 | TBD at Gemba
1175 | Hyaluronic Acid - DBYHAS001 - 76NYA - Two no reportable data sites and failed zero filter | Craig.Stevenson@sanofi.com | 76 New York Ave, Room 355 | 2023-01-24 | 2023-02-21 | See Gemba email
1174 | Temperature excursion on Cerezyme DGHDT010 per TempTale data | Steven.Servais@sanofi.com | 80 New York Av. Framingham, MA. | 2023-02-15 | 2023-02-17 | Temperature should be from 2-8 DegC
1173 | hTSH, Missing time entry on Cold Room 109 Attachment 1 of FBL-SOP-000790 | brian.mcgee@sanofi.com | 51NYA/Rm 109 | 2023-02-15 | 2023-02-20 | It is expected for the logbook entry to be completed fully prior to signing |  | Deviation
1172 | 8NYA-Fz-MCC X-2543 experiencing hold conditions due to Load Controller Low Low alarms. After automation support, additional support requested | john.lytle@sanofi.com | Rm1333 | 2023-02-18 | 2023-02-18 | Load flow is within processing limits (low low 6.7rpm)
1171 | 8NYA-Fz-X2210 pC02 probe 1 deviating from probe 2 indicating potentially inaccurate reading | john.lytle@sanofi.com | Rm1333 | 2023-02-18 | 2023-02-18 | probe 1 and 2 should have similar readings as they are used in tandem for an average mmHg
1170 | 8NYA-Fabrazyme- Primary vent filter on X-2210 has potential contact with foam after doubling perfusion rate | john.lytle@sanofi.com | Rm1333, 8NYA | 2023-02-18 | 2023-02-18 | it is intended that foam does not make contact with the sterile filter
1169 | Cz Media Tote Possible Leak | michael.knapp@sanofi.com | 8NYA 1403 | 2023-02-17 | 2023-02-17 | Media tote shows no evidence of a leak.
1168 | CNC Upstream activities being performed by operators not signed off | robert.morawski@sanofi.com | 8NYA CNC | 2023-02-17 | 2023-02-17 | CNC operators should have just been assisting with these operations, or performing with a trainer. | No | Deviation
1167 | Missed weekly EM site 226 (LFH) 27Nov2022 – 03Dec2022 | Craig.Stevenson@sanofi.com | 76 New York Ave, Room 364 | 2022-12-03 | 2023-02-17 | Per FBL-SOP-000941 V13.0, Step 6.2.2, all EM sites should be collected at their required frequency. | Yes | Deviation
1166 | SQRP report interval missed in 2022 per BQ-GOP-00274 | joshua.gray@sanofi.com | Not Applicable | 2022-12-31 | 2023-02-16 | Per BQ-GOP-000274 step 5.5.1.2, Site Quality Risk Profile Reports are written and documented in Geode+ twice yearly (at a minimum). The site risk report will be approved by the site head, senior operational management, and senior quality management.
1165 | Particles found in stability aliquots | alexander.decoste@sanofi.com | 68 NYA, Room 4790 (bioburden room), QC Micro | 2023-02-16 | 2023-02-16 | Foreign objects should not be present in this sample. Could not find any applicable passages as our appearance test only checks for color and clarity per FBL-MTD-000028. | Yes
1164 | 74NYA-Room PU1-24-Dept PF74-FB-3663 (WS-3201) and FB-3664 (WS-3202) weight indicators Found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 74-PU1-24 | 2023-02-16 | 2023-02-16 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results |  | Deviation
1163 | TSA lot information not recorded for dynamic EM during an HA fermentation operation | megan.botti-marino@sanofi.com | 76 NYA | 2023-01-24 | 2023-02-15 | It is expected that the information of all materials used during a process be recorded on the associated documentation.
1162 | Incorrect Pharmacopia grade listed in Cz Dossier | Sarah.Francis@sanofi.com | N/A | Unsure | 2023-02-06 | VV-QUAL-0659179, section 3.2.S.2.3
1161 | 8 NYA, all CNC rooms were not configured correctly in MES to complete the b-monthly facilities cleaning | timothy.hoard@sanofi.com | 8 New York Ave / CNC areas | unknown | 2023-02-15 | It is expected that the cleaners are prompted to perform the appropriate cleanings in MES furthermore the facility cleanings should be performed and documented in MES per FBL-EWI-000005.
1160 | 705672 buffer for Fz Butyl out of usages, and tubing stuck in satorious welder | robert.morawski@sanofi.com | 8NYA CNC | 2023-02-15 | 2023-02-15 | Welder does not get stuck, and welds successfully.
1159 | 45/51NYA plan Materials-Thryogen dataloggers for shippers FB-4523-005, FB-4523-006, FB-4524-005 and FB-4524-006 missing | robert.bauch@sanofi.com | 80NYA | 2022-11-21 | 2023-02-07 | BQ-GOP-000248 Chart 7
1158 | Metaflex analyzer would not process sample in Rm 140 and RM 152 | adam.hildebrand@sanofi.com | 51 NYA RM 140, RM 152 | 2023-02-15 | 928 | It is expected that the analyzer processes the sample accordingly. |  | No Deviation
1157 | Harvest Goods receipt labels printed with no barcode or HU identification. | adam.hildebrand@sanofi.com | 51 NYA | 2023-02-10 | 2023-02-13 | It is expected the bar code and corresponding HU number would print |  | No Deviation
1156 | OneLIMS lots 16245 and 16246 (Falcon Bioreactor) are unable to be dispositioned by QC | valerie.staras@sanofi.com | 74 NYA | 2022-03-17 | 2023-02-14 | It is expected that the lots can be dispositioned
1155 | 74NYA-Room PU1-22-Dept APT SS- F-17756-006 Weight Transmitter and F-17756-007 Delta V Weight Indicator found non-operational, no as found data taken. | robert.bauch@sanofi.com | 74-PU1-22 | 2023-02-14 | 2023-02-14 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1154 | Cracked filter leading to incomplete SUM bag inflation | Richard.Manasian@sanofi.com | 8 NYA room 1350 | 2023-02-14 | 2023-02-14 | Per FBL-EWI-000019, it is expected that inflation will be complete allowing for batching of the buffer.
1153 | Daily LAL water was collected and missing | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2023-02-11 | 2023-02-13 | For the sample to be placed in the "to be tested" bin. | Yes | Deviation
1152 | Nano Load Sum ran empty during buffer chase | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-02-13 | 2023-02-13 | Sum should not run empty during chase.
1151 | Cz Harvest Guard Filter Pressure Alarm | michael.knapp@sanofi.com | 8NYA 1403 | 2023-02-13 | 2023-02-13 | Guard filter does not foul. | Yes | No Deviation
1150 | 8NYA Upstream Harvest SUM TCU alarms | michael.knapp@sanofi.com | 8NYA 1403 | 2023-02-13 | 2023-02-13 | TCU remains operational.
1149 | MCC skid x2543. hold due to eluate SUM HIHI temp alarm | Yves.Antoine@sanofi.com | Room 1333 | 2023-02-11 | 2023-02-11 | FBL-EWI-000032
1148 | hTSH Room 152A was observed to have had an electrical fire at the outlet behind staging area table overnight on 11FEB23 | brian.mcgee@sanofi.com | 51NYA/Rm 152A | 2023-02-11 | 2023-02-11 | It is expected for the room integrity to remain intact during all operations.
1147 | SUB X2220 pH probe 2 behaving erratically | caleb.coppola@sanofi.com | 8NYA Rm 1403 | 2023-02-11 | 2023-02-11 | per FBL-CST-000029 Table 7.8 pH is a KPP with PAR of 6.8-7.0 excursions less than or equal to 6 hours allowed | No | Deviation
1146 | Pump Tubing leak during Zinc Filtration | nathaniel.barnes@sanofi.com | 8 NYA 1323 | 2023-02-10 | 2023-02-10 | Continue filtration of Zinc eluate with new SUT
1145 | Recipe went into hold due to air in the bubble trap. | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1323 | 2023-02-10 | 2023-02-10 | its expected that load will pause once the recipe sees less than 5-7
1144 | Not able to find Portable Media storage and usage form. | steven.servais@sanofi.com | 80 New York Av. Framingham, MA. | 2023-02-10 | 2023-02-10 | Per FBL-SOP-000802 6.3.1 When a media vessel or media bag is moved from one storage location to another, the operator will record the date, time and initials/date on the Removed line of ATTACHMENT 1.
1143 | ASM Stability studies frozen in error upon receipt | cleiton.lopes@sanofi.com | 68 QC ICB Chem lab | 2023-02-09 | 2023-02-10 | Stability samples are expected to be stored per Stability storage requirement.
1142 | No Lots - Cleaning Log (FBS-SOP-000664 Attachment 10) missing information | charlie.houle@sanofi.com | 76 NYA; Room 364 (HA Fermentation) | 2023-02-03 | 2023-02-09 | Per FBS-SOP-000664, section 6.3, a monthly cleaning must be performed once per calendar month for each manufacturing area and documented on the appropriate attachment.
1141 | Fz MCC HBOD Waste Line to top tote Leak 2-3 Liters, spill procedure followed | michael.knapp@sanofi.com | 8NYA CNC | 2023-02-09 | 2023-02-09 | There is sufficient space in the waste totes for operation. |  | Deviation
1140 | R&D media inadvertently discarded during the decommissioning activities for Falcon vaccine. | michael.montiverdi@sanofi.com | 74 CR #164 | 2023-02-08 | 2023-02-08 | Proper materials to be discarded.
1139 | Cz ATF Sample Line Failed Weld | michael.knapp@sanofi.com | 8NYA 1403 | 2023-02-09 | 2023-02-09 | Welder works correctly. |  | Deviation
1138 | 8NYA-PF8-Room 1703-FB-4080-007 weight indicator was found out of tolerance during an unscheduled calibration due to equipment change over | robert.bauch@sanofi.com | 8-1703 | 2023-02-09 | 2023-02-09 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results |  | Deviation
1137 | Canada IFU MFG received does not match the NBO-SPC | alan.rousseau@sanofi.com | room 1017 74 NYA | 2023-02-09 | 2023-02-09 | Expected result is that Sanofi SPC document would match the Baxter IFU p/n
1136 | RF613C010 part 713972 not used during processing | richard.manasian@sanofi.com | 1350 at 8 NYA | 2022-11-19 | 2023-02-08 | Part 713972 would be used as required in the BOM
1135 | Fz SP Chrom Skid X3031 leaking from triclamp connection during Setup Recipe. | robert.morawski@sanofi.com | 1323 | 2023-02-08 | 2023-02-08 | Fz SP X3031 doesn't leak while running.
1134 | 51 NYA UHU 7 unexpectedly shutdown at approximately 0120 | Alex.Biskaduros@Sanofi.com | 51 NYA room 149 to 84 and additional rooms | 2023-02-08 | 2023-02-08 | FBL-SOP-000316 alarm response
1133 | 45NYA Room 169A Dept Facilities Engineering-F-15836-023 Conductivity Indicator on parts washer found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 45-169A | 2023-02-07 | 2023-02-08 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1132 | Parts referenced in record cannot be traced. | bailey.gleason@sanofi.com | 8NYA RM1350 | 2023-02-07 | 2023-02-07 | Hold Gemba to see path forward.
1131 | hTSH DGFSB002 CO2 BGA result was out of range on 12L D1, 07FEB23 | brian.mcgee@sanofi.com | 51NYA/Rm 152A | 2023-02-07 | 2023-02-07 | It is expected for BGA data from daily sampling to be within the acceptable ranges.
1130 | 68NYA Room 4700 Dept QCM68-TOC's FB-2424 and FB-2425 were found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 68-4700 | 2023-01-30 | 2023-02-07 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1129 | 74NYA Room 1021 Dept SPRA-F-6261 Magnehelic Gauge found out of tolerance during a scheduled Calibration. | robert.bauch@sanofi.com | 74-1021 | 2023-02-01 | 2023-02-07 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1128 | 17 Late PMs for Jan2023 | tanni.arif@sanofi.com | Sitewide | 2023-02-01 | 2023-02-06 | Complete all PM tasks prior to PM closure (otherwise approved by QA) or a deviation is required per site specific SOP
1127 | Incorrect EM program selected during Formulation process for CGF2B11 | sylvester.kegbeh@sanofi.com | 51 NYA ROOMS 88 and 87 | 2023-02-06 | 2023-02-06 | It is expected that the program NV_GAS be used during EM for formulation per FBL-S0P-000755.
1126 | HTSH: Small drip detected from clamp during media additon for refeed of CGF2B17 on H9, 06FEB23 | brian.mcgee@sanofi.com | 51NYA/Rm 140 | 2023-02-06 | 2023-02-06 | It is expected for the connections to be intact without any drips or leaks during media addition to the reactor.
1125 | MAHU 1 and 2 at 45 NYA shutdown | patrick.hosegood@sanofi.com | 45 NYA | 2023-02-06 | 2023-02-06 | System should have kept running
1124 | Bioburden EM Swabs from Tested on R2A instead of TSA | kyle.lebherz@sanofi.com | Discovered at 68NYA, QCM. Sampling took place in 45, 51. | 2023-01-25 | 2023-02-06 | Per ALF-MTD-000034 Step 7.5.1, swab samples for bioburden are to be tested using TSA as the media. | Yes | Deviation
1123 | MES/SOP Do Not Align with New Standards, SOP Not Complete for New Standards | jeff.hill@sanofi.com | 45RM190 | 2023-02-06 | 2023-02-06 | FBL-MBR-000496
1122 | 8NYA-Fz-Harvest Overflow Pump F008-VSC-25550002 not responding during SUB recipe start | john.lytle@sanofi.com | Rm1333 | 2023-02-04 | 2023-02-04 | During SUB recipe launch, the associated Harvest equipment is acquired in DeltaV. The cGMP equipment was found in a disconnected state, and Mfg has no guiding SOP on initial configuration for this equipment as it is assumed the equipment remain configured for cGMP processing
1121 | CZ Harvest SUM TCU Failure | michael.knapp@sanofi.com | 8NYA 1403 | 2023-02-05 | 2023-02-05 | TCU alarm does not occur
1120 | Missed monthly water samples in 68 and 45 in January | Stellasotheara.long@sanofi.com | 68 NYA, 45 NYA | 2023-02-04 | 2023-02-04 | FBL-SOP-000478. Step 6.14.1 | Yes | Deviation
1119 | FZ Aber Probe Cross-Threaded in Bellows | michael.knapp@sanofi.com | 8NYA 1333 | 2023-02-04 | 2023-02-04 | Probes are installed properly into their bellows and sterile following autoclave cycle.
1118 | SUT transacted into multiple records | stanley.durand@sanofi.com | 8NYA room 1323 | 2023-01-29 | 2023-02-04 | It's expected that SUT is transacted into one correct process order.
1117 | Flaky Particulate in Validation BIs | Rebecca.Schmidt@sanofi.com | QC Micro Counting Bench | 2023-01-27 | 2023-02-04 | Media should clear and free of any particulate or turbidity per ALF-SOP-000186
1116 | EM Plates Batch#: 202300378-1 INC-04516 expired before delivery on 03FEB22 | karen.ramos2@sanofi.com | 68 NYA QC Sample Receipt Rm 4720 | 2023-02-03 | 2023-02-03 | Plates to be delivered and received within adequate time before they expire.
1115 | Hot water reheat leak in 76NYA room. 235 | rebecca.Cunningham@sanofi.com | 76 NYA | 2023-02-03 | 2023-02-03 | No water leaks in controlled space
1114 | Autoclave Kuhlman monthly sanitization not recorded | valeria.moreno@sanofi.com | 76-352 | 2023-01-31 | 2023-02-03 | FBS-SOP-000601, Attachment 5.
1113 | Multiple Inexecutable Biosurgery Documents | rebecca.Cunningham@sanofi.com | 76 NYA | 2023-02-02 | 2023-02-02 | Documents be executable per BQ-GOP-000181 | Yes | Deviation
1112 | Prompt to start Cz thaw using HOFF was not answered | zachary.rickard@sanofi.com | 8NYA/1365 | 2023-02-02 | 2023-02-03 | Once shells are loaded onto the HOFF the prompt will be acknowledged starting the thawing process.
1111 | HA Lot - DBYHAH002 - Operator signed as trainer, when not a qualified trainer | charlie.houle@sanofi.com | 76 NYA; Room 3102 (HA Purificaiton) | 2023-01-31 | 2023-02-02 | Only operators who have completed the Train-the-Trainer curricula and class should be signing as trainer.
1110 | HA Lot - DBYHAH002 - Missing Verification and Witness Steps | charlie.houle@sanofi.com | 76 NYA; Room 3102 (HA Purificaiton) | 2023-01-31 | 2023-02-02 | Per FBL-SOP-000942, verified by steps require signature contemporaneously. Per FBL-SOP-000942, witnessed by steps must be signed at time of execution.
1109 | Low Viability on Fz Culture | marcelo.dossantos2@sanofi.com | 8NYA/1340 | 2023-02-02 | 2023-02-02 | Viability results higher than 80% per FBL-CST-000028
1108 | Bioburden incubation sample started after 24 hours, PS-1476-00 rhTSH harvest fluid | Alyssa.Barrett@sanofi.com | 68 NYA/Room 4700 | 2023-01-25 | 2023-02-02 | Per FBL-SOP-000019 V. 18.0 Attachment 2 Page 1 of 16, the sample has a 24hr expiry, and should be tested within that timeframe.
1107 | Discrepancy between local load cell transmitter and DeltaV weight for top media tote for Fz 8NYA SUB X2200 | Brandon.Steff@sanofi.com | 8NYA CNC | 2023-02-02 | 2023-02-01 | Media tote switches at 50kg to new tote
1106 | HTSH: Leak detected during harvest removal for refeed of CGF2B17 on H5, on 02FEB23 | brian.mcgee@sanofi.com | 51NYA/Rm 140 | 2023-02-02 | 2023-02-02 | It is expected for there to be no leaks during the transfer of harvest material from the reactor to the harvest bag.
1105 | Identified at QC that Bioburden and Endotoxin Sample Never Received into QC Micro | kyle.lebherz@sanofi.com | 68NYA QCM Lab | 2022-10-10 | 2023-02-01 | For these samples to be received into QC Micro and tested.
1104 | 3 Run PQ was not performed after change from qualified state for Seprafilm Dose Mapping | david.zhong@sanofi.com | 76NYA, | 2023-02-02 | 2023-02-02 | 3 run PQs are required
1103 | 51 NYA - Thyrogen - DGFSB002 - EM batch 202300373-1 unable to be incubated | kelvin.leo@sanofi.com | QC Micro | 2023-02-02 | 2023-02-02 | EM batch plates to be incubated within 24 hours of sampling
1102 | 76- HA Bulk (DBYHAH002) endotoxin submitted in incorrect container | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2023-02-01 | 2023-02-01 | Per MSAT, batch record states "Collect a 15 mL LAL sample in a pyrogen free tube"
1101 | No reportable Non-Viable EM result for site 1204A-5 on 12Jan2023 | Craig.Stevenson@sanofi.com | NBO Room 1204A | 2023-01-12 | 2023-02-01 | It is expected that the non-viable data is parsed to onelims so that all 3 sampling occasions would be resulted.
1100 | LAL sample repeat test not tested within expiry | Stefanie.Erickson@sanofi.com | QC Micro LAL NEXUS bay | 2022-11-14 | 2023-02-01 | for the sample to be repeat tested.
1099 | Sartorius Welder 2009138 In Use in Mfg, but deactivated in INFOR EAM | sheila.pandolfo@sanofi.com | 8 NYA Room 1403 | 2022-11-17 | 2023-02-01 | It is expected that the equipment would be reactivated after a repair is completed and the equipment is put back into use in Mfg.
1098 | 3 Oxygen Critical Utility sites not tested by QC in 2022 per FBL-SOP-001506 | Cstevenson8130@gmail.com | 45 NYA Room 112 | 2022-12-31 | 2023-02-01 | Per FBL-SOP-001506 V2.0, Attachment 2 (page 2 of 2), the Oxygen sites should be sampled Biannually.
1097 | hTSH Growth media (609616, DGFMP001) was placed in the warm room on 29JAN23 and not documented on the transfer log | brian.mcgee@sanofi.com | 51NYA/Rm 148 | 2023-01-29 | 2023-01-31 | It is expected for all transfers of media to be documented on the applicable log sheet
1096 | Process pH is identified to be out of range during chromatogram review from PI.. | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 138 | 2023-01-28 | 2023-01-31 | It is expected that the process pH be between 4.8-5.6 per FBL-SOP-000788.
1095 | 8NYA-Room 1333 Dept CC*-24 GMP Critical items on MCC Chrome skid X-2510 overdue for calibration | robert.bauch@sanofi.com | 8-1333 | 2023-02-01 | 2023-02-01 | BQ-GOP-000046 Calibration Program Version 5 Section 5.7.5 units should have been calibrated before it's due date
1094 | Unable to create 15kg bench top scale Calibration EBR | robert.morawski@sanofi.com | 8NYA RM1303 | 2023-01-31 | 2023-01-31 | Successfully creating EBR for 15kg bench top scale calibration.
1093 | Thyrogen/Buffer prep - unable to consume P/N 686157 to PO RG012D080 in Athena for DGFB1080/1288 | Jeremy.Wooden@Sanofi.com | 45 NYA/190 | 2023-01-31 | 2023-01-31 | It is expected that the quantity of material weighed out and used for the process could be consumed in Athena. | Yes | No Deviation
1092 | 51 NYA- Thyrogen- E3 Bioreactor QCA for low pH (CGF3B16) | ian.odell@sanofi.com | 51 NYA Room 140 | 2023-01-31 | 2023-01-31 | Per FBL-CST-000007 (step 9.4.4), the pH range, which is classified as a Process Control (PC), is 6.5 to 7.1
1091 | “FBS-MBR-000058, multiple lots, chart records didn't record appropriately the incubator temperature | dena.elmesalamy@sanofi.com | 76NYA-364 | 2023-01-08 | 2023-01-31 | The incubator chart temperature must stayed in range 36 -38 degree Celsius
1090 | Missing Action Notice and CV paperwork | vincent.palumbo@sanofi.com | 74 and 76 NYA | 2023-01-31 | 2023-01-31 | Action Notices were executed, and paperwork was returned to be filed.
1089 | Missing Action Notice and CV paperwork | vincent.palumbo@sanofi.com | 74 NYA | 2023-01-31 | 2023-01-31 | Action Notices were executed, and paperwork was returned to be filed.
1088 | Missing Yeast Extract performance qualification data in NBO release packet and Engineering MBR | michael.purple@sanofi.com | 76 New York Ave | 2023-01-27 | 2023-01-27 | It would be expected that the original Attachment 1 would be in the release package for yeast extract lot 21MN00321 and that a copy of this attachment would also exist with the performance qualification documentation.
1087 | Cz Resin Prep Recipe did not populate in Delta V. | robert.morawski@sanofi.com | 8NYA RM1303 | 2023-01-30 | 2023-01-30 | Recipe makes onto the equipment in Delta V from recipe launch BF. | Yes | No Deviation
1086 | Fz Nanofiltration Setup | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-01-30 | 2023-01-30 | Tubing should fit skid properly. | Yes | No Deviation
1085 | Fz Nanofiltration Setup | zachary.rickard@sanofi.com | 8NYA/1323 | 2023-01-30 | 2023-01-30 | Tubing should fit skid properly. | Yes | No Deviation
1084 | 76-231 Low temp EMS | john.beaule@sanofi.com | 76-231 | 2023-01-27 | 2023-01-27 | FBS-SOP-000747
1083 | Weepage on column 2 (X2560)/X2560 Column 2 Inlet Line Weld Weepage | marcelo.dossantos2@sanofi.com | 8NYA/1403 | 2023-01-29 | 2023-01-29 | welding should have not failed. | No | Deviation
1082 | Agitator seal on SUM bag unseated after initial WFI addition and will not reseal. Cerezyme/ 8NYA/ 760741 DTT media | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-01-28 | 2023-01-28 | Determination at Gemba
1081 | Weld failure on outlet side of column 3/Weld Failure | marcelo.dossantos2@sanofi.com | 8NYA/1403 | 2023-01-27 | 2023-01-27 | Lines should weld without failing. | No | Deviation
1080 | Temperature excursion for package containing TSH DP Lot Numbers: DW0011 DW0006 | cleiton.lopes@sanofi.com | 45 NYA Rm1111 | 2023-01-27 | 2023-01-27 | Temperature expected to stay within temperature range of 2-8 during transit.
1079 | 51 NYA- Thyrogen- Data not recorded in MBR from D1 Bioreactor pH probe calibration event (DGF1B01) | ian.odell@sanofi.com | 51 NYA Room 140 | 2023-01-26 | 2023-01-26 | Per FBL-SOP-000718, Step 6.2.2, the mV reading of pH 4 should be recored in FBL-MBR-000093.
1078 | Low level switch alarm on X2560 during buffer prime. | marcelo.dossantos2@sanofi.com | 8NYA/1403 | 2023-01-27 | 2023-01-27 | Recipe should run without hold
1077 | Temptale Alarm for ASM-ICB Rabbit Pyrogen Sample Shipment to Charles River Ireland | michael.searles@sanofi.com | N/A | 2023-01-21 | 2023-01-26 | Samples should be at 2-8C throughout shipment
1076 | IPA Cans left in Mechanical Rooms located in 8NYA, 74NYA & 76 NYA | brizaida.medina-vega@sanofi.com | 74 NYA, 76 NYA & 8 NYA | 2023-01-26 | 2023-01-26 | No trash (including water), n is prohibited in all controlled areas which include the production, testing, warehouse areas and controlled support areas (e.g. utility mechanical rooms, record control). Expired cans must be properly disposed.
1075 | 68NYA-room 4700 QCM-product impact unknown-FB-3494 Thermometer found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-4700 | 2023-01-20 | 2023-01-27 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
1074 | Capacitance probes on SUB X2220 (8NYA) exceeded deviation alarm limits | Brandon.Steff@sanofi.com | 8RM1403 8NYA | 2023-01-27 | 2023-01-27 | Capacitance probes remain withing 20pF of each other | Yes | Deviation
1073 | 51NYA-room 129-APT SS-Product impact unknown- FB-0577 Pressure Gauge was found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 51-129 | 2023-01-26 | 2023-01-27 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1072 | Cz MCC Hex Glycol buffer connection failure | michael.knapp@sanofi.com | 8NYA 1403 | 2023-01-26 | 2023-01-26 | the 1" line on the hex glycol bag is weldable.
1071 | FBL-FRM-000270 Preventative Maintenance Extension Form Error: Failure to Follow GDP (most recent version of form not used) + Incorrect PM Extension Date | dlframbio-f&ecompliance@sanofi.com | 76NYA | 2022-10-28 | 2023-01-25 | FBL-SOP-000942 Step 6.10.1: All documents or records intended for GxP use must be maintained in a controlled quality system; FBL-SOP-000081 Step 6.5.2 TABLE: Maximum allowed extension time is based on frequency of PM
1070 | Inadequate Benchtop equipment (UPLCs) documentation, including PMs missing complete documentation uploaded into the WO, documentation missing sanofi/vendor signatures, approvals, reviews, pagination, blank pages and no WO references. | rondik.ahmed@sanofi.com | 8NYA/1333/IPT | 2021-06-28 | 2023-01-25 | Complete vendor documentation per FBL-SOP-000128
1069 | Bioburden sample 2300002185 rhTSh Harvest Fluid H-05 incubated after expiry | alexander.decoste@sanofi.com | 68 NYA, Room 4790 (bioburden room), QC Micro | 2023-01-25 | 2023-01-25 | Sample should have been tested before 1344 on 25JAN22
1068 | 2021 Lixisenatide (Adlyxin) NBO PQR not written per GOP-000089 | kelly.desousa@sanofi.com | Northborough | 2022-05-01 | 2023-01-20 | Per GOP-000089 Step 5.7, a 2021 NBO PQR would have been written for Adlyxyin | Yes | Deviation
1067 | 8NYA-Room 1333 Dept CC8-product impact unknown-FB-4773-047 and FB-4773-048 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 8-1333 | 2023-01-25 | 2023-01-26 | BQ-GOP-000046 Calibration Program Version 5 Section 5.8.3 unit should have produced passing results
1066 | HA - DBYHAS001 – 76NYA room 355 – 2 sets of HA Fill exit RODACS with 1 sample time | Craig.Stevenson@sanofi.com | 76 New York Ave, Room 355 | 2023-01-24 | 2023-01-25 | Per FBL-SOP-000607, it is expected that there is a start and end time recorded for each set of Exit RODACs.
1065 | Multiple Phase II deviations (Phenix type 13, 14) closed without an investigation and/or without proper QC manager specific approver | claire.fisher@sanofi.com | Multiple Buildings (APUs) | 2021-12-20 | 2023-01-24 | It is expected all Phase II records are processed and approved per BQ-GOP-000249 and BQ-GOP-000248 requirements
1064 | Training - Bioburden and Endotoxin Samples X and y left at incorrect storage temperature for z hours | paul.consolmagno@sanofi.com | 68 NYA, QCM, Sample receipt | 2023-01-24 | 2023-01-25 | Samples would be delivered and received via normal processes per FBL-SOP-000964
1063 | Climet cone cap was on during HA Fill sampling | valeria.moreno@sanofi.com | 76-358 | 2023-01-24 | 2023-01-24 | FBS-SOP-000607, step 6.3.4
1062 | HA Lot - DBYHAH001 - Load Cells on V-3532 | charlie.houle@sanofi.com | 76 NYA; Room 3102 (HA Purificaiton) | 2023-01-25 | 2023-01-25 | Per FBS-MBR-000056 step 7.8.15, we must tare V-3532. The expected result would be an accurate tare of the system to ensure correct transfer amount.
1061 | Conductivity sample tested at incorrect temperature | Jeremy.Wooden@Sanofi.com | 45 NYA Room 190 | 2023-01-15 | 2023-01-24 | Per FBL-MBR-000027 step 4.30 the temperature needs to be between 24.9 and 25.1 degrees C.
1060 | GEMBA- 8NYA- Expiration on SP low pH exposure time | cory.lefort@sanofi.com | 8 NYA Downstream / 1323 | 2023-01-24 | 2023-01-24 | The expected results were that SP process would of ran within the low pH exposure time.
1059 | TSH lot CGF2B17, QCA alarm for DO control caused during T1 refeed on D2 at 51NYA | brian.mcgee@sanofi.com | 51NYA/Rm 140 | 2023-01-24 | 2023-01-24 | It is expected for the 02 sparge condensate valve to be left closed during the refeed SIP.
1058 | Missing Stop signs for Action Notices at 51NYA | daniel.bati@sanofi.com | IN QA - RM room | 2021-09-02 | 2023-01-24 | Scheduled a Gemba
1057 | TSH CGF3B16 pH probe A faulty reading for E3 causing QCA | brian.mcgee@sanofi.com | 51NYA/Rm 140 | 2023-01-24 | 2023-01-24 | It is expected for the pH probes to continue to read accurately during the entire run. |  | No Deviation
1056 | Triple Labeled Bioburden Sample | joseph.napoli@sanofi.com | 68NYA/QCM | 2023-01-24 | 2023-01-24 | for there to be only two labels on the bioburden bag, with no endotoxin labels
1055 | Lost NaOH due to pump Asset #9 damaging the tubbing and producing a leak. | dena.elmesalamy@sanofi.com | 76NYA-364 | 2023-01-24 | 2023-01-24 | Pump #9 is to run properly.
1054 | GEMBA- 8 NYA: Failed Swab#3 CV results for MMC Column | cory.lefort@sanofi.com | 8 NYA | 2023-01-22 | 2023-01-22 | FBL-EWI-000016 and FBL-FRM-000776 |  | No Deviation
1053 | Magnehelic Gauge 45 NYA Rooms 189B to 188 DP Out Of Range | jeff.hill@sanofi.com | 45 NYA 189 to 188 | 2023-01-24 | 2023-01-24 | FBL-SOP-000915, FBL-LBK-000019
1052 | No reportable data will be generated for Fz Endotxoin BioLims# 612709, P/N 762729, Lot# DGHBE003 | paul.consolmagno@sanofi.com | 68 NYA, QCM, Sample receipt | 2023-01-21 | 2023-01-23 | BioLims# 612709 would be stored at 2-8C
1051 | GEMBA- 8 NYA: MMC bottom bed supports missing bolts. | cory.lefort@sanofi.com | 8 NYA | 2023-01-20 | 2023-01-23 | FBL-EWI-000016
1050 | Outstanding old QCAs are not acknowledged on time. | Subhashini.Satishkumar@Sanofi.com | 45NTA, 51NYA, 76NYA, 80NYA | 2023-01-23 | 2023-01-23 | It is expected that all QCAs are unacknowledged, and comments made on a timely manner per FB-SOP-000316
1049 | CZ SUB Perfusion/ATFs Not Starting | michael.knapp@sanofi.com | 8NYA 1403 | 2023-01-22 | 2023-01-22 | ATFs operate correctly when perfusion is initiated.
1048 | Thyrogen-CGF3B16-51NYA- Daily Offline BGA Sample Result for pCO2 Outside of Range | ian.odell@sanofi.com | 51 NYA Room 140 | 2023-01-22 | 2023-01-22 | Per FBL-MBR-00091 step 4.25.8, this must be 100-230mmHg
1047 | Leak on SUM bag during initial WFI fill from clamp | richard.manasian@sanofi.com | 1350 at 8 NYA | 2023-01-22 | 2023-01-22 | WIF fill would complete without leaking from the clamp
1046 | Thyrogen-CGF3B16-51NYA- Daily Offline BGA Sample Result for pCO2 Outside of Range | ian.odell@sanofi.com | 51 NYA Room 140 | 2023-01-21 | 2023-01-21 | Per FBL-MBR-00091 step 4.25.8, this must be 100-230mmHg
1045 | 8NYA-Fz-Viability results do not meet criteria after scale up to 500mL flask stage | john.lytle@sanofi.com | 8NYA- Rm1303 | 2023-01-21 | 2023-01-21 | Minimum Viability after scale up to 500mL stage is 80%
1044 | CZ 761217 SUB D0 Samples Not Collected | michael.knapp@sanofi.com | 8NYA 1403 | 2023-01-18 | 2023-01-20 | Day 0 samples are collected
1043 | Vicell Allocated after going Weekly Clean Expired 18DEC22 | michael.knapp@sanofi.com | 8NYA 1343 | 2022-12-18 | 2023-01-20 | Vicell is weekly clean during use.
1042 | Warm Room 153 Acknowledged Alarm will not return | christopher.holmes@sanofi.com | 51-153 | 2023-01-09 | 2023-01-19 | It is expected that an alarm will clear when it has returned into range.
1041 | 51 NYA Thyrogen- Bioreactor E3 QCA for hardware node configuration failure. | ian.odell@sanofi.com | 51 NYA Room 140 | 2023-01-20 | 2023-01-20 | Per FBL-SOP-000576, we won't see QCA's associated with process recipes.
1040 | 51 NYA Thyrogen Purification-NRSR 22SRF020 sample storage not documented per FBL-SOP-000010 ATTACHMENT 1 | jacob.kallon@sanofi.com | 51 NYA Room 108C | 2023-01-13 | 2023-01-19 | See attached copy of NRSR 22SRF020 | Yes | No Deviation
1039 | 45NYA Room 1131-No Product impact-QCC45 Chart Recorder F-15828-002 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 45-1131 | 2023-01-04 | 2023-01-20 | BQ-GOP-000046 Calibration Program Version 4 Section 5.8.3 as found should have been in tolerance
1038 | Cold room 51-87 Quarterly PM went overdue, Work order. 2909322 | patrick.hosegood@sanofi.com | 45 NYA | 2023-01-15 | 2023-01-19 | System operating properly. Pm Completed on Time
1037 | GEMBA for Cz Train 1 Harvest SUM PM Not Assigned | erica.davis@sanofi.com | 8 NYA 1403 | 2023-01-19 | 2023-01-19 | N/A
1036 | Cz - 8 NYA - Growth media spill in CNC area at 8 NYA | kelvin.leo@sanofi.com | 8 NYA | 2023-01-19 | 2023-01-19 | Integral |  | No Deviation
1035 | 8NYA Solution Prep WFI Valve FV-9257025 Dosing Setpoint Changed | stephanie.aceto@sanofi.com | 8NYA, Solution Prep Room 1350 | 2023-01-19 | 2023-01-18 | The expected result would have been the tuning parameters would have remained the same for TP011, TP012, and TP013 as documented under IS_1483875 Version 22.0. |  | Deviation
1034 | 76NYA WFI Manufacturing loop distribution pump found not running | scott.brouillard@sanofi.com | 76NYA Manufacturing suites | 2023-01-19 | 2023-01-19 | Uninterupted supply of WFI for manufacturing activities.
1033 | False alarm conditions for Conductivity (F045-AI-9500-008 in 8NYA) | scott.brouillard@sanofi.com | 45NYA Mech Room | 2022-10-30 | 2023-01-18 | Only true alarm conditions to be captured and logged.
1032 | QCM Cold Room F-15591 Fell Below 2-8°C Qualified Temp. Range | krystal.masseri@sanofi.com | 68 NYA QCM | 2023-01-18 | 2023-01-18 | Per FBL-SOP-000884, cold room F-15591 must be between 2-8°C with the set temperature being 5°C.
1031 | Incorrect instrument chosen during Dynamic EM sampling in Onelims | Adriana.Alban@sanofi.com | 51NYA 152A | 2023-01-18 | 2023-01-18 | it is expected that the Instrument is available for use in the Onelims system.
1030 | Sample bag 7082-01 for production media doesn't have disconnect | nicholas.shoenfelt@sanofi.com | 51-124 | 2023-01-18 | 2023-01-18 | It is expected that the sample bag will be assembled with the required disconnect.
1029 | HPLC Instrument used for GMP assay runs while Annual PM/PV was overdue. | NANCY.WATKINS@SANOFI.COM | 68NYA, QC Chemistry, 2nd floor, room2700 | 2023-01-10 | 2023-01-17 | The expectation is that instrument be within PM expiry during use
1028 | missing daily LAL sample | Stefanie.Erickson@sanofi.com | QC Micro LAL bay | 2023-01-17 | 2023-01-18 | for sample 2300001590 to be in the LAL bin to be tested.
1027 | Observed Fermenter low temperature during DBYFRM004 operation | dena.elmesalamy@sanofi.com | 76 NYA- room 364 | 2023-01-17 | 2023-01-17 | Chiller should runs properly
1026 | Meeting to Discuss SP1 FBL-MBR-00047 Step 8.4.7 buffer entry did not occur at the time of operation. | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 138 | 2023-01-13 | 2023-01-18 | It is expected that buffer information be recorded at time of use per FBL-MBR-00047
1025 | Covid Positive Personnel at 51 NYA | conner.booker@sanofi.com | 51 NYA, 45 NYA, Penthouse hallway, cafe room 140, Hallway cafe, dock room 108, 108C, 128 | 2023-01-12 | 2023-01-14 | It is expected that no individual with Covid-19 enters classified Manufacturing spaces.
1024 | rhTSH Growth Phase Lot CGF1B14 mycoplasma samples bacterially contaminated | michael.searles@sanofi.com | 68 NYA | 2023-01-17 | 2023-01-17 | No bacterial contamination for TSH Growth Day 0 Lot CGF1B14 and valid mycoplasma testing.
1023 | FBS-LBK-000030, Lot DBYHAS001 missing verification | dena.elmesalamy@sanofi.com | 76NYA-3101 | 2023-01-13 | 2023-01-13 | All steps need to be verified immediately.
1022 | No verification signature for manual entry during Cerezyme Phenyl buffer setup. | stanley.durand@sanofi.com | 8NYA RM1212 | 2022-12-09 | 2023-01-16 | Per FBL-SOP-001320 sec. 6.8 Record Review Execution 'For manual entries, confirm additional verification was performed and documented by second operator.
1021 | FZ SUB X2200 pCO2 probe 2 reading failure | michael.knapp@sanofi.com | 8NYA 1333 | 2023-01-17 | 2023-01-17 | Probes remain accurate in their readings.
1020 | 8NYA-Cerezyme- X-2220 experienced multiple failures including the load cells not taring properly and agitator failing to register on Delta V while in use | john.lytle@sanofi.com | Rm 1403 | 2023-01-16 | 2023-01-16 | SUB should be returned to a state of service acceptable to cGMP operations prior to use after remediation
1019 | No reportable Viable Air for routine EM on 14Oct2022 | Craig.Stevenson@sanofi.com | 45 NYA | 2022-10-14 | 2023-01-15 | FBL-SOP-000376 V3.0 step 6.3.7 references SD-000463, which is the LIMS EM Sampling document.
1018 | 8NYA-Cerezyme-X2220 SUB weight not responding during initial Media load operations. | john.lytle@sanofi.com | Rm 1403 | 2023-01-16 | 2023-01-16 | the recipe automation is expected to tare the SUB weight prior to introduction of media and reflect introduction of media by increasing weight accurately
1017 | Missing Zero Filter for EM Batch #: 202202120_1 | Craig.Stevenson@sanofi.com | 51NYA | 2022-06-23 | 2023-01-12 | Per FBL-SOP-001039 V5.0 step 6.4, a zero filter is required to be performed prior to Non-Viable EM sampling.
1016 | QCA alarm caused during refeed on G7/T1 for CGF3B16, agitation timer of 45 min was exceeded | brian.mcgee@sanofi.com | 51NYA/Rm 140 | 2023-01-15 | 2023-01-15 | It is expected that the N2 rotameter will be turned on during the refeed activties in order to maintain pressure
1015 | Dehumidifier for cold room 87 was found by Facilities to be not functioning during PM 2909322 and PM cannot be completed. | alyson.belleville@sanofi.com | 51-87 | 2023-01-14 | 2023-01-14 | It is expected that the unit would be functioning as it should and that PM would be able to be completed.
1014 | Gemba for unverifiable sampling time provided for rhTSH sample | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2023-01-14 | 2023-01-15 | Per FBL-SOP-000964 step 6.3, a sampling time must be provided to receive a sample in iLIMS.
1013 | hTSH CGF2B17 D2 "SIP PASS" Status lost post Production Pressure Hold SIP in Delta V | brian.mcgee@sanofi.com | 51NYA/Rm 140 | 2023-01-14 | 2023-01-15 | It is expected for the SIP Pass status to remain on the reactor if positive pressure was held throughout the activity.
1012 | QA was not able to test Bioburden sample causing sample to expire | Pavel.dobrovolskiy-wynes@sanofi.com | Column Pack *RM1303 | 2023-01-14 | 2023-01-14 | Its expected that bioburden samples are tested within the 24 hour time frame
1011 | hTSH CGFSB017 Wave Reactor C has mix gas pressure failure in Unicorn on 14JAN17 | brian.mcgee@sanofi.com | 51NYA/Rm 152A | 2023-01-14 | 2023-01-14 | It is expected for the Unicorn method to run without issue, and there to be no warning messages for failures.
1010 | Atypical PH noted during Chromatogram review for Q Sepharose column operation. | Subhashini.Satishkumar@Sanofi.com | on 09Jan23 | 2023-01-09 | 2023-01-14 | It is expected that the PH should be within the range during the Q operation.
1009 | 4715 instances of paper forms unable to be reconciled to electronic records for training documentation | erin.young@sanofi.com | 68 New York Ave | 2020-03-12 | 2023-01-13 | FBL-SOP-000835 step 6.1.3- Records must beMaintained on the appropriate media, in a manner that retains the record’s authenticity, reliability, readability, and protection from accidental and/or unauthorized modification throughout the record retention period, based on data sensitivity, tracked when removed and returned from their storage location. FBL-SOP-000942- 6.10.2 all controlled documents or records must be properly identified and maintained in a secure location. FBL-SOP-001486- step 6.2.2.2 All completed GxP records signed via Adobe Sign must be stored accordingly, to preserve original data by acting as specified in the source procedure of the activity, or must be submitted to Records Management. FBL-SOP-001486 Step 6.2.2.6 Retain completed record containing all required signature/entries. Retain original data per the appropriate Quality Management system of parent record. Wet signatures/entries on forms are generally exempt from the "original data" requirement.
1008 | Samples from 2 buffers (832316 and 832320) at 8NYA for ASM Endotoxin were not received for testing | aidan.mula@sanofi.com | 8NYA RM1350 | 2023-01-13 | 2023-01-13 | Determination at Gemba
1007 | Batch number not issued per FBL-SOP-000953 | Edward.kozatek@sanofi.com | 51NYA | 2023-01-03 | 2023-01-13 | The batch issuance should follow FBL-SOP-000953 | Yes | No Deviation
1006 | HA-CMC Manufacturing - Room 222- Labels were found resting on top of Flat Pouch with zipper pouch part number GZ5482-01 | jan.stepanik@sanofi.com | 76NYA/room 222 | 2023-01-10 | 2023-01-10 | Labels should have been attached.
1005 | Moblie Phase Spool Pieces on columns C-3116010 and C-3116020 were qualified without a third party weld inspection | stephanie.aceto@sanofi.com | 8NYA, Suite 1 Downstream Room 1323 | 2023-01-12 | 2023-01-12 | A nonconformance would have been opened to address the spool pieces not being inspected by a third party weld inspector.
1004 | AN 2022-C-076 never filled out for work execution/cleaning and release | dennis.bazinet@sanofi.com | 74NYA/ CC1-13 | 2022-12-15 | 2023-01-13 | Per AN page 5, work execution complete, Facilities cleaning, MFG cleaning and QA release
1003 | Incorrect Incubation for Pompe Plates | Rebecca.Schmidt@sanofi.com | QC Micro, F-15594 (30-35 degree incubator) | 2023-01-09 | 2023-01-13 | FBL-BSP-000012, step 6.6.4
1002 | 8NYA-Room 1170-Product impact unknown-Facilities TOC Analyzer FB-4002-010 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1170 | 2023-01-09 | 2023-01-13 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
1001 | 68NYA-Room 1845-product impact unknown-F-17505 Pressure Indicator found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1845 | 2023-01-12 | 2023-01-13 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
1000 | ASM ASM Initial Stability Window | Kaylyn.Manoogian@sanofi.com | 68NYA/QC Chemistry Lab | 2023-01-06 | 2023-01-11 | Samples would've been testing within the 18 day stability window
999 | Can't verify correct SUT charged against process order | stanley.durand@sanofi.com | 8NYA rm 1323 | 2022-11-23 | 2023-01-11 | It is expected for correct SUT to captured in batch record |  | No Deviation
998 | Automation intervention needed to adjust VFD to get pump speed adjusted to right set point. | stanley.durand@sanofi.com | 8NYA rm 1323 | 2023-01-12 | 2023-01-12 | It is expected per FBL-EWI-000033 during set up recipe for the pump to ramp up to 33LPM. Pump was only able to reach set point of 6LPM. Automation had to intervene to get pump to ramp up to expected set point of 33LPM.
997 | Foil pouches containing reserve containers were found compromised during annual examination | paul.consolmagno@sanofi.com | 76 NYA, QCBP | 2023-01-11 | 2023-01-11 | Reserve material would not be compromised
996 | 500L SUB goes into hold during inoculation of 2k for DO | adam.hildebrand@sanofi.com | 74 NYA/ CC1-03 | 2022-11-18 | 2023-01-12 | It is expected the recipe wouldn't monitor DO parameters when transferring out
995 | During the QS WFI fill portion of the Formulation recipe, the recipe alerted operators that the batch exceeded the 1% variance for final batch size. | richard.manasian@sanofi.com | 8RM1350 | 2023-01-11 | 2023-01-11 | That the final batch size will be within the process variance limits. | Yes | No Deviation
994 | Filter part#793979 not charged against correct MO in Athena | stanley.durand@sanofi.com | 8NYA Room 1323 | 2022-11-27 | 2023-01-11 | It's expected when an item in not in the BOM it should be taken out against the MO in Athena per FBL-SOP-000897. |  | Deviation
993 | Fabrazyme Gen2 _8NYA_Automation intervention needed twice during butyl operations. | stanley.durand@sanofi.com | 8NYA 1333 | 2023-01-11 | 2023-01-11 | It is expected that agitator turns on at the 15 minutes mix and also turn off once the SUM in Upstream reaches 100 kgs.
992 | Discrepancy discovered during Stability 2023 annual count- Sterile HA Powder mock samples in plastic bottles | yekaterina.zavgorodniy2@sanofi.com | NBO Room 1324 Freezer | 2022-03-21 | 2023-01-11 | Per FBL-PTM-000004, 20 mock samples are sent to vendor (West Pharmaceutical Services) at the 24 month time point for container closure testing and a container closure result is generated. | Yes | Deviation
991 | SUB 2200 recipe went into hold | edgar.vale@sanofi.com | 8NYA/room 1333 | 2023-01-10 | 2023-01-10 | pCO2 probe would maintain parameters within the 20mmHg limit
990 | Resuspension/Casting at 76 NYA SAP Discrepancy for HA-CMC (2139-32) | jan.stepanik@sanofi.com | 76NYA/room 209 | 2023-01-08 | 2023-01-09 | There should not have been such abnormal discrepancy.
989 | Filter failed 3x during 1716 buffer batching | nicholas.shoenfelt@sanofi.com | 45-190 | 2023-01-10 | 2023-01-10 | It is expected that the filter integrity test will pass per FBL-MBR-000010.
988 | Operator was in the fermentation room and after, entered the aseptic core room | valeria.moreno@sanofi.com | HA Fill 76NYA room 355 | 2023-01-10 | 2023-01-10 | FBS-SOP-00066
987 | Missing end time on the AM sample run; credo logbook # 2023QCC0006 on 06Jan2023 | michael.montiverdi@sanofi.com | unknown | 2023-01-06 | 2023-01-06 | End time to be entered. | Yes | No Deviation
986 | 68NYA-Metrology-no product impact-Metrology Standard FBS-0097 no as found data taken during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2022-12-20 | 2023-01-10 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been able to be collected
985 | 76NYA-Room 209-Product Impact unknown-Balance F-21199 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 76-209 | 2023-01-09 | 2023-01-10 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
984 | Site 212 not sampled for non-viable air on Day 3 of Action Notice 2022-N-055 | valeria.moreno@sanofi.com | HA Fill 76NYA | 2023-01-08 | 2023-01-09 | Validated, passing EM non-viable particulate results from 3 consecutive days of static monitoring after cleaning for all sites including BSCs/LFHs
983 | X-4069 Incubator RQ not performed for temperature | adam.hildebrand@sanofi.com | 74 NYA | 2023-01-09 | 2023-01-09 | It is expected that the RQ would be performed within specified timeframes.
982 | Olipudase - Lot CGHSD001 -8 NYA - Loss of Batch Traceability for A2B Straight Tube Set (P/N 756380) | sheila.pandolfo@sanofi.com | 8 NYA Room 1403 | 2022-11-03 | 2023-01-06 | It is expected that there is lot traceability to all materials used in a Sanofi batch.
981 | hTSH Unicorn Method Lost for Lot CGFSB017 | brian.mcgee@sanofi.com | 51NYA 152A | 2023-01-08 | 2023-01-08 | It is expected for the method to run during the entire stage of the process, until it is scaled up to the next method.
980 | FZ SUB X-2200 Media Switchover error resulting in inaccurate Rate of Change reading | michael.knapp@sanofi.com | 8NYA 1333 | 2023-01-08 | 2023-01-08 | Perfusion EM calculates accurate Rate of Change based on active tote. | Yes | Deviation
979 | Shipping bag transfer recipe | zachary.rickard@sanofi.com | 8NYA/1703 | 2023-01-07 | 2023-01-07 | Recipe should end after filling second barrel and not look to fill a third.
978 | 8NYA-Pompe-Recipe automation calling for incorrect Media pinch valve, automation intervention requested to begin Perfusion | john.lytle@sanofi.com | Rm 1434/ 8NYA | 2023-01-06 | 2023-01-06 | recipe automation should be verified prior to production start, correct Media pinch valve should be associated with SUB X-2240
977 | GEMBA for 3 Unreviewed Entries in QC Micro Reagent Logbook | jessica.kychler@sanofi.com | QC Micro 68 New York Avenue, floor 4, room 4700 | 2022-11-03 | 2023-01-06 | It is expected that Quality Control Reagent logbook entry will be reviewed prior to the reagent, microbiological media, control, or standard being used | Yes | Deviation
976 | Reserve vials 161931, 161932, and 161933 (56726) for Olipudase UF/DF Final concentrate burst in Sample Management freezer | paul.consolmagno@sanofi.com | 68 NYA, Freezer 98963 | 2022-12-20 | 2023-01-05 | Vials 161931, 161932, and 161933 would be kept as reserve samples until 3 years after the product had expired. | No | Deviation
975 | 714022 part missing from RF1BHC122 | richard.manasian@sanofi.com | 1350 at 8 NYA | 2022-11-30 | 2023-01-06 | Part would be used per EWI
974 | Improperly parsed EM Non-viable sample data for CGFSB015 | Adriana.Alban@sanofi.com | 152A | 2022-12-03 | 2023-01-05 | It is expected that operators parse non-viable data from the climet to the parsing tablet per FBL-SOP-000207 and after check in biolims for successful parsing of non-viable data
973 | Growth Media LAL sample found outside of cold room | nicholas.shoenfelt@sanofi.com | 51-108 | 2023-01-05 | 2023-01-06 | LAL samples are to be stored in cold room 108C until picked up for delivery to QC. | Yes | No Deviation
972 | 68NYA-Metrology-no product impact-Metrology Standard FBS-0026 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2022-12-21 | 2023-01-06 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
971 | No second verify per CCR. | bailey.gleason@sanofi.com | 8 NYA | 2023-01-05 | 2023-01-05 | to be determined |  | No Deviation
970 | 74NYA room UTIB-02 Dew point indicator overdue for calibration | robert.bauch@sanofi.com | 76-UT1B-02 | 2023-01-01 | 2023-01-05 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 unit should have been calibrated before their due date.
969 | asks performed prior to approval and posting of Action Notice for 76NYA New Still tie-in. | brizaida.medina-vega@sanofi.com | 76 NYA DSMB WFI distribution | 2022-12-16 | 2023-01-05 | No work should begin prior to action notice approval and postings.
968 | 76NYA-Room 227-Facilities-EMS Temperature for 76-AHU 2A found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-227 | 2023-01-03 | 2023-01-05 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
967 | Overdue RQs for Isolator | Cassandra.Lotito@sanofi.com | 76 NYA 3104 | 2022-12-01 | 2023-01-04 | RQs are to be completed within requalification window
966 | Approved version of FBL-SOP-000379 used in batch record, V6.0 became effective on 18OCT22 and was used on 09OCT22 | pamela.mills2@sanofi.com | 74 | 2022-10-09 | 2023-01-04 | effective version should have been used
965 | Required Minimum Pressure on PI-06350 Unachievable | james.letourneux@sanofi.com | 51RM140 | 2023-01-04 | 2023-01-04 | A pressure regulator capable of achieving the required batch record specifications should be installed on all reactors.
964 | Process pH is identified to be out of range during chromatogram review from PI | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 132 | 2022-12-20 | 2023-01-03 | It is expected that the process pH perimeter is in range per attachment 6 of FBL-SOP-000788
963 | Process pH is identified to be out of range during chromatogram review from PI | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 132 | 2022-12-20 | 2023-01-04 | It is expected that the pH trend be within parameters set per attachment 6 in FBL-SOP-000788
962 | Low recovery for TSH Formulated Bulk lot CGF2B08 of 92%, range 93-105% | alyson.belleville@sanofi.com | 51 NYA | 2023-01-04 | 2023-01-04 | It is expected that the recovery would be in range
961 | 22 Late PMs Not Completed/Extended By 31Dec2022 | tanni.arif@sanofi.com | 51NYA, 74NYA, 76NYA, 80NYA | 2023-01-01 | 2023-01-03 | Complete all PM tasks prior to PM closure (otherwise approved by QA) or a deviation is required per site specific SOP.
960 | No product - No Lot - *NYA- Multiple unreviewed Dynamic EM / routine EM packets for Non-Viable sampling during 2021 | Craig.Stevenson@sanofi.com | 8NYA | 2021-02-27 | 2022-12-29 | All review should be completed by the functional area within 21 days of the last entry.
959 | Unknown Product - Unknown Lot - 8NYA - Unacceptable Zero filter obtained prior to Dynamic EM | Craig.Stevenson@sanofi.com | TBD | 2022-12-03 | 2023-01-03 | Fifth particle count on particle counter FB-4529 for zero count filter check result 0 Ft3 for filter size 0.5 µm
958 | Tare weight discrepancy during batching of Thyrogen media 609615 at 51NYA | nicholas.shoenfelt@sanofi.com | 51-124 | 2022-12-27 | 2023-01-01 | It is expected that the bag weight + the fill weight = the gross weight (within a 1kg range).
957 | Bowie Dick parameters from printout do not align with SOP | Valeria.moreno@sanofi.com | 76 NYA M-61 | 2023-01-03 | 2023-01-03 | FBS-SOP-000419, Attachment 2
956 | Cz Zinc temperature excursion during Production DV recipe | stanley.durand@sanofi.com | 8NYA rm1503 | 2022-12-14 | 2023-01-02 | Per FBL-CST-000029 process temperature must be 2-13 degree C
955 | Sampling for ASM Blank Run done prior to mixing of eluate sum | robert.morawski@sanofi.com | 1323 | 2023-01-02 | 2023-01-03 | Mix recipe to be located in the record prior to needing the completion of the CIP of the Skid.
954 | 8NYA-Fz-low DO on SUB X2200 causing system hold | john.lytle@sanofi.com | Rm1333 | 2023-01-02 | 2023-01-02 | %DO >/= 30
953 | ASM AEX Blank Run ran out of 848078 buffer during Elution | robert.morawski@sanofi.com | 8NYA 1323 | 2023-01-02 | 2023-01-02 | Enough 848078 buffer to finish elution.
952 | Buffer 705663 run out of solution during Reagen and storage | Christian.Venegas@Sanofi.com | 8NYA Downstream Phase 2 Room 1503 | 2022-12-31 | 2022-12-31 | We will need to replace SUT's and a Buffer 705663 replacement
951 | 51 NYA- Thyrogen- 10L Wave tubing damaged during media fill (CGFSB016) | ian.odell@sanofi.com | 51 NYA Room 152A | 2022-12-29 | 2022-12-29 | Tubing would not be damaged by using of peristaltic pumps set up for Wave use.
950 | No product-no lots-68NYA-10mL PBS for mL water bioburden samples not selected in BioLIMS | larissa.fairweather@sanofi.com | 68NYA/4700/QCM | 2022-11-11 | 2022-12-28 | ALF-MTD-000018 V.12.0 Step 9.1.8.7. Samples volumes <10mL must be aseptically dispersed in at least 10mL of a sterile dilution/rinse fluid (10mL volume of dilution/rinse fluid completely submerges the bioburden filter, ensuring even distribution of the sample over the filter surface).
949 | 51 NYA- Thyrogen- Climet (NV) stopped before 11th sample reading during Grade A operations (CGFSB016) | ian.odell@sanofi.com | 51 NYA Room 152A | 2022-12-28 | 2022-12-28 | Climet will be stopped at or after the 11th (Climet) reading per FBL-SOP-000755
948 | Incorrect Storage Unit automatically updated in the post viral freeze record. | cory.lefort@sanofi.com | 8 NYA | 2022-12-26 | 2022-12-26 | The record was suppose to update the correct storage unit with the correct volume that was pulled in from the RoSS. | Yes | Deviation
947 | No negative control result recorded for routine EM, 8NYA | megan.botti-marino@sanofi.com | 68 NYA | 2022-12-14 | 2022-12-26 | It is expected that a negative control be included with routine EM samples
946 | Incorrect version of form FBL-FRM-000796 used | megan.botti-marino@sanofi.com | 8 NYA | 2022-12-04 | 2022-12-26 | It is expected the most recent version of forms and SOPs be used at time of operations.
945 | Unverified sampling time for expired Bioburden sample | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2022-12-24 | 2022-12-25 | Bioburden samples must be tested within 24 hours per FBL-SOP-000019
944 | Counts not recorded for Grade C plates from 74NYA | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2022-12-01 | 2022-12-25 | Per FBL-SOP-000604 step 6.6.2 data entry must be performed by counting analyst.
943 | Label discrepancy with TOC Swab samples | michael.aiello@sanofi.com | 68 NYA RM 4720 | 2022-12-23 | 2022-12-24 | Duplicate Swab 3 (TOC) for same CV pull | Yes | No Deviation
942 | Incorrect HOFF recipe launched | zachary.rickard@sanofi.com | 8NYA/1365 | 2022-12-24 | 2022-12-24 | Correct recipe will launch and product will freeze accordingly |  | Deviation
941 | 8NYA FZ SUB TCU alarm due to low level | michael.knapp@sanofi.com | 8NYA 1333 | 2022-12-24 | 2022-12-24 | TCU for SUB has sufficient glycol and does not hold for low level. when performing equipment changeover and installing a new SUB jacket, the glycol line and jacket are primed. |  | Deviation
940 | 51 NYA- Thyrogen- 10L Wave bag not issued to process order (CGFSB008) | ian.odell@sanofi.com | 51 NYA Room 152 | 2022-06-20 | 2022-12-23 | It is expected that all materials used in cGMP batches be issued to the associated process order.
939 | An operator used footnotes out of sequence | dena.elmesalamy@sanofi.com | 76NYA-3102 | 2022-10-25 | 2022-12-22 | Footnotes to be entered in a sequential number
938 | Using footnotes out of sequence | dena.elmesalamy@sanofi.com | 76NYA-3102 | 2022-12-25 | 2022-12-22 | Footnotes to be entered in a sequential number
937 | Partial load material take over of CECX eluate during AEX load check step | michael.daly@sanofi.com | 8 NYA Room 1323 | 2022-12-03 | 2022-12-22 | It is expected the eluate is completely taken over during the eluate take out step
936 | Recording error for Brix value for during batching of GZ1713 buffer | nicholas.shoenfelt@sanofi.com | 80-102 | 2022-12-04 | 2022-12-22 | It is expected that the value be recorded as seen correctly.
935 | MFG breach of rooms 1330, 1331, 1333, and 1334 in 8NYA prior to QC sampling and action notice release (2022-I-067) | marissa.lentoni@sanofi.com | 8 NYA, Rooms 1330, 1331, 1333, and 1334 | 2022-12-22 | 2022-12-22 | QC Micro would be notified that the cleanings were completed so that sampling could be performed. QC Micro would perform sampling and provide non-viable results in order to release the areas.
934 | 2 QC Micro EM/CU samplers tested positive for Covid | Craig.Stevenson@sanofi.com | TBD | 2022-12-17 | 2022-12-21 | N/A
933 | Organism of Concern (Bacillus Cereus) detected in 74 NYA Cleaning Verification X-2650 Depth Filtration CIP Lot 2022-CV-0739 | Craig.Stevenson@sanofi.com | 74NYA, Falcon | 2022-11-21 | 2022-12-22 | It is expected that the Falcon CV generate a result of 0CFU for Bioburden
932 | Incorrect FIT program used to test manifold filter for FB CGF2B05 and CGF1B06 | alyson.belleville@sanofi.com | 51-87 | 2022-12-08 | 2022-12-21 | it is expected that part number and test 6906 would be used to test this filter
931 | V-2561 Load Recipe not advancing | cindy.guerra@sanofi.com | 74NYA / PU1-09 | 2022-12-22 | 2022-12-22 | V-2561 Load Recipe would not go into hold
930 | CV paperwork missing and not reconciled | adam.hildebrand@sanofi.com | 74 NYA CC2-01 | 2022-05-22 | 2022-12-21 | It is expected that CV paperwork would not be generated, requested and issued for an asset that is cleaning validated.
929 | Falcon Bioburden Sample Leaking/ Insufficient Volume to Test | jessica.kychler@sanofi.com | QC Microbiology, 68 NYA floor 4 room 4700 | 2022-12-21 | 2022-12-21 | It is expected that bioburden samples will be secure and will not be leaking/ compromised
928 | GEMBA: Calculation Error on TSH Stability 30M Sample, Lot AW0145 for Native Page Assay | Kaylyn.Manoogian@sanofi.com | 45NYA, QC Chemistry Lab | 2022-09-29 | 2022-12-20 | Calculation would've been performed correctly.
927 | At the end of a buffer transfer at 8NYA, a leak was detected in the SUT tote bag. | thomas.loik@sanofi.com | 8NYA RM 1202, outside of solution prep | 2022-12-21 | 2022-12-21 | That the tote bag collecting buffer product will not leak.
926 | On 20DEC22, X-3051 (MMC skid) experienced a Hi Hi pressure alarm (PI-305012A) causing Capto MMC production recipe to go into hold during sanitization phase. | craig.valery@sanofi.com | 74 NYA, PU1-09 | 2022-12-20 | 2022-12-20 | It is expected that the recipe would continue without any hold conditions.
925 | GEMBA 8NYA: Daily Fz wave reactor ViCell samples performed in inoc room during AN 2022-I-067 with inoc room not part of the AN | lolav.ahmed@sanofi.com | 8NYA room 1333 & 1343 | 2022-12-17 | 2022-12-20 | N/A
924 | E/L and SFU documents don't include 691805 for use in Cz | Sarah.Francis@sanofi.com | 8 NYA | 2021-09-30 | 2022-12-19 | An assessment for assembly PN 691805 would have been performed and documented in FBL-RPT-001164 and FBL-RPT-001310
923 | UFDF production recipe won't advance due to Temp EM | robert.morawski@sanofi.com | 8NYA 1323 | 2022-12-20 | 2022-12-20 | Recipe advances into diafiltration phase. |  | Deviation
922 | Seprafilm - CBYSEP087 - 76 NYA Room 233- COVID positive employee in Seprafilm DHT | rebecca.Cunningham@sanofi.com | 76 NYA/Room 233 | 2022-12-02 | 2022-12-19 | It is expected that personnel would not enter classified manufacturing space when they have an illness per FBS-SOP-000667.
921 | Multiple instances where QA specialist performed as both "Recorded By" and "Verified By" | julie.souza@sanofi.com | 74 NYA, Room 1017 | 2022-12-02 | 2022-12-20 | FBL-SOP-000942 step 6.11.6 defines proper use of signatures which requires verification to be performed by a second individual.
920 | ID not submitted for EM site 210, Room 355 (2200051436) | Craig.Stevenson@sanofi.com | 76 New York Ave, Room 355 | 2022-12-01 | 2022-12-20 | Per FBL-SOP-001496 V 1.0, step 6.2.19, all plates with growth should be saved until results have been evaluated for excursion level results and plates requiring ID have been submitted.
919 | Incorrect version of FBL-FRM-000796 used 07NOV22 | megan.botti-marino@sanofi.com | 8 NYA | 2022-11-07 | 2022-12-19 | That the effective version of an SOP be used. | Yes | Deviation
918 | 74NYA-Room PU1-24-Dept PF74-FB-3659-014 Retentate Flow Indicator on the TFF Skid X-3211 Found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 74-PU1-24 | 2022-12-16 | 2022-12-19 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
917 | CZ Q membrane filtration recipe was stuck due to buffer tote not reading correctly | Pavel.dobrovolskiy-wynes@sanofi.com | 1503 | 2022-12-18 | 2022-12-18 | Its expected the buffer weight sensors will work properly allowing the recipe to advance
916 | Cz phenyl equilibration bioburden samples expired | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA Room 1503 | 2022-12-18 | 2022-12-18 | Its expected that samples dropped off for pickup will be delivered to QC within the timeframe. | Yes | Deviation
915 | Post-Viral Freeze Bag Quantity Not Updated | cory.lefort@sanofi.com | 8 New York Ave / 1313 | 2022-12-17 | 2022-12-17 | The quantity in storage unit 79232 (bag 19) would of been updated in the record to reflect that volume that was weighed out on the RoSS.
914 | High high alarm for DO was incurred on 01DEC22 upon | adam.hildebrand@sanofi.com | 74 NYA CC2-01 | 2022-12-01 | 2022-12-17 | It is expected that the recipe for the 500L SUB would be built accordingly when transferring out and not monitoring these process parameters when transferring out.
913 | Leak manifold during post-viral freeze operation. | cory.lefort@sanofi.com | 8 New York Ave / 1313 | 2022-12-16 | 2022-12-16 | No leak in the manifold line and bag to fill to the set point.
912 | Sample Apparatus Detached From Buffer Bag During Fill | jeff.hill@sanofi.com | 45RM190 | 2022-12-16 | 2022-12-16 | Per FBL-MBR-000010 Step 4.17 Fill the Containers
911 | TSH CGF4B15 QCA for Low O2 on 16DEC22 | brian.mcgee@sanofi.com | 51NYA/Rm 140 | 2022-12-16 | 2022-12-16 | It is expected for the p02 to be within allowable range.
910 | Mixer Setting is out of range | dena.elmesalamy@sanofi.com | 76NYA-3102 | 2022-10-12 | 2022-12-15 | The Mixers setting range is between 89.5 to 90.4, per FBS-MBR-000075, Step 7.5.30.1 pg 40
909 | Missing to record performing the CIP for V-3060 in LBK FBS-LBK-000029 during CBYENG015 and CBYENG016. | dena.elmesalamy@sanofi.com | 76NYA-364 | 2022-11-04 | 2022-12-16 | Record performing the CIP in the associated LBK
908 | Pre-Use Vent filter failed 3 tests | charles.hart@sanofi.com | 51NYA, Rm.143 | 2022-12-06 | 2022-12-15 | Per FBL-SOP-001011 step 6.7.2.6 For pre-use filters, if a filter fails a total of three integrity tests per the decision tree in ATTACHMENT 5, initiate a deviation per BQ-GOP-000248, place the filter into a bag, attach a Sanofi approved label, and provide the filter to the MSAT group for further investigation. • Obtain a new filter, wet and test the new filter per the appropriate steps listed
907 | Automation Support needed for TFF Skid X3211 | alvaro.rueda@sanofi.com | 74 NYA Falcon | 2022-12-16 | 2022-12-16 | HMI switched back on with no issues
906 | MES/Athena comm issue- transactions unable to be processed at 8NYA | Shiful.Kohli@sanofi.com | 8NYA | 2022-12-15 | 2022-12-15 | MES/Athena transactions are processed as per design
905 | No Product - No Lots - 51NYA - No Zero Filter was not parsed into LIMS prior to QC Micro sampling on 23Jun2022 | Craig.Stevenson@sanofi.com | 51NYA | 2022-06-23 | 2022-12-15 | Per FBL-SOP-001039 step 6.4 and step 6.8, the zero filter should have been performed prior to sampling and parced into Lims.
904 | 8NYA-Cz-MCC X-2811 outlet ports leaking past closed pinch valve during CIP | john.lytle@sanofi.com | 1403 | 2022-12-15 | 2022-12-15 | it is expected that these outlet ports remain integral for the entirety of the CIP |  | Deviation
903 | No Product - No Lot - 68NYA - QC Micro missed weekly water site 74-RO-2 | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2022-12-10 | 2022-12-15 | Per FBL-SOP-000336, V13.0, Attachment 1 Page 1, CU site 74-RO-2 should collect Bioburden, Coliform and Endotoxin on a weekly frequency
902 | Incorrect SUT in Chrom CIP Record | robert.morawski@sanofi.com | 8NYA RM1503 | 2022-12-15 | 2022-12-15 | SUT in record would be correct |  | Deviation
901 | Nano filter FIT step not executed; filter missing: RF1C3C023/CGHDS023 | alejandro.martinez2@sanofi.com | 8NYA Solution Prep, 1350 | 2022-12-15 | 2022-12-15 | Determine path forward on missing filter.
900 | Pompe samples arrived to QC expired | elizabeth.couturier@sanofi.com | 8 NYA | 2022-12-03 | 2022-12-15 | It is expected that samples are transported to QC at the first pickup after sampling to ensure bioburden testing can be performed within the 24-hour time limit
899 | Settle Plate Open/Close Time not Recorded in BioLIMS | jessica.kychler@sanofi.com | QC Microbiology, 68 NYA floor 4 room 4700 | 2022-12-02 | 2022-12-15 | Expected results: open/close times are entered for settle plates when performiing spread plate bioburden | Yes | Deviation
898 | 51 NYA- Thyrogen- Bioreactor sample allocation signature missing. | ian.odell@sanofi.com | 51 NYA Room 140 | 2022-06-22 | 2022-12-14 | All performed steps will be signed for as required per MBR.
897 | 8NYA-Room 1350-Dept FE8 & Solution Prep-product impact unknown- one HEPA filter in a bank of 38 filters found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1350 | 2022-12-13 | 2022-12-15 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
896 | 8NYA-MES-Steps not complete in MES at time of execution for P/N GZ75000040, WFI (unfiltered) for multiple Batches | john.lytle@sanofi.com | 8NYA-Upstream, multiple | 2022-05-01 | 2022-12-14 | it is expected that steps be executed in MES at the time of completion
895 | ASM Nano Post Viral Yield Calculation out of Range | stanley.durand@sanofi.com | 8NYA rm1323 | 2022-12-14 | 2022-12-14 | Per FBL-CST-000032 the yield is expected to be above at or above 90%
894 | ASM Nano Formulation storage unit lock at take out start material step in the record | stanley.durand@sanofi.com | 8NYA rm 1323 | 2022-12-14 | 2022-12-14 | Upon performing the start material take out step SU should be available | Yes | No Deviation
893 | 705661 Removed From MMC Skid Prior to Regen/Store | zachary.rickard@sanofi.com | 8NYA/1503 | 2022-12-13 | 2022-12-14 | Buffers will not be broken down until all processes are completed.
892 | GEMBA 74NYA - TFF FBL-MBR-000185 Integrity Check failed | William.Bradley@sanofi.com | PU1-24 | 2022-12-14 | 2022-12-14 | Integrity check should pass
891 | No Products – No Lots - 76NYA, HA - Incomplete Non-Viable EM sampling of site 221. | Craig.Stevenson@sanofi.com | 76 New York Ave, Room 358 | 2022-10-03 | 2022-12-14 | Per FBL-SOP-000207 Attachment 5, The following setting are to be used to collect a routine EM sample. • [VOLUME]: 1.0 cubic feet • [INITIAL DELAY]: 00:01:00 • [SAMPLE DELAY]: 00:00:15 • [STOP AFTER N]: Enabled • [SAMPLE NUMBER 'N']: 3
890 | Nanofiltration formulation conc out of specification/ Drug substance conc low end of range | michael.daly@sanofi.com | 8 NYA Room 1323 | 2022-12-08 | 2022-12-13 | it is expected the formulation conc is between 4.0-5.0. It is expected the
889 | Unable to get passed the Bag ID to unload in the Cz bag freezing HOFF step | stanley.durand@sanofi.com | 8NYA rm1703 | 2022-12-13 | 2022-12-13 | Bag ID to unload and Set Bag CX value should be executed automatically
888 | Multiple errors detected in HA fill packets | megan.botti-marino@sanofi.com | 76 NYA Fill room | 2022-09-07 | 2022-12-14 | 1. That all samples be taken and submitted with the fill paperwork. 2. that lims numbers be pulled for all samples taken. 3. That plates not exceed the four hour exposure time per 6.3.3 FBS-SOP-000607. 4. That a viable air would have been taken for hour ten. 5. For the action level steps should have been taken Per FBS-SOP-000607 Step 6.3.9, and a note made on the paperwork.
887 | No Product - No Lots - Caps required per FBL-SOP-000478 are not being used | paul.consolmagno@sanofi.com | 45 NYA, 51 NYA, 76 NYA | 2022-12-12 | 2022-12-12 | Sample ports that step 6.6 is applicable to would have a red cap fit snugly over them after santizing | No | Deviation
886 | ASM-ICB Drug Substance Rabbit Pyrogen Samples submitted in wrong conical tube and frozen at <-60oC | michael.searles@sanofi.com | 68 NYA, QC Microbiology, 4740 | 2022-12-08 | 2022-12-12 | It is expected that the ASM-ICB drug substance samples (2 x 25 mLs) for rabbit pyrogen testing (RPT) would be submitted in 50 mLs polystyrene conical tubes and then stored at 2-8C upon sample receipt.
885 | ASM-ICB Drug Substance Rabbit Pyrogen Samples submitted in wrong conical tube and frozen at <-60oC | michael.searles@sanofi.com | 68 NYA, QC Microbiology, 4740 | 2022-12-08 | 2022-12-12 | It is expected that the ASM-ICB drug substance samples (2 x 25 mLs) for rabbit pyrogen testing (RPT) would be submitted in 50 mLs polystyrene conical tubes and then stored at 2-8C upon sample receipt.
884 | A step was not signed in the batch record by performer but only the verifier | alvaro.rueda@sanofi.com | 74 NYA Falcon | 2022-12-11 | 2022-12-12 | Performed signs performed by step
883 | DO high high alarm limit reached on X2210 during inoculation transfer to 2k. | adam.hildebrand@sanofi.com | 74 NYA CC1-02 | 2022-12-12 | 2022-12-13 | It is expected that the parameters would be in a state of control during inoculation of the 2k reactors.
882 | Grade A fill areas in 76 NYA sampled without obtaining an acceptable zero filter reading | megan.botti-marino@sanofi.com | 76 NYA, site 211 and 214, room 355 | 2022-10-28 | 2022-12-12 | That the 5th reading of the zero filter be zero prior to sampling in a Grade A space.
881 | Cz HOF Recipe in Alarm "Overpressure Compressor" | jarad.kukla@sanofi.com | 8NYA, 1365 | 2022-12-12 | 2022-12-12 | HOF recipe should not go into alarm for "Overpressure Compressor"
880 | 8 NYA LAL samples 509338, 509340 found expired upon reconciling sample pendings | Arpan.christian2@sanofi.com | 68 NYA, QC Microbiology, 4700 | 2022-12-09 | 2022-12-11 | It is expected that samples should have been delivered to the correct laydown area for pick-up and delivery to respective testing department | Yes | Deviation
879 | TSH CGFSB015 Delta V Prompt Acknowledged By Mistake on 12DEC22 | brian.mcgee@sanofi.com | 51NYA/Rm 140 | 2022-12-12 | 2022-12-12 | It is expected for the prompt to not being acknowledged until the media has been equilibrated
878 | Two hood operations performed in 76 after an unacceptable zero filter result was obtained | megan.botti-marino@sanofi.com | 76 NYA, laminar flow hood F-0704, Room 364 | 2022-11-21 | 2022-12-12 | According to table 2 of FBL-SOP-000755 "Environmental Monitoring Under Dynamic Conditions of Critical Manufacturing Processes for the Framingham Campus Manufacturing Areas," the only acceptable limits for a zero filter in a Grade A space are counts of 0.0.
877 | Olipudase - L/N: CGHAE004 and CGHAE005 - Multiple Bioburden ILIs | Craig.Stevenson@sanofi.com | 8NYA | 2022-12-12 | 2022-12-12 | It is expected that product does not generate bioburden results.
876 | Incorrect butyl Eluate barrel was used for CZ MMC pooling | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA Room 1503 | 2022-12-12 | 2022-12-12 | Its expected that manufacturing used appropriate MCC butyl eluates for MMC pooling (76548)
875 | 68NYA-QCM-product impact unknown-3 Alpha Moisture Hygrometers (F-20754, F-20755, F-20756) found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-4700 | 2022-11-29 | 2022-12-12 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
874 | Two (2) AHU Preventative Maintenance Work Orders (2893253, 2893254) Were Not Completed by Extended Due Date 30Nov2022 | tanni.arif@sanofi.com | 76NYA | 2022-12-01 | 2022-12-11 | Complete all PM tasks prior to PM closure (otherwise approved by QA) or a deviation is required per site specific SOP.
873 | 8NYA-Cz-SUB X2250 DO probes not reading accurately (follow up) | john.lytle@sanofi.com | Rm 1403 | 2022-12-11 | 2022-12-11 | MSAT is not able to give meaningful guidance for manual pH adjustments based on the current state of the SUB
872 | SARS-CoV2-CGHCE024-74 NYA-CAPTO MMC sample leak (COV_MMC_827588_EQUIL_BIOB) | craig.valery@sanofi.com | 74 NYA, PU1-09/QC 68NYA | 2022-12-10 | 2022-12-10 | It is expected that samples are to be completely sealed and in the appropriate container when received by QC
871 | 8NYA-Cz- X2250 ATF flow unstable following process hold conditions | john.lytle@sanofi.com | Rm 1434/ 8NYA | 2022-12-10 | 2022-12-10 | After the unexpected hold on the SUB and ATF system and addressing the cause of the hold, the system should be able to be restarted and return to expected operational parameters
870 | ASM Ross Fill recipe did not end properly | stanley.durand@sanofi.com | 8NYA room 1313 | 2022-12-09 | 2022-12-09 | Per FBL-EWI-000043 it is expected for Ross Fill recipe to end after last bag has been filled.
869 | High pressure of Capto Column operations during sanitization phase in 74NYA | peter.darcy@sanofi.com | PU1-09 in 74NYA | 2022-12-09 | 2022-12-09 | The pressure shall be below 30psi during operations
868 | 8NYA-Cz-SUB X2250 DO probes not reading accurately | john.lytle@sanofi.com | Rm 1434 | 2022-12-09 | 2022-12-09 | it is expected that we make accurate adjustments to our p02 based on a combination of online and offline analytical comparisons
867 | Cz MMC sampled but not picked up | stanley.durand@sanofi.com | 8NYA room 1410 | 2022-12-03 | 2022-12-09 | Per FBL-EWI-000009 samples taken should be placed on correct shelf in sample pass through fridge and picked up.
866 | 51 NYA- Thyrogen- Quality Critical alarms discovered during ANMS review (CGFSB006) | ian.odell@sanofi.com | 51 NYA Room 152A | 2022-05-25 | 2022-12-09 | No quality critical alarms during cGMP production process.
865 | ASM UFDF production recipe will not advance due to low tank volume | robert.morawski@sanofi.com | 8NYA 1323 | 2022-12-09 | 2022-12-09 | Recipe advances into diafiltration phase.
864 | Reactor clean expiry time reached during PHT | paul.mclaughlinjr@sanofi.com | 74NYA | 2022-12-09 | 2022-12-09 | Reactor would be steamed prior to CHE
863 | 08DEC221030 805006 Glycol recipe not launching | richard.mansian@sanofi.com | 1350 at 8 NYA | 2022-12-09 | 2022-12-09 | delta V would promt for a glycol addition before showing mix timer
862 | Cz Phenyl BHT containing 0.1M 705661 buffer leaking from bag | stanley.durand@sanofi.com | 8NYA room 1212 | 2022-12-09 | 2022-12-09 | Tote bags are expected to be leak free.
861 | Skid did not have adequate time to reach target pressure during Nano Filtration | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1323 | 2022-12-08 | 2022-12-08 | Its expected that the recipe will allow enough time for skid to reach target pressure
860 | Falcon - Invalid TOC Swab 2 Results for 2022-CV-0724 | molly.pomeroy@sanofi.com | 74 NYA | 2022-11-26 | 2022-11-26 | It is expected that TOC swab results be found passing and validated
859 | Cracked CV Vial | Rebecca.Schmidt@sanofi.com | QC Micro Nexus bench | 2022-12-07 | 2022-12-08 | FBL-SOP-000960
858 | Incorrect thermometer for conductivity recorded | Rebecca.Schmidt@sanofi.com | QC Micro Water Chem Bench | 2022-12-07 | 2022-12-07 | FBL-MTD-000009 | Yes | Deviation
857 | No reportable viable and surface results for Site 132-1 51 NYA (467094, 467016, 466960) | marissa.lentoni@sanofi.com | 51 NYA, Room 132, Site 132-1 | 2022-08-09 | 2022-12-07 | Samples would be present on batch they are intended to be linked to and results would be recorded for these samples.
856 | No reportable conductivity results for 51-STM-3 (duct 432046 or retain 432057) | marissa.lentoni@sanofi.com | 51 NYA, Room 146 | 2022-09-23 | 2022-12-07 | Conductivity would be tested for this site.
855 | Particles observed in T-111 during media prep at 51NYA | nicholas.shoenfelt@sanofi.com | 51NYA Room 124 | 2022-12-08 | 2022-12-08 | It is expected that T-111 is particulate free.
854 | ASM Nano Equil / Chase buffer stored in the fridge | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1323 | 2022-12-08 | 2022-12-08 | Its expected that buffer 76284 will be at 16-24 degrees for processing
853 | TSH Nano filter 714715 failed post-use integrity test 3 times | alyson.belleville@sanofi.com | 51 NYA, Room 88 | 2022-12-07 | 2022-12-07 | It is expected that the filter would pass post-use integrity test
852 | 74 EMS Alarm Call Outs Delayed 48 Hours | gary.dodakian3@sanofi.com | 74 | 2022-12-07 | 2022-12-07 | ems call outs will go out as soon as possible
851 | Walk-through performed without training | valeria.moreno@sanofi.com | HA Fill 76NYA | 2022-12-06 | 2022-12-06 | FBL-SOP-000709, attachment 9A: Manufacturing Walk-Through
850 | UFDF CIP record unable to be generated from production record | robert.morawski@sanofi.com | 1323 | 2022-12-06 | 2022-12-06 | Record would create child MO to perform UFDF cleaning.
849 | SUM unable to deallocate from UFDF record | robert.morawski@sanofi.com | 1323 | 2022-12-06 | 2022-12-06 | record properly reallocates SUM
848 | FALCON LAL Cell Boost cracked sample vial | emily.devaney@sanofi.com | 68 NYA 4th Floor, Micro Lab - LAL bench | 2022-12-06 | 2022-12-06 | Per Quality Control Sample receipt SOP FBL-SOP-000964, Step 6.1.7 damaged vials and or containers are to be brought to management and Submitting department and not received or tested until issue is resolved.
847 | Atypical pH trend identified during chromatogram review for TSK lot CGF2B08B. | jacob.kallon@sanofi.com | 51 NYA Room 132 | 2022-12-02 | 2022-12-06 | It is expected that the pH trend be within specifications set forth in FBL-SOP-000788 ATTACHEMNT 5
846 | CZ SUB X-2250 Excessive Air Flow Sound Near Sparge Line | michael.knapp@sanofi.com | 8NYA 1403 | 2022-12-06 | 2022-12-06 | Sparge flows uninterrupted and without concern.
845 | Routine static environmental monitoring was not performed in the HA Fill (354,355, 358) area from 16Oct-22Oct | paul.consolmagno@sanofi.com | Biosurgery/76 NYA/HA Fill area 354, 355, 358 | 2022-10-22 | 2022-12-06 | FBL-SOP-000941, step 6.2.1
844 | 4 Late Nov PMs: 2453120, 2846359, 2846360, 2851310 | kelly.alanmurphy@sanofi.com | 74NYA, 51NYA | 2022-12-01 | 2022-12-06 | Per FBL-SOP-000081 Step 6.4.8: Complete all PM tasks prior to PM closure (otherwise approved by QA) or a deviation is required per site specific SOP.
843 | 832327/RF1QJC004- Hex Glycol below target in record but in range per GENDOC_021887 | alejandro.martinez2@sanofi.com | 8NYA Solution Prep, 1350 | 2022-12-05 | 2022-12-05 | Confirm discrepancy in record, update tolerance in record?
842 | TSH Part Supplier Deviation: GZ7082-01 Manifold 500mL & 100mL Sample Bags | andrew.ott@sanofi.com | 51 NYA | 2022-11-04 | 2022-11-04 | Open a supplier deviation to ensure correctly assembled parts are received in future
841 | 76NYA DSMB Rooms 406, 406A and 407 were not sampled for EM during November 2022. | Craig.Stevenson@sanofi.com | 76 NYA room 406, 406A, 407 | 2022-11-30 | 2022-12-05 | Per FBL-SOP-000277, it is expected that Grade D areas are sampled for EM on a monthly frequency.
840 | Pompe Engineering Seed Record Unable to Advance due Thaw Re-Execution | michael.knapp@sanofi.com | 8NYA 1403 | 2022-12-05 | 2022-12-05 | Record is able to advance and identify correct flask
839 | Accugenix account accessed by former employee | kyle.lebherz@sanofi.com | QC Micro | 2022-11-10 | 2022-12-05 | Former employee to have account deactivated.
838 | Unlabeled vials transported at sample pick-up | cleiton.lopes@sanofi.com | 8NYA drop off station to 68NYA 4fl Sample Management | 2022-12-03 | 2022-12-03 | Samples to be submitted with proper labeling/documentation and rejected in the event conditions are not met per steps 6.4.6
837 | Missing Bi-Weekly EM site CC1-117 in CC1-03 / CC1-07 | Craig.Stevenson@sanofi.com | 74NYA, Falcon | 2022-08-13 | 2022-12-05 | Per FBL-SOP-000604 step 6.1.5, EM sites in room CC1-07 are required to be sampled on a Bi-Weekly basis.
836 | Balance F-18255 stored in cabinet when not in use and moved to point of use when needed. | charlie.houle@sanofi.com | 76 NYA; Room 364 (HA Fermentation) | 2022-12-02 | 2022-12-05 | "As Founds" taken for the balance before it is moved and when it put back in place. (Metrology to provide standard information
835 | FZ SUB X2210 Perfusion and Weight Control Alarms | michael.knapp@sanofi.com | 8NYA 1333 | 2022-12-05 | 2022-12-05 | System is able to operate correctly without alarms.
834 | Pompe Engineering Run Thaw performed with incorrect vial | michael.knapp@sanofi.com | 8NYA 1343 | 2022-12-04 | 2022-12-04 | Batch Record requests the correct PN and thaw is performed with appropriate vial
833 | No Lot - No Product- 76NYA CRM - No EM activity level recorded for 76-424-33 on 20Oct2022 | Craig.Stevenson@sanofi.com | 76 New York Ave, Room 424 | 2022-10-20 | 2022-12-05 | Per FBL-SOP-000277, it is expected that Activity Level be documented for each EM site during sampling.
832 | QA did not record line clearance or equipment usage log | Cassandra.Lotito@sanofi.com | 74 NYA 1017 | 2022-12-01 | 2022-12-02 | It is expected all equipment usage and line clearance would be recorded per FBS-SOP-000383 step 6.1, FBS-SOP-000454, step 6.9.1
831 | Incorrect Post Test Selected | jeff.hill@sanofi.com | 45RM190 | 2022-12-03 | 2022-12-03 | Correct Post FIT program would be selected
830 | UFDF volume did not reach appropriate volume to trigger TUC to turn on | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1323 | 2022-12-02 | 2022-12-03 | Its expected that the parameters in the recipe will match actual volume
829 | UFDF Recipe Miscalculation | nathaniel.barnes@sanofi.com | 8NYA 1323 | 2022-12-02 | 2022-12-02 | MES communication fixed to populate correct calculations into the Delta V recipe
828 | Logbooks 2019QCC0034 and 2019QCC0021 pages missing | john.kapaniris@sanofi.com | QC Chemistry 45 NYA | 2022-11-22 | 2022-12-01 | All pages need to be accounted for
827 | 51 NYA Thyrogen Purification-Blue eluate lot CGF2B08 Storage not documented Per FBL-SOP-000371 | jacob.kallon@sanofi.com | 51 NYA room 131 | 2022-11-28 | 2022-12-01 | Per FBL-SOP-000371 step 6.4.1 "Record the location, temperature, time, initial/date, and complete check mark for within TOR entry on ATTACHMENT 1".
826 | UF Record Inspection Lot Canceled After Record Used | michael.daly@sanofi.com | 8 NYA | 2022-11-30 | 2022-11-30 | Records
825 | 8NYA-Fz- Automation intervention needed on SUB X2210 Perfusion EM prior to start of Perfusion | john.lytle@sanofi.com | Rm1333 | 2022-11-29 | 2022-11-29 | The perfusion countdown timer is currently set to begin immediately following Innoculation of the SUB, allowing for 24hrs to complete Perfusion setup. This timeframe should be reassessed to allow more time between Innoculation and Perfusion setup
824 | COVID + Gemba at 68 NYA | molly.pomeroy@sanofi.com | 68 NYA | 2022-11-29 | 2022-11-30 | It is expected that no employee tests positive for COVID while onsite
823 | 51 NYA- OneLims Collection Batch ID 20220433-1 negative control plates reversed | ian.odell@sanofi.com | 51 NYA Room 152A | 2022-11-30 | 2022-11-30 | Per SD-000463 step 4.1, collection for sample type / sample number should align with expected media (plate) type used.
822 | CZ Eluate SUM Bag Unable to Hold 21 Eluations | michael.knapp@sanofi.com | 8NYA 1403 | 2022-11-30 | 2022-11-30 | Bag is properly deployed and able to accept all samples and able to mix properly.
821 | 8NYA-room 1150-product impact unknown-GMP Critical dew point indicator and transmitter overdue for calibration | robert.bauch@sanofi.com | 8-1150 | 2022-12-01 | 2022-12-01 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been able to be collected
820 | 68NYA-Room 2700- QCC68-No as found data collected on Balance FB-2468 | robert.bauch@sanofi.com | 68-2700 | 2022-11-30 | 2022-12-01 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been able to be collected | Yes | Deviation
819 | 8NYA-CC8 room 1403-product impact unknown-FB-4095-029, FB-4095-032 on 500L Bioreactor overdue for calibration | robert.bauch@sanofi.com | 68-1403 | 2022-12-01 | 2022-12-01 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been able to be collected
818 | 8NYA-Room 1403-PF8 7 GMP Critical Instruments overdue for calibration | robert.bauch@sanofi.com | 68-1403 | 2022-12-01 | 2022-12-01 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been able to be collected
817 | Manway gasket for production reactors perpetually fail inspection prior to installation | paul.mclaughlinjr@sanofi.com | 74NYA | 2022-12-01 | 2022-12-01 | Manway gasket would be intact prior to being installed
816 | Gemba for Missed Water Level Check on CO2 Incubator in CBA | nicholas.koester@sanofi.com | QC Micro CBA | 2022-11-25 | 2022-11-28 | Per FBL-SOP-000368 Step 6.4, Verify incubator water level each week, 6-8 days from last recording. Step 6.4.1., Open the incubator and observe water levels in the bottom of the unit. Step 6.4.2., Verify the water levels between the front ridge and the small ridge/sensor
815 | 8NYA-Fz- leak observed on both ATF Harvest lines A&B located at the Pendotech connection | john.lytle@sanofi.com | Rm1333 | 2022-11-30 | 2022-11-30 | no leak, integral system
814 | QC micro plate count data validated in iLIMS by manufacturing | megan.botti-marino@sanofi.com | 8 NYA/68 NYA | 2022-11-11 | 2022-11-30 | Per 6.2.22 of FBL-SOP-000881, "Quality Control Review of Laboratory Data and Associated Data Systems," The review performs the following: After data packet has been reviewed and accepted, sign and date "Reviewed by" section and validate results in iLIMS, if applicable.
813 | Incorrect amount of stock created for UF cassette install | michael.daly@sanofi.com | 8 NYA 1323 downstream ballroom | 2022-11-29 | 2022-11-29 | It is expected 1 storage unit is created during stock creation
812 | 8NYA-Cz-MCC X2811 Chromatography cycles are out of sync | john.lytle@sanofi.com | 1434 | 2022-11-30 | 2022-11-30 | The MCC runs 3 column operations simultaneously and is expected to maintain a sequence of process operations suitable for forward processing
811 | Cz MMC pooling eluate in 200L barrel leaking | robert.morawski@sanofi.com | 8NYA RM1503 | 2022-11-30 | 2022-11-30 | Eluate bag would have no leaks.
810 | Missing Initial/Date on Balance Printout | joseph.napoli@sanofi.com | 68NYA QCM | 2022-11-08 | 2022-11-30 | analyst should have recorded initial, date, and SOP number on printout
809 | potential contamination for SDA lot 22GN01579 | Stefanie.Erickson@sanofi.com | QC Micro bioburden room | 2022-11-23 | 2022-11-23 | For there to be no growth on media within the packaging.
808 | Manway Gaskets are splitting while in Parts Washer | duane.messier@sanofi.com | 74NYA/CC1-07 | 2022-11-29 | 2022-11-29 | parts should not be damaged by washer cycle
807 | Missing entry on BIRF for storage of SP2 eluate CGF2B05A. | alyson.belleville@sanofi.com | 51 NYA | 2022-11-18 | 2022-11-29 | It is expected that there would be an entry on the BIRF for storing the eluate in a cold room within the allotted time out of refrigeration after it is removed from a cold room.
806 | SDA plates put in incorrect bag labeled TSA, resulting in incorrect incubation | Stefanie.Erickson@sanofi.com | QC Micro bioburden room | 2022-11-25 | 2022-11-29 | for the SDA plates to be in F-15593 and TSA plates to be in F-15594
805 | Fz- Lot CGHZE021 - Shift in tote weight for Zinc process Buffers | vraj.patel@sanofi.com | 8NYA Downstream Suite1 | 2022-10-26 | 2022-11-17 | There should be no shift in weight for the process buffer totes during ideal hold.
804 | Q Membrane Recipe Freeze | nathaniel.barnes@sanofi.com | 8NYA 1503 | 2022-11-29 | 2022-11-29 | Recipe continued to complete buffer drain phase and skid WFI rinse |  | No Deviation
803 | Correct Cycle Count on MMC Column | cory.lefort@sanofi.com | Suite 2 Downstream 8 NYA 1503 | 2022-11-29 | 2022-11-29 | The results were that there should be 40 cycle counts on the column per FBL-CST-000029. | Yes | No Deviation
802 | Barrel transfer setup was performed incorrectly. 832326, MO# RF1QIC002, ASM/8NYA/1350 | aidan.mula@sanofi.com | 8NYA/RM1350 | 2022-11-28 | 2022-11-28 | Gemba to determine if buffer can still be used |  | No Deviation
801 | 45NYA-Room 190-Thyrogen shared services-product impact unknown-F-1890 BSC had no as found data collected during an unscheduled repair and calibration | robert.bauch@sanofi.com | 45-190 | 2022-11-17 | 2022-11-29 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been able to be collected
800 | ASM UFDF Membrane install BOM missing p/n 714013 | jarad.kukla@sanofi.com | 1323 | 2022-11-25 | 2022-11-25 | 714013 should be included on the BOM
799 | AEX elution SUM was not released after recipe ended | michael.daly@sanofi.com | 8 NYA Room 1323 downstream ballroom | 2022-11-27 | 2022-11-27 | It is expected the recipe releases the equipment after completion
798 | Recipe went into hold multiple times due to pump communication errors | michael.daly@sanofi.com | 8 NYA 1323 Downstream Ballroom | 2022-11-27 | 2022-11-27 | It is expected the pump starts during the adjust
797 | Recipe went into hold high high retentate pressure alarms during the WFIA flush post membrane install. | micahel.daly@sanofi.com | 8 NYA | 2022-11-25 | 2022-11-25 | It is expected the WFI flushes
796 | AEX Regeneration buffer temperature was out of specification during use | michael.daly@sanofi.com | 8 NYA Room 1323 downstream ballroom | 2022-11-28 | 2022-11-28 | It is expected the regeneration buffer is 16-24 deg C during use
795 | Incorrect version of logbooks used | alvaro.rueda@sanofi.com | 74 NYA Falcon | 2022-06-08 | 2022-11-28 | Communication of new version of logbook available, old version of logbook removed from use and new version of logbook obtained.
794 | Butyl Eluate SUM agitator did not turn on during mixing recipe | michael.daly@sanofi.com | 8 NYA room 1323 downstream ballroom | 2022-11-28 | 2022-11-28 | It is expected the agitator turns on during the mixing recipe
793 | Multiple automation issues encountered during downstream ASM PPQ operations (Olipudase – CGHAM001 – 8 NYA – Short Description of Problem) | michael.daly@sanofi.com | 8 nya room 1323 downstream ballroom | 2022-11-27 | 2022-11-27 | It is expected the automation and equipment runs as intended
792 | Material 713975 was assembled incorrectly by Supplier | bailey.gleason@sanofi.com | 8NYA RM1350 | 2022-11-28 | 2022-11-28 | Hold Gemba for supplier deviation and request new part.
791 | Operator witnessed a step that they performed. | alvaro.rueda@sanofi.com | 74 NYA Falcon | 2022-11-12 | 2022-11-28 | Separate individuals perform and witness work performed.
790 | During the set-up to transfer a buffer at 8NYA, a small leak was detected from the bottom agitator portion of the SUM bag. | thomas.loik@sanofi.com | 8RM1350 | 2022-11-25 | 2022-11-25 | Per FBL-EWI-000019, that the installed SUM bag will not leak.
789 | FZ SUB X-2210 TCU Not Reading Properly on DeltaV | michael.knapp@sanofi.com | 8NYA 1333 | 2022-11-26 | 2022-11-26 | TCU operates correctly.
788 | ASM CEX Eluate FIT Not In Delta V | stanley.durand@sanofi.com | 8NYA rm 1290 | 2022-11-26 | 2022-11-26 | Correct filter to be in Delta V
787 | Load cells on 500L SUB are producing erratic readings during media load | paul.mclaughlinjr@sanofi.com | 74NYA | 2022-11-26 | 2022-11-26 | Load cells would accurately show mass as it is added to SUB
786 | TCU of AEX Load SUM Not Chilling | stanley.durand@sanofi.com | 8NYA room 1323 | 2022-11-25 | 2022-11-25 | It is expected that the CEX eluate/AEX load SUM to chill following eluate transfer filtration between 2-8 degree C.
785 | FZ SUB No Base Buffer Available for Setup | michael.knapp@sanofi.com | 8NYA 1333 | 2022-11-26 | 2022-11-26 | Buffer is available.
784 | ASM Eluate SUM Profibus Cable Damaged | michael.knapp@sanofi.com | 8NYA 1333 | 2022-11-24 | 2022-11-25 | Cable is not damaged.
783 | LAL WFI Sample 517941 Tested Twice | Rebecca.Schmidt@sanofi.com | QC Micro Nexus bench | 2022-11-24 | 2022-11-25 | WFI sample 547363 should have been retested instead of WFI sample 517941
782 | CZ Bioreactor X2250 Offline Sample Result and DO Probe Drift | michael.knapp@sanofi.com | 8NYA 1403 | 2022-11-25 | 2022-11-25 | Probe readings and offline samples are close enough to not warrant concern. There is minimal drift between DO probes.
781 | Missing initial pH for Media Lot CGFMP014 | daniel.bati@sanofi.com | 51NYA | 2022-11-25 | 2022-11-25 | Final pH 6.8 to 7.2
780 | Bubble trap sensor failed causing bubble trap value to over flow, wetting the filter | Pavel.dobrovolskiy-wynes@sanofi.com | 8RM1323 | 2022-11-24 | 2022-11-24 | Its expected that recipe will stop filling the bubble trap once hi sensor is seeing liquid.
779 | Insufficient volume submitted for COVID Bioburden PSFM Media | Pushti.Shah@sanofi.com | 68 NYA/ QC Micro/ Bioburden room | 2022-11-23 | 2022-11-23 | For there to be 200mL of volume readily available per sampling plan.
778 | ASM filter label not scanning | Christian.Venegas@Sanofi.com | 8NYA upstream Room 1333 | 2022-11-23 | 2022-11-23 | Get HU fix to be able to be consumed in the record
777 | Sample labels and the test requests are different for 2022-45CV-0026 and 2022-45CV-0030 | Subhashini.Satishkumar@Sanofi.com | 45 NYA | 2022-10-29 | 2022-11-23 | It is expected that the results of CVs match with the corresponding CVs.
776 | growth outside of the filter on negative control R2A media | Stefanie.Erickson@sanofi.com | QC Micro bioburden room | 2022-11-17 | 2022-11-23 | for the media to have no growth
775 | 8NYA ASM Pooling SUM Recipe Error | michael.knapp@sanofi.com | 8NYA 1333 | 2022-11-22 | 2022-11-22 | Recipe progresses correctly.
774 | X2311 in hold due to multiple alarms. pCO2 is reading around 200mmHg, Nitrogen sparge is open 100% causing foaming, which is wetting out exhaust filters. Both vessel pressure and vessel O2 is oscillating | paul.mclaughlinjr@sanofi.com | 74NYA | 2022-11-22 | 2022-11-22 | Reactor readings would be stable | Yes | Deviation
773 | Equipment not recorded and step not verified on HA fill packet | megan.botti-marino@sanofi.com | 76 NYA | 2022-08-16 | 2022-11-22 | All fields be filled in and not left blank. That verified by steps are filled out by a qualified verifier and not left blank.
772 | 76 NYA COVID positive employee in Biosurgery controlled manufacturing space. | kenneth.morgan@sanofi.com | 76 NYA, multiple rooms | 2022-11-09 | 2022-11-22 | It is expected that personnel would not enter classified manufacturing space when they have an illness per FBS-SOP-000667.
771 | No initial WFI. Recipe went straight to Mix time. | alex.tello@sanofi.com | 8NYA RM1350 | 2022-11-22 | 2022-11-22 | path forward
770 | Door to refrigerator RF-8161001 was found left open and in alarm. | Stanley.Durand@sanofi.com | Room 1410 at 8 NYA | 2022-11-21 | 2022-11-21 | Samples should remain within the required temperature range.
769 | Door to refrigerator RF-8161001 was found left open and in alarm. | Stanley.Durand@sanofi.com | Room 1410 at 8 NYA | 2022-11-21 | 2022-11-21 | Samples should remain within the required temperature range.
768 | Door to refrigerator RF-8161001 was found left open and in alarm. | Stanley.Durand@sanofi.com | Room 1410 at 8 NYA | 2022-11-21 | 2022-11-21 | Samples should remain within the required temperature range.
767 | Tote Identification for 832328/RF1QKC001 calls for chilled tote, but BOM doesn't call for chilled tote bag | alejandro.martinez2@sanofi.com | 8NYA CNC | 2022-11-22 | 2022-11-22 | Path forward whether these buffers need to be chilled or not?
766 | Incorrect incubator selected in ONELIMS for Liquid Bioburden Batch | arpan.christian2@sanofi.com | 68 NYA, QC Microbiology, 4700 | 2022-10-24 | 2022-11-21 | It is expected that the correct incubator (F-15594) should have been selected for Bioburden sample type
765 | ASM Eluate Pooling SUM not controlling to temperature | michael.knapp@sanofi.com | 8NYA 1403 | 2022-11-19 | 2022-11-21 | SUM is set to correct temp control.
764 | ASM Butyl Column Bottom Mobile Phase piping bent during column transport | robert.morawski@sanofi.com | 8NYA 1323 | 2022-11-21 | 2022-11-21 | Column transport is completed with no accidents.
763 | 8NYA-ASM-650L Eluate SUM 4010010 weight found to be incorrect, reading zero prior to handoff to Downstream operations | john.lytle@sanofi.com | Rm1333 | 2022-11-19 | 2022-11-21 | the SUM should maintain an accurate reading of the weight of Eluate during use
762 | Pages missing from logbooks 2021QCT0049 and 2021QCT0050 | jacqueline.legere@sanofi.com | 68NYA Rm 1825 | 2022-11-21 | 2022-11-21 | Log books 2021QCT0049 and 2021QCT0050 should contain 53 pages total
761 | 8NYA ASM SUB Media Disposal/take-out not performed in batch record | michael.knapp@sanofi.com | 8NYA 1333 | 2022-11-20 | 2022-11-21 | Media is taken out against batch record and media disposal is performed.
760 | VOID - practice GEMBA ONLY | kevin.southwick@sanofi.com | 45 NYA / | 2022-11-21 | 2022-11-21 | Per SOP, air handler cannot be down for more than 30 minutes, per step 6.2.3
759 | EM Site 220 in room 357 not sampled Weekly per FBL-SOP-000941 | Craig.Stevenson@sanofi.com | 76 New York Ave, Room 357 | 2022-11-21 | 2022-11-21 | Per FBL-SOP-000941 Step 6.2.1, Site 220 is expected to be sampled for EM on a weekly basis. | Yes | Deviation
758 | 8NYA Cell Bank Cryofreezer Communication Loss | michael.knapp@sanofi.com | 8NYA 1270 | 2022-11-19 | 2022-11-21 | Cryostat does not experience communication loss.
757 | Status was changed on a Downflow booth that was OOS and operations were performed. 760726, MO# RF608C010 Cerezyme/8NYA 1350 | aidan.mula@sanofi.com | 8NYA, RM1350 | 2022-11-20 | 2022-11-20 | TBD at Gemba
756 | Record for 832319, MO# RF1Q9C001, ASM/8 NYA/RM1350 is asking for a chilled tote but not a recirculation bag | aidan.mula@sanofi.com | 8NYA, RM1350 | 2022-11-20 | 2022-11-20 | Determination whether the bag is incorrect or the tote and path forward |  | No Deviation
755 | Record does not have Handling Unit reference comment to material used that is not on the BOM. 760741, MO# RF1BHC108 Cerezyme | aidan.mula@sanofi.com | 8NYA, RM1350 | 2022-11-19 | 2022-11-19 | Determination at Gemba
754 | TSH Lot CGFSB013, Unicorn Method for Wave Reactor B ended inadvertently, Process is now being run in Manual mode | brian.mcgee@sanofi.com | 51NYA/Rm 152A | 2022-11-19 | 2022-11-19 | It is expected for the Method to continue running until ended by the operator at the end of the run or end of the stage
753 | Multiple DP alarms for Purification areas at 51 | alyson.belleville@sanofi.com | 51 NYA, rooms 134/135, 136/139, and 88/87. | 2022-11-18 | 2022-11-18 | It is expected that there should be no alarms if there are no issues, and if there are alarms that the issue would be addressed.
752 | Incorrect version of FBL-EWI-000005 effective in EWI server | Amy.McIntyre@sanofi.com | 8 NYA Facility | 2022-11-15 | 2022-11-17 | Per BQ-GOP-000292 , that FBL-EWI-000005 would have the same effective version in both the EWI server and in GEODE+/Gibraltar.
751 | Bubble trap sensor failed causing bubble trap filter to overflow. | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA1503 | 2022-11-18 | 2022-11-18 | Its expected that bubble trap sensor works appropriately and is able to sense liquid vs air.
750 | 8NYA Upstream Media tote weight drift prior to use | michael.knapp@sanofi.com | 8NYA 1403 | 2022-11-16 | 2022-11-17 | Media Tote load cells accurately provide mass loss for perfusion calculations.
749 | ANMS call out service interruption | Shiful.Kohli@sanofi.com | site wide | 2022-11-16 | 2022-11-16 | TBD
748 | Sample Tee (776567) not documented per 2022FRACC0062 | jarad.kukla@sanofi.com | 8NYA | 2022-09-18 | 2022-11-17 | Sample Tee information should have been recorded in MBR (p/n, lot number, expiration) and reference made to 2022FRACC0062, per SD-FRA-22FRACC0062.
747 | X-3051 (MMC skid) experienced a Hi Hi pressure alarm (PI-305012A) causing Capto MMC production recipe to go into hold during sanitization phase 1. | craig.valery@sanofi.com | 74 NYA, PU1-09 | 2022-11-17 | 2022-11-17 | It is expected that the recipe would continue without any hold conditions.
746 | Daily SporKlenz cleaning missed for 16NOV22 for action notice | paul.mclaughlinjr@sanofi.com | 74NYA | 2022-11-16 | 2022-11-17 | Daily cleaning would be performed per action notice
745 | Part 808713 is needed for batching 805015 | richard.manasian@sanofi.com | 1350 at 8 NYA | 2022-11-17 | 2022-10-28 | BOM would have part 808713 instead of part 692315
744 | Entered incorrect wave run in vi-cell sample data. Unable to fix or add a comment in vi-cell. No guidance in SOP on how to resolve this type of error | adam.rowe@sanofi.com | 74NYA, CC2-01 | 2022-11-17 | 2022-11-17 | Correct wave run should be entered
743 | 8NYA CZ Harvest SUM agitation noise | michael.knapp@sanofi.com | 8NYA 1403 | 2022-11-17 | 2022-11-17 | Agitation continues without interruption. |  | No Deviation
742 | 832318 buffer calling for chilled tote but should be ambient | richard.manasian@sanofi.com | 1350 at 8 NYA | 2022-11-17 | 2022-11-17 | Tote chill step and associate tote checks in MBR |  | No Deviation
741 | Bag leak and WFI over fill for 760741-RF1BHC112 | richard.mansian@sanofi.com | 1350 at 8 NYA | 2022-11-17 | 2022-11-17 | Value should run under recipe and close WFI valve as intended
740 | During the initial WFI fill portion of the Formulation recipe, interference from the nightly sani recipe put the buffer recipe into hold but did NOT shut the WFI valve. | richard.manasian@sanofi.com | 8RM1350 | 2022-11-17 | 2022-11-17 | That a recipe that goes into hold will shut all related activities associated with the recipe.
739 | The CNC transfer SFO for 832318 was set up as a chilled transfer for an ambient buffer | richard.manasian@sanofi.com | 8RM1202 | 2022-11-16 | 2022-11-16 | There was buffer specific document listed in the ASM control document (FBL-CST-000032) and it could not be verified if the buffer should be chilled or ambient. If the buffer should be ambient, then the BOM is correct and multiple steps in the EBR are not. Conversely, if this buffer should be chilled then the BOM is incorrect.
738 | Raw materials were mistakenly brought to the cleaning staging area | santiago.noriega@sanofi.com | 80 Warehouse | 2022-11-11 | 2022-11-16 | TBD
737 | 8NYA-ASM-unable to install new eluate SUM bag due to SUM Setup step logic in MES | john.lytle@sanofi.com | Rm1333 | 2022-11-16 | 2022-11-16 | MES loop that accommodates the need for emergency replacement of SUT without intervention
736 | 8NYA Gemba SUB X2220 Media Install not complete | jeffrey.boldizar@sanofi.com | 8NYA | 2022-11-15 | 2022-11-16 | Steps should be executed during time of completion
735 | Drug Product shipment left on loading dock at room temperature from 11/14/22 through 11/16/2022 | cleiton.lopes@sanofi.com | 45 New York Ave | 2022-11-14 | 2022-11-16 | Upon clearance, temp tales will be referenced to verify temperature range
734 | 100L SUB went into hold due to pCO2 probe deviation | paul.mclaughlinjr@sanofi.com | 74NYA | 2022-11-16 | 2022-11-16 | 7
733 | Drug Product shipment left on loading dock at room temperature from 11/14/22 through 11/16/2022 | cleiton.lopes@sanofi.com | 45 New York Ave | 2022-11-14 | 2022-11-16 | Upon clearance, temp tales will be referenced to verify temperature range
732 | 8NYA-ASM-Harvest SUM TCU not functioning properly | john.lytle@sanofi.com | Rm1333 | 2022-11-16 | 2022-11-16 | TCU controlling temperatures as expected
731 | 51 NYA- Thyrogen- Missing Dissolved Oxygen probe information in BRX prep record (CGF4B07) | ian.odell@sanofi.com | 51 NYA Room 140 | 2022-06-13 | 2022-11-16 | Per FBL-SOP-000718, the DO baseline number needs to be recorded from the transmitter during calibration.
730 | Environmental Monitoring Site 53 in room 224 was not collected biweekly | Craig.Stevenson@sanofi.com | 76 New York Ave, Room 224 | 2022-11-12 | 2022-11-16 | Per FBL-SOP-000941 step 6.2.1, EM site 53 should be collected on a biweekly basis.
729 | Media EM Low-Low alarm for rate of change on X2250 | Brendan.Dickhaut@sanofi.com | 8NYA, Room 1403 | 2022-11-15 | 2022-11-15 | Expected result is that media rate of change stays within the set High and Low level alarm range
728 | Wooden Pallets at 76NYA Cooridor | rebecca.Cunningham@sanofi.com | 76 NYA | 2022-11-15 | 2022-11-15 | Per FBL-SOP-000661 step 6.4.6 pallets are to be re-palletized, and wood is not to enter 76NYA
727 | Expired IPA used for routine EM monitoring of the HA fill suite | megan.botti-marino@sanofi.com | 76 NYA | 2022-08-25 | 2022-11-15 | It is expected that reagents/cleaning agents used for GMP work be within expiry.
726 | Missing documentation- quarterly walkthrough (Q2 2022) for learning and development | erin.young@sanofi.com | N/A | 2022-11-15 | 2022-11-15 | Per BQ-GOP-000260 step 5.2.1, department management for each functional area related to GMP activities ensures a walk-through inspection of their area is completed quarterly. Per step 5.2.3 walk-through inspector reviews the previous quarter's walk-through inspection documentation before conducting the current walk-through inspection of the designated area. Per step 5.7.1 department management conduct a review, evaluation, and approval of the walk-through inspection documentation within 30 calendar days of the walk-through inspection.
725 | Overflowing and leaking Endotoxin sample 220049683(CV-0732) submitted wrapped in parafilm | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2022-11-14 | 2022-11-14 | TBD
724 | Totes OOS, need to be worked on to be put back into service. | Sarah.flattich@sanofi.com | 8 NYA CNC location | 2022-11-15 | 2022-11-15 | To have the totes worked on and put back into service for production.
723 | Personnel Monitoring Negative Control Rodac plate not Submitted for Spread Plate Batch | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2022-07-27 | 2022-08-03 | Per Attachment 3 of FBL-SOP-001327 a negative control for the rodac plates must be submitted. | Yes | Deviation
722 | 500L SUB setup inflated and Pressure held without temp probe | adam.hildebrand@sanofi.com | 74 NYA, CC2-01 | 2022-11-13 | 2022-11-14 | It is expected that the temp probe is installed per FBL-SOP-001444 step 6.2.2.8. See attached snip from SOP.
721 | No reportable results for 76 NYA EM sites 227-2, 229-2, 229-3, 229-4 and 3102-2 | marissa.lentoni@sanofi.com | 76 NYA Room 227, 229, 3102 | 2022-08-10 | 2022-11-14 | Sites would be sampled biweekly per FBL-SOP-000941 upon implementation on 10Aug2022.
720 | No filter recipe available for record RF649C026 761222 690770 FF. | bailey.gleason@sanofi.com | 8NYA | 2022-11-14 | 2022-11-14 | N/A
719 | Rollup MAL door 1701 is stuck closed | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1701 | 2022-11-13 | 2022-11-13 | Its expected that MAL door will operate normally
718 | 51 NYA rHTSH Purification individually wrapped propylene not available for sampling. | jacob.kallon@sanofi.com | 51 NYA rooms 88 and 132. | 2022-11-14 | 2022-11-14 | 50 mL, individually wrapped polypropylene sample tubes are required for A280 buffer blank samples per FBL-MBR-000109 and Myco blanks per FBL-MBR-000051.
717 | 1 Logbook and 4 74NYA Batch Records Not Located during Q3 Logbook & Batch Record Audits | jennifer.lopez@sanofi.com | 74 NYA | 2022-09-30 | 2022-10-05 | Per FBL-SOP-000681 and FBL-SOP-000441; audits are performed to account for all outstanding records.
716 | Activity sample 380348 is unable to be found in the QC Chemistry lab | paul.consolmagno@sanofi.com | 45 NYA, QC Chemistry | 2022-07-27 | 2022-11-19 | OneLIMS Sample 380348 would be available for testing | Yes | No Deviation
715 | For Cerezyme buffer 760729 MO#RF611C009 both the pH and Conductivity measurements are OOS | thomas.loik@sanofi.com | 8RM1350 | 2022-11-13 | 2022-11-13 | That the pH and Conductivity will be within process ranges.
714 | EMS Call out signals missing for incubators (X4070, X4069, X4071) at 74NYA | adam.rowe@sanofi.com | 74NYA, CC1-18, CC1-21 | 2022-11-10 | 2022-11-10 | Incubators should have EMS call outs
713 | 51 NYA- Thyrogen- Wave 12L Day 1 offline reading out of range pCO2 | ian.odell@sanofi.com | 51 NYA Room 152A | 2022-11-12 | 2022-11-12 | Offline pCO2 result should be between 30-100mmHg.
712 | Failed Visual Inspection - T-111 | nicholas.shoenfelt@sanofi.com | 51-124 | 2022-11-11 | 2022-11-11 | It is expected that T-111 would pass visual inspection (no particulates)
711 | Bioburden and TOC tests needed, LAL tested instead of TOC | Rebecca.Schmidt@sanofi.com | QC Micro LAL bench | 2022-07-19 | 2022-11-11 | FBL-SOP-000189 Attachment 1 Bottom of Table | No | Deviation
710 | 8NYA Bioreactor excursions following site power outage | michael.knapp@sanofi.com | 8NYA 1403 | 2022-11-10 | 2022-11-10 | SUB continues to perform without excursions.
709 | MCC X-2811 bubble trap sensor alarm | michael.knapp@sanofi.com | 8NYA 1403 | 2022-11-10 | 2022-11-10 | The MCC requires constant buffer flow and held once it read low flow as a result of running out of buffer.
708 | Phenyl Eluate Recovery Below Normal Range (80.298%) - Normal Range >85%! | stanley.durand@sanofi.com | 8NYA room 1503 | 2022-11-05 | 2022-11-11 | It is expected that the Phenyl eluate recovery calculation to be at or greater than 85%.
707 | Missing IQ Report ID for F-17553 Mettler Toledo Floor scale in 51-138 | damian.barbato@sanofi.com | 51 NYA Room 138 | 2022-11-10 | 2022-11-10 | IQ Report ID for F-17553 Mettler Toledo Floor scale in 51-138 could be found
706 | Cerezyme media osmolality failing. PN: 760741, MO# RF1BHC106 | aidan.mula@sanofi.com | 8NYA, RM1350 | 2022-11-10 | 2022-11-10 | Determination on viability of the buffer
705 | Stability used approved Attachment 1, Stability Sample Transaction Form from FBL-SOP-001262, Stability Sample Inventory but not effective yet. | Aldona.Szostek-Pontchek@sanofi.com | 80NYA stability chamber area | 2022-09-16 | 2022-11-10 | An effective document would be used for transactions. | Yes | Deviation
704 | pratice GEMBA - accidentally clicked submit 0 please cancel | elizabeth.cook@sanofi.com | 45NYA | 2022-11-08 | 2022-11-10 | pratice GEMBA - accidentally clicked submit 0 please cancel | No | No Deviation
703 | TSH- pH Probes Not Adjusted with a Difference Greater Than 0.05 From Offline BGA | christopher.holmes@sanofi.com | 51-140 | 2022-06-24 | 2022-11-09 | It is expected that the bioreactor probes would be adjusted to match the offline BGA readings if a difference of > 0.05 was calculated.
702 | Sartorius Welder 2009968 Deactivated While Used | erica.davis@sanofi.com | 8 NYA | 2022-11-08 | 2022-11-10 | N/A
701 | 45 NYA MAHU-1/MAHU-2 unexpected shutdown | Jason.silva@sanofi.com | 45 NYA | 2022-11-09 | 2022-11-09 | MAHUs would be operating properly
700 | 10-hour mix timer exceeded on buffer batch RF543C004 | Brendan.Dickhaut@sanofi.com | 8NYA, Room 1350 | 2022-11-10 | 2022-11-10 | Expect results are that buffer batches are completed within the allotted time of 10 hours from initial WFI fill | Yes | No Deviation
699 | Record calls for SUT that is incompatible with small scale media transfer set up for 200L batches / EWI does not detail media cart transfer for 200L batch. | bailey.gleason@sanofi.com | 8NNYA RM 1350 | 2022-09-24 | 2022-11-09 | Change of EWI to specify 200L media cart transfers.
698 | Step witnessed by performer | alvaro.rueda@sanofi.com | 74 NYA Falcon | 2022-09-02 | 2022-11-09 | Second operator should have signed Witnessed by step
697 | 74 NYA-Room CC1-13 was used and stop signs were removed without QA release of Action Notice 2022-C-068 | timothy.pyne@sanofi.com | 74 NYA CC1-13 | 2022-10-31 | 2022-11-09 | it is expected that the action notice would have been approved and released prior to use of the area.
696 | Improper fitment of TFF liner upon installation (P/N 797020) (HU's 1001048843, 1001039844, 1001045993) | alvaro.rueda@sanofi.com | 74 NYA Falcon | 2022-11-08 | 2022-11-08 | Liner should fit as intended
695 | Incorrect version of FBL-FRM-000796 used multiple times | megan.botti-marino@sanofi.com | 8NYA, 68 NYA | 2022-09-01 | 2022-11-08 | per 5.1.1 of Good Documentation Practices (BQ-SOP-000269) "When using a controlled document to record data, confirm that the most current effective version of the document (e.g. form or log sheet) is being used. Per 5.1.1.1 of BQ-SOP-000269 "Forms should be printed on the day of intended use." | Yes | Deviation
694 | Leak identified on post- ATF guard filter | jeffrey.boldziar@sanofi.com | 8 NYA 1403 | 2022-11-08 | 2022-11-08 | expected that the single use tubing remain integral throughout the process
693 | Thyrogen Stability Sample Tested at Incorrect Volume for Bioburden | krystal.masseri@sanofi.com | QC micro bioburden room | 2022-06-27 | 2022-11-08 | Stability samples should be tested at 10mL volume
692 | incorrect volume tested for falcon bioburden | Rebecca.Schmidt@sanofi.com | QC micro bioburden room | 2022-10-22 | 2022-11-08 | expected results for 100ml to be tested for both TSA and SDA | Yes | Deviation
691 | Pump tubing for SUM ruptured during transfer from small scale SUM to 25L bags. | kelvin.leo@sanofi.com | 8 NYA | 2022-11-08 | 2022-11-08 | Pump tubing stays integral | Yes | No Deviation
690 | 8NYA Gemba MCC Skid Hex Glycol buffer tote tube welding failed | lolav.ahmed@sanofi.com | 8NYA | 2022-11-08 | 2022-11-08 | N/a
689 | Pump tubing leak on Cz Zinc WFI dilution | zachary.rickard@sanofi.com | 8NYA/1503 | 2022-11-08 | 2022-11-08 | Pump tubing should not leak or get cut during pumping operation.
688 | X-2230 500 SUB , recipe holding due to spike on pCO2 probe A. | otto.cordon@sanofi.com | 74NYA / CC2-01 | 2022-11-08 | 2022-11-08 | Per FBL-CST-000030 page 43, Dissolved CO2 Concentration should operate ≤ 120 mmHg
687 | 805592 failed to charge to SUM | richard.manasian@sanofi.com | 1350 at 8 NYA | 2022-11-07 | 2022-11-07 | Hex glycol would charge SUM to required amount automatically without error
686 | 8NYA-ASM-Setpoint deviation alarm on X2260 pC02 probes resulting in process hold | john.lytle@sanofi.com | Rm1333 | 2022-11-06 | 2022-11-07 | SUB pC02 readings are checked daily using the BGA. When off line results vary by more than 10mmHg from the BGA reading, Mfg performs a recalibration of the probes per FBL-EWI-000010 for alignment
685 | 45 NYA Dynamic Monitoring RODAC Negative Control Mold Recovery | elizabeth.couturier@sanofi.com | 45 NYA | 2022-10-27 | 2022-11-05 | it is expected that negative control plates do not have growth. | Yes | Deviation
684 | 8NYA ASM Bleed Setup SUT not included within BOM | shawn.nguyen@sanofi.com | 8NYA Room 1333 | 2022-11-07 | 2022-11-07 | Additional SUT added to BOM. All items to be accounted for.
683 | 8 NYA MCC Skid X2510 0.1M NAOH, 50% Hex Glycol offline Cond sample off SPC | Lolav.ahmed@sanofi.com | 8 NYA skid in room 1333 | 2022-11-07 | 2022-11-07 | FBL-CST-000032, FBL-EWI-000010, FBL-EW-000022, & FBL-EWI-000032
682 | Capto MMC HIHI pressure hold during equilibration cycle 2 | craig.valery@sanofi.com | 74 NYA, PU1-09 | 2022-11-07 | 2022-11-07 | It is expected that the recipe would continue without any hold conditions.
681 | Media Flow Low-Low alarm on X2260 | Brendan.Dickhaut@sanofi.com | 8NYA, Room 1333 | 2022-11-06 | 2022-11-06 | Expected result is that RV swiches to 2 when required without intervention
680 | Pin hole leak in Q Membrane filter | cory.lefort@sanofi.com | 1503 phase 2 | 2022-11-06 | 2022-11-06 | FBL-CST-000029
679 | BioLims Batch 202203898 INC-03683 / Arrived after Expiry | michael.aiello@sanofi.com | Sample receipt room / QC Micro | 2022-11-05 | 2022-11-05 | For the plates to be within expiry at time of receipt
678 | Not enough volume to test for bioburden | Stefanie.Erickson@sanofi.com | QC Micro bioburden room | 2022-11-05 | 2022-11-05 | for there to be enough volume to test for bioburden
677 | GEMBA: 74NYA: Pro Active approach if HI alarm happens | William.Bradley@sanofi.com | 74NYA PU1-09 | 2022-11-06 | 2022-11-06 | Process should not exceed the HI HI alarm limit of 30psi
676 | Nano Guard filter change out with out QA observation | devin.parks@sanofi.com | 8 NYA / 1343 | 2022-11-04 | 2022-11-04 | QA should be observe guard filter change out
675 | SARS CoV2 – Lot (or “Multiple Lots”) – 74NYA 2022-CV-0660 returned bioburden recovery of 42 CFU/200mL which fails the criteria for FBL-SOP-000397 for final release. | katherine.milliken@sanof.com | CC1-13 CC1-14 | 2022-10-29 | 2022-11-04 | <10 CFU/100mL is the acceptance criteria FBL-TSP-000333 | No | Deviation
674 | HA - CBYHAS006 - 76 NYA Room 355 - HA Aspect fill was not observed for the required duration | Juan.Ramos-Torres@sanofi.com | 76 NYA Room 355 | 2022-06-29 | 2022-11-04 | HA Aspect fill must be observed at least 2 hours
673 | Fz Nano Pressure Spike during Equilibration | nathaniel.barnes@sanofi.com | 1323 | 2022-11-04 | 2022-11-04 | Equilibration to complete successfully with no pressure spikes |  | Deviation
672 | Small hole discovered in the phenyl Load SUM bag | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1503 | 2022-11-04 | 2022-11-04 | Its expected that SUM bags will come from manufacturer without any damage
671 | Gemba for Incorrect CX values in Mod Enzyme Prep Record MO#RF618C019/ Cerezyme/ 8NYA | aidan.mula@sanofi.com | 8NYA/RM1350 | 2022-11-04 | 2022-11-04 | Correct CX values or otherwise be allowed to continue with batching operations
670 | Falcon CV with Counts/Rinse volume not Recorded at Time of Bioburden Testing | jessica.kychler@sanofi.com | QC Microbiology, 68 NYA floor 4 room 4700 | 2022-10-29 | 2022-11-04 | It is expected that CV bioburden would have 0 cfu and that rinse volumes will be recorded and verified prior to testing
669 | 74NYA-Room PU1-05-PF74-product impact unknown- 2 Temperature indicators FB-3668-002 and FB-3672-002 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 74-PU1-05 | 2022-11-03 | 2022-11-04 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
668 | 8NYA-AMS-damaged Aseptikwik X membrane found after A2B Straight Tube set was installed on retentate elbow | john.lytle@sanofi.com | Rm1333 | 2022-11-03 | 2022-11-03 | A successful Aseptikwik X connection results in no damage to membrane during removal, and a pre-use inspection of the membranes ensures there is an integral system ready to receive the connection
667 | INC 03658 - Environmental Monitoring Bioburden plates expired before receipt into QCM Lab | cleiton.lopes@sanofi.com | 68 NYA, QCM, Sample receipt | 2022-11-02 | 2022-11-03 | Samples would be delivered with 24 hours.
666 | CU RO for 68, 45, 51 NYA limits for COND not configured correctly in ONELIMS | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2022-11-03 | 2022-11-03 | It is expected that limits in our SOPs are aligned with limits configured in OneLIMS
665 | Missing Q2 2022 Quarterly Nitrate sample for 51-WFI-9 | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2022-07-01 | 2022-11-03 | It is expected that 51-WFI-9 be collected every quarter for Nitrates | Yes | Deviation
664 | P/N 692174 3K Tote bag leak. | alex.tello@sanofi.com | 8 NYA | 2022-11-03 | 2022-11-03 | Path forward. |  | No Deviation
663 | Creedo Cooler Over Capacity During Sample Drop Off | Robert.petersjr@sanofi.com | 68NYA Sample Drop off Room, 4th floor | 2022-11-02 | 2022-11-02 | Per FBL-SOP-000172, Step 6.4.9, the amount of liquid in the Creedo will not exceed 750ml.
662 | BioLims Sampling Issue | katherine.men@sanofi.com | 8 NYA , 1323 | 2022-11-02 | 2022-11-02 | per FBL-RPT-000376 2x 20mL qc process impurities retains and 3x 2mL qc retains samples should be taken at the UF concentrate sampling point.
661 | CX value missing from MBR 760739/ Timer exceeded | richard.manasian@sanofi.com | 1220 cold room and CNC area | 2022-11-02 | 2022-11-02 | The mod enzymes should be able to be scanned into the MBR without adding Cx value
660 | 8NYA-ASM-Backup Vent filter wetted (Primary wetted previously) | john.lytle@sanofi.com | Rm1333 | 2022-11-03 | 2022-11-03 | functional vent filter |  | No Deviation
659 | 74NYA-Room UT2-06-Falcon-on AHU 8131-product impact unknown-HEPA Banks FB-3106 and FB-3107 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 74NYA-Room UT2-06 | 2022-11-02 | 2022-11-03 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
658 | Expired bioburden rhTSH Q load sample | kathryn.lentoni@sanofi.com | 68NYA 4700, QCM | 2022-11-02 | 2022-11-02 | Sample would be submitted within an appropriate timeframe for QC to perform bioburden testing. | Yes | Deviation
657 | Not Fully Chilled WFI Was Used During 1 to 1 WFI Dilution During Cz Zinc Production | stanley.durand@sanofi.com | 8NYA rm1503 | 2022-11-02 | 2022-11-02 | It is expected that chilled WFI is used during 1to1 WFI dilution per FBL-EWI-000033.
656 | No Products – No Lots - 68NYA, QC Microbiology - RQ for CO2 Incubator 75-QC-0115 (F-16140) was not executed by 30Sep2022. | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2022-11-01 | 2022-11-01 | It is expected that the CO2 Incubator have the required equipment in order for the RQ be completed on time.
655 | Loss of AHU in PU1-20, 21, 22, 23, 24, PU2-01 | craig.valery@sanofi.com | 74 NYA PU1-20, 21, 22, 23, 24, PU2-01 | 2022-11-02 | 2022-11-02 | It is expected AHU system operates with out loss
654 | 8NYA-ASM-needed automation intervention to prime base line, as it was owned by SUB pH controller | john.lytle@sanofi.com | Rm1333 | 2022-11-02 | 2022-11-02 | Prime base line using Base pump
653 | Requalification of cold room 147 at 51NYA missed per FBL-SOP-000994 step 6.5.2 | angela.sauro-havalotti@sanofi.com | 51-147 | 2022-11-01 | 2022-11-01 | RQ to be completed by due date
652 | Batch Record safety label issuance performed by operator that had not trained on FBL-SOP-000379 version 6.0. | jennifer.lopez@sanofi.com | Buffer/Media Batch Records for 45, 51, 74 NYA | 2022-10-18 | 2022-11-01 | Per FBL-SOP-000176 - Step 6.6.1: Complete and document all required training within specified time frame and/or prior to performing tasks in accordance with BQ-GOP-000009.
651 | Step 8.8 of FBL-MBR-000185 incomplete | craig.valery@sanofi.com | 74 NYA PU1-24 | 2022-11-02 | 2022-11-02 | it is expected that all applicable steps within FBL-MBR-000185 are completed/documented
650 | Damage to wall in room 130B | jacob.kallon@sanofi.com | 51 NYA | 2022-11-01 | 2022-11-01 | Area clearance is performed per FBL-SOP-000367
649 | Quality Control Microbiology Analyst not trained on Endotoxin Curricula (QCMF-LABGEN8_Nexus Robot Endotoxin Detection System) | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2022-11-01 | 2022-11-01 | It is expected that the curricula be assigned to the Analyst and the Analyst be up to date on the current effective SOP, prior to assay initiation.
648 | 2 Month agitator PM unable to be performed for X2311 | adam.hildebrand@sanofi.com | 74 NYA | 2022-11-01 | 2022-10-26 | It is expected that 2 month PM's will occur in a timely manner
647 | The RQ on chamber FB-7536 will not be performed in the month due | kathy.tourville@sanofi.com | 80 NYA, stability cage area room 148A | 2022-11-01 | 2022-11-01 | FBL-SOP-000994. step 6.5.2 |  | Deviation
646 | 8NYA-Room 1364-PF8-product impact unknown-FB-3039 GMP Critical Filter Integrity Tester found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8NYA room 1364 | 2022-10-17 | 2022-11-01 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
645 | No reportable results for 76NYA Critical Utility site 144 on 31Oct2022 | Craig.Stevenson@sanofi.com | 76 New York Ave, Room 352 | 2022-10-31 | 2022-11-01 | It is expected that the correct sampling site be collected per FBL-SOP-001050.
644 | 1YR PM DHT OVEN 4 not completed for OCT22 | john.beaule@sanofi.com | 76-352 | 2022-11-01 | 2022-11-01 | Complete 1YR PM for DHT OVEN 4
643 | Post FIT of Eluate vessel Vent Filter not performed | alvaro.rueda@sanofi.com | 74 NYA Falcon | 2022-10-06 | 2022-10-31 | Post FIT of filter would be performed
642 | 1YR PM DHT OVEN 3 not completed for OCT22 | john.beaule@sanofi.com | 76-233 | 2022-11-01 | 2022-11-01 | Complete 1YR PM for DHT OVEN 3
641 | Overdue GMP PM | john.beaule@sanofi.com | 76-3103 | 2022-11-01 | 2022-11-01 | MO PM completed in OCT22
640 | 76NYA-HAFM Biosurgery-product impact unknown-F-188693 Digital Temperature Gauge overdue for calibration | robert.bauch@sanofi.com | 76-Varies | 2022-11-01 | 2022-11-01 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.7.3 Instrument should have been received into Metrology before it's due date
639 | 51NYA-Room 134-51 Purification-product impact unknown-FB-1852 conductivity meter overdue for calibration | robert.bauch@sanofi.com | 51-134 | 2022-11-01 | 2022-11-01 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.7.3 Instrument should have been received into Metrology before it's due date
638 | 68NYA-QCC-product impact unknown-7 GMP Critical Digital thermometers overdue for calibration | robert.bauch@sanofi.com | 68-2700 | 2022-11-01 | 2022-11-01 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.7.3 Instrument should have been received into Metrology before it's due date
637 | Late PM WO 2919718 QT, PM; FBL-SOP-000719 | brizaida.medina-vega@sanofi.com | 76-PURE STEAM SYSTEM | 2022-10-14 | 2022-11-01 | PM completed or extended by P&S Only, expediting a PM extension
636 | expired rodac selected in biolims for EM spread plates | Stefanie.Erickson@sanofi.com | QC Micro bioburden room | 2022-10-20 | 2022-10-30 | for the rodac lot within expiry to be selected | Yes | Deviation
635 | ASM butyl column out of spec on bed height per MES record | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1303 | 2022-10-31 | 2022-10-31 | its expected that packing factor and bed height align to be with in range on both | Yes | No Deviation
634 | Incorrect tote was used for 706210. 8NYA/Fabryzyme | aidan.mula@sanofi.com | 8NYA, RM1350 | 2022-10-31 | 2022-10-31 | TBD at Gemba
633 | 51 RM108C cold room went into alarm for high temp at 2247 on 30OCT22 | 51NYAActionNoticeCoreGroup@Sanofi.com | 51NYA/RM108C | 2022-10-30 | 2022-10-30 | That 51 RM108C would remain between 2-8C.
632 | Distillate Pump on WFI Still X-9251 at 74 failure | Scott.Brouillard@sanofi.com | 74 | 2022-10-29 | 2022-10-29 | NA
631 | Cz MMC Load SUM Agitator Generating Error Code. | stanley.durand@sanofi.com | 8NYA room 1503 | 2022-10-29 | 2022-10-29 | It is expected for agitator to operate as intended and not generate error codes that place recipe in alarm.
630 | 8NYA-ASM-50L bag containing ASM equil media (PN 832317/ CGHMP004 exp 21Dec22) found to be leaking, no back up 50L available | john.lytle@sanofi.com | 8NYA/ cold room 1220 | 2022-10-29 | 2022-10-29 | integral media bag |  | No Deviation
629 | 8NYA-ASM-SUB recipe held at start of pH control due to Hi-Hi pH (7.2). Automation intervention needed to correct air sparge settings | john.lytle@sanofi.com | Rm1333 | 2022-10-29 | 2022-10-29 | at the start of pH control, the pH of the warmed media should be in acceptable range |  | No Deviation
628 | Path forward on determining the cleaning status of UFDF skid after pump seal replacement. | stanley.durand@sanofi.com | 8NYA room 1323 | 2022-10-28 | 2022-10-28 | Per FBL-EWI-000028 during equil. pump should perform as intended without leak, | Yes | Deviation
627 | A leak was found on the agitator seal of a 3k SUM bag at 8NYA | christopher.arnold@sanofi.com | 8RM1350 | 2022-10-28 | 2022-10-28 | That the agitator motion would not create a leak in the bag.
626 | Disk Rupture X3081 during skid WFI rinse | craig.valery@sanofi.com | 74 NYA PU1-24 | 2022-10-28 | 2022-10-28 | no disk rupture
625 | 805015- Record Missing BOM item (p/n:692312) | alejandro.martinez2@sanofi.com | 8NYA Solution Prep, 1350 | 2022-10-28 | 2022-10-28 | BOM item needed to be able to transfer batch
624 | 8NYA-Cz-X2220 DO probes 1 and 2 are starting to diverge, pO2 reading trending upwards | john.lytle@sanofi.com | Rm 1403/ 8NYA | 2022-10-28 | 2022-10-28 | both controlling probes functioning in unison to control SUB sparge
623 | Sterile HA Powder Lot ABYHAS002-A 30M LOD EP OOT | yekaterina.zavgorodniy2@sanofi.com | QC Biopolymers Lab | 2022-09-25 | 2022-10-27 | It is expected that 30M LOD EP result is not OOT
622 | Zinc Wash 2 buffer 705669 | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1323 | 2022-10-27 | 2022-10-27 | Ite expected that buffer ph will be within range during operations |  | Deviation
621 | Incorrect Test Requested Used (CBYSEP054 and CBYSEP055) | alexandre.silva2@sanofi.com | 76 NYA. | 2022-09-30 | 2022-10-27 | It's expected that Films for Test been submitted to QC Lab. with the corrected Lot#.
620 | INC-03606 Samples unable to be received in OneLIMS | paul.consolmagno@sanofi.com | 68 NYA, QCM, Sample receipt | 2022-10-27 | 2022-10-27 | Samples are received and documented contemperaneously | Yes | No Deviation
619 | Zinc skid went into hold due load SUM not being tared during product filtration | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1323 | 2022-10-27 | 2022-10-27 | Its expected that column loads without issues | Yes | Deviation
618 | defective JPN C-section pouch labels p/n 849138 lot 22MN00499 | alan.rousseau@sanofi.com | room 1009 74 NYA Framingham | 2022-10-27 | 2022-10-27 | Expected result is that the label would be usable for printing and packaging of Seprafilm.
617 | Incorrect Hexylene Glycol conversion factor pulled from PI/8NYA/832327/ASM | aidan.mula@sanofi.com | 8NYA, RM1350 | 2022-10-27 | 2022-10-27 | Record needs to be fixed to the correct conversion factor | Yes | Deviation
616 | Hexylene Glycol did not dispense for 832327 buffer/8NYA/ASM | aidan.mula@sanofi.com | 8NYA/RM1350 | 2022-10-27 | 2022-10-27 | Automation will need to walk back recipe to transfer and hex glycol needs to be correctly transferred from the drop |  | Deviation
615 | Contamination Found on SDA Media Prior to Use | jessica.kychler@sanofi.com | QC Microbiology, 68 NYA floor 4 room 4700 | 2022-10-26 | 2022-10-26 | Supplier Deviation
614 | Microbial Isolate tested twice | gretchen.brunner@sanofi.com | QC Micro lab | 2022-09-21 | 2022-10-24 | It is expected that isolates will have only one valid result.
613 | A gemba is needed for: "GEMBA 74NYA -CAPTO MMC- CGHCE021 - Proactive approach if system goes into HI HI ALARM". | richard.tadlockjr@sanofi.com | 74NYA PU1-09 | 2022-10-27 | 2022-10-27 | Process should not exceed the HI HI alarm limit of 30psi
612 | GEMBA 74NYA -CAPTO MMC- CGHCE021 - Proactive approach if system goes into HI HI ALARM | william.bradley@sanofi.com | 74NYA PU1-09 | 2022-10-26 | 2022-10-26 | Process should not exceed the HI HI alarm limit of 30psi
611 | Leak coming from support pole in ceiling above BRX E-Skid. | james.letourneux@sanofi.com | 51NYA Room 140 | Unknown | 2022-10-26 | Controlled area boundaries should be sealed from external environment.
610 | Quarterly Review of New Equipment in INFOR Not Performed | tanni.arif@sanofi.com | sitewide | 2022-09-21 | 2022-09-21 | FBL-SOP-000910 Section 6.7.1: Review all assets with a status of ‘New Equipment’ quarterly, at a minimum.
609 | Power fault on FB-3664 for X3202 scale causing reset, clearing tare. Shortly after, a rupture was noticed on the pump tubing rendering the remaining buffer untransferable. | craig.valery@sanofi.com | 74 NYA PU1-24 | 2022-10-25 | 2022-10-25 | no power fault and no leak
608 | 8NYA CZ SUB X-2250 Aber Probes Drifting | michael.knapp@sanofi.com | 8NYA 1403 | 2022-10-25 | 2022-10-25 | Probes remain consistent with each other and do not deviate.
607 | Errors obtained during FIT testing at 45NYA | nicholas.shoenfelt@sanofi.com | 45-190 | 2022-10-23 | 2022-10-24 | It is expected that the test will complete without errors.
606 | Falcon LAL sample spilled out of vial during assay performance | emily.devaney@sanofi.com | 68NYA 4th floor Micro LAb LAL plate base Bay 4 | 2022-10-24 | 2022-10-24 | Per FBL-SOP-000817 FBL-SOP-001492 | Yes | Deviation
605 | COVID+ Personnel Entry Into 51NYA from 11Oct2022 to 24Oct2022 [Potential 5W2H] | tanni.arif@sanofi.com | 68 NYA: contractor’s desk #1700-C2, cafeteria. 51 NY: – moving equipment from clean suite to storage area. Clean suite on shutdown. | 2022-10-11 | 2022-10-25 | QPP-00031: SARS-COV-2 Risk evaluation of contamination of drug products manufactured by Sanofi |  | No Deviation
604 | 8NYA CZ SUB Post-ATF Guard Filter leak | michael.knapp@sanofi.com | 8NYA 1403 | 2022-10-25 | 2022-10-25 | There is no leak
603 | 76NYA-Room 107-Biosurgery QCBP-GMP Critical Balance F-20435 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-107 | 2022-10-24 | 2022-10-25 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
602 | OOL Bioburden Cleaning Verification 74 NYA Hoses Lot 2022-CV-0645 | craig.valery@sanofi.com | 74 NYA | 2022-10-25 | 2022-10-25 | <5 CFU/100mL is expected
601 | 51NYA-Room 146-FFO3-product impact unknown-Distillate Resistivity on clean steam generator #1, X-9130, F-10964-012 and F-10964-015 found out of tolerance during an unscheduled calibration | robert.bauch@sanofi.com | 51-146 | 2022-10-24 | 2022-10-25 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
600 | Missing September logbook FBS-LBK-000033 | dena.elmesalamy@sanofi.com | 76NYA-3101 | 2022-10-21 | 2022-10-21 | N/A
599 | Endotoxin ASM samples tested concurrently with product samples. | roodline.bertrand@sanofi.com | 68NYA, QC Micro | 2022-10-21 | 2022-10-24 | Per FAD22-003 ASM samples should be tested separately. | Yes | Deviation
598 | 8NYA CZ SUB Recipe Hold requiring Automation Intervention | michael.knapp@sanofi.com | 8NYA 1403 | 2022-10-24 | 2022-10-24 | Recipe Holds are able to be restarted.
597 | Expired Bag Used During Fz Butyl Load Adjust | stanley.durand@sanofi.com | 8NYA room 1323 | 2022-09-01 | 2022-10-24 | Per FBL-EWI-000015 and FBL-EWI-000033 it is expected to use unexpired bags during Fz Butyl Load Adjust.
596 | Aseptic qual failure - 74 NYA | elizabeth.couturier@sanofi.com | 74 NYA CC1-20, F-13851 | 2022-10-17 | 2022-10-24 | See FBL-SOP-000755 Table 3 for alert/action levels
595 | Post-dry Relative humidity not recorded for Loss on Drying | aleen.merzoian@sanofi.com | 76-107 | 2022-10-20 | 2022-10-24 | FBS-MTD-000065 section 6.6.17 %RH must be checked with calibrated hygrometer during post-dry weighout. | Yes | Deviation
594 | Magnehelic gauge readings are not at an action level in the cold room 174 at 45 NYA. | Subhashini.Satishkumar@Sanofi.com | 174 Cold room | 2022-10-22 | 2022-10-22 | It is expected that Magnehelic gauge would function normally with acceptable readings.
593 | 8NYA CLEAN STEAM ALARMED , CANNOT RESTART. FACILTIES WORKING ON IT. | MACJONES.OSEIKUFFOUR@SANOFI.COM | 8NYA MECHANICAL ROOM | 2022-10-21 | 2022-10-22 | FBL-SOP-001326
592 | MFG signed as QA on Action notice 2022-C-067 | dianne.moustafa@sanofi.com | 74NYA Media | 2022-10-16 | 2022-10-21 | QA should sign spot releasing the action notice
591 | Vante Tube Sealer Recipe Issue | erica.davis@sanofi.com | 8 NYA | 2022-10-19 | 2022-10-21 | For C-Flex 3/4x1 and 1/4x7/16 tubing, the sealing times should be 180 sec and 100 sec respectively. It was found to be 140 sec and 90 sec on certain sealers.
590 | MBR corrections 74 NYA 2 Issues: unverifiable batch id and verified issuance not performed | craig.valery@sanofi.com | 74 NYA | 2022-10-20 | 2022-10-20 | accurate information to be recorded
589 | No As-Found data collected for Environmental Chamber F-7536 Prior to work by Shon's | todd.greenough@sanofi.com | 80 NYA | 2022-09-16 | 2022-10-19 | Stability would have scheduled metrology to perform As-Found data collection prior to the vendor performing work on the unit.
588 | The inoculum from the maintenance wave that was supposed to be used to inoc the virus wave was mistakenly drained. Normally it would be drained, if it wasn't going to be used for the virus wave. | adam.rowe@sanofi.com | 74NYA, CC2-01 | 2022-10-20 | 2022-10-20 | The residual material in the maintenance wave would have been used to inco the virus wave
587 | Skid setup step missing from SKID cleaning EUR | michael.daly@sanofi.com | 8 NYA room 1303 - column packing room | 2022-10-20 | 2022-10-20 | It is expected the record directs operators to setup of the skid for CIP prior to performing the CIP |  | Deviation
586 | No as found data performed on chamber prior to work being performed | kathy.tourville@sanofi.com | 80 NYA, stability cage area room 148A | 2022-10-20 | 2022-10-20 | BQ-GOP-000046, step 5.7.2.1
585 | 8NYA-Room 1170-Dept FFO3-no product impact-FB-4002-005 found out of tolerance during a scheduled calibration for conductivity | robert.bauch@sanofi.com | 8-1170 | 2022-10-17 | 2022-10-20 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
584 | 8NYA-ASM- The setpoint for the Wave Rocker speed was found to be incorrect | john.lytle@sanofi.com | Rm1333 | 2022-10-20 | 2022-10-20 | Rocker setpoint should be 25rpm. This is controlled by the recipe automation |  | Deviation
583 | No reportable result for endotoxin sample 512196 Buffer_705675 CGHBN015 | marissa.lentoni@sanofi.com | 68 NYA/4700/QC Micro lab | 2022-10-19 | 2022-10-19 | Sample acceptance criteria would pass at the minimum valid dilution of 1:40.
582 | 74NYA-Room CC1-21-Dept CC74-Product impact unknown-NO EMS alarms for Tag # F074-TI-406009 on X-4069 incubator | robert.bauch@sanofi.com | 74-CC1-21 | 2022-10-04 | 2022-10-19 | there should have been an EMS alarm when the temperature went low by 6 degrees C
581 | Late entry for parts washer logbook | william.lee@sanofi.com | 74NYA | 2022-10-19 | 2022-10-19 | For COP wash run operations the verifier should sign on the same day as the performer. | Yes | No Deviation
580 | 8NYA CZ MCC Blank Run Eluate SUM Acquisition Error | michael.knapp@sanofi.com | 8NYA 1403 | 2022-10-18 | 2022-10-18 | Recipe is able to acquire the equipment it needs. Equipment is not still attached to prior recipes.
579 | Missing Witness by. | alexandre.silva2@sanofi.com | 76 NYA (Purification). | 2022-05-16 | 2022-10-19 | It's expected that Witness by Steps been performed at time of execution.
578 | Leak observed near PIT02 on X3211 | craig.valery@sanofi.com | 74 NYA PU1-24 | 2022-10-19 | 2022-10-19 | It is expected that no leaks are observed
577 | Supplier Deviation needed for: 3D TOTE BAG W/ PUMP-AROUND LOOP, 2000L (P/N: 692178) | alejandro.martinez2@sanofi.com | 8NYA CNC | 2022-10-19 | 2022-10-19 | B114121-I
576 | CZ SUB X2250 Rate of Change Alarm | michael.knapp@sanofi.com | 8NYA 1403 | 2022-10-19 | 2022-10-19 | Media switchover works correctly and Rate of Change is calculated off of current active tote.
575 | 68NYA-Metrology-no product impact-Metrology Standard S-1405 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2022-10-05 | 2022-10-19 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
574 | QCA Alarms at 08NYA not reviewed in time. | timothy.ryley@sanofi.com | 08 NYA | 2022-10-18 | 2022-10-18 | Alarms are commented on within 72 Hours
573 | Test version of MES record utilized for GMP Operation | kyle.tesorero@sanofi.com | 8NYA | 2022-10-12 | 2022-10-18 | PROCOLPK v10 should reference 8EURSKD, not 8EURSKT.
572 | FBL-LBK-000338 used after new revision date of 30SEP22 | dianne.moustafa@sanofi.com | 74 NYA Media CC1-08 | 2022-10-01 | 2022-10-17 | It is expected that new logbook should have been used.
571 | GEMBA: 74NYA - Automation movement on load from V-2561 | william.bradley@sanofi.com | 74NYA CC1-07 | 2022-10-17 | 2022-10-17 | Delta V recipe to run based on the procedure
570 | GEMBA: 74 NYA - DSP- V-2561 Issues - Cannot load to HyperCel | william.bradley@sanofi.com | 74NYA CC1-07 | 2022-10-17 | 2022-10-17 | SUT to be closed once Capto Eluate has completed load.
569 | 8NYA CZ SUB pH Control Setpoint Error | michael.knapp@sanofi.com | 8NYA 1403 | 2022-10-16 | 2022-10-16 | Process Air is able to be sparged into the SUB without incident.
568 | No Product – No Lots – 45 NYA – WFI Loop Gasket Leak | sheila.pandolfo@sanofi.com | Leak was observed at the outside wall of Room 189B at 45 NYA by the access door. The leak was at a point of use cooler located a few feet from Room 189B. The WFI entered Room 189B (MSAT Pilot Lab, non GMP). | 2022-10-16 | 2022-10-16 | It is expected that there are no leaks on the WFI loop distribution lines. | No | Deviation
567 | Automation intervention needed 2 times during Fz SP Production | robert.morawski@sanofi.com | 8NYA RM1323 | 2022-10-15 | 2022-10-15 | On the first instance, the agitator turning on at the start of the pH adjust recipe. On the second one, the bubble trap sensor to read correctly. |  | No Deviation
566 | pH adjustment of Fz SP load went below target range | robert.morawski@sanofi.com | 8NYA RM1323 | 2022-10-15 | 2022-10-15 | pH probe in SUM read the correctly, and the pH to be within range. |  | Deviation
565 | 8NYA-Cz, ASM-Covid positive employee in cGMP areas prior to testing positive | john.lytle@sanofi.com | 8NYA, all Upstream and CNC space leading to Upstream suites | 2022-10-14 | 2022-10-14 | Covid positive employees remain out of classified areas |  | Deviation
564 | X-3071 NWP Failure due to Incorrect Baseline value | jarad.kukla@sanofi.com | 8NYA / 1323 | 2022-10-14 | 2022-10-14 | It is expected that NWP passes and NWP baseline saved in DeltaV and used for recovery calculation.
563 | Insect found in 74 NYA main material airlock hallway | marissa.lentoni@sanofi.com | 74 NYA | 2022-10-14 | 2022-10-14 | An insect would not be found in this area.
562 | Libero Data Logger Temperature Excursion | Jonathan.Lucas@sanofi.com | Patheon | 2022-09-28 | 2022-10-14 | FBL-SOP-000638, step 6.5.2.
561 | Sampling for CGHMP002 | richard.manasian@sanofi.com | 1350 8NYA | 2022-10-04 | 2022-10-14 | Samples should be taken per FBL-EWI-000009
560 | Thyrogen- Shutdown Daily Cleanings and Inspections not Perfomed per FBL-SOP-0001060 | christopher.holmes@sanofi.com | 51 NYA | 2022-10-11 | 2022-10-14 | It is expected that the facility would be inspected and remediated daily per step 6.2.3 of FBL-SOP-001060.
559 | Cancelled CO2 and Alcohol deliveries impact assessment and path forward | Subhashini.Satishkumar@Sanofi.com | 51, 74 and 8 NYA | 2022-10-13 | 2022-10-13 | Bulk deliveries supposed to happen as per the schedule.
558 | Fz Nano guard filter changeout done incorrectly | robert.morawski@sanofi.com | 8NYA RM1503 | 2022-10-14 | 2022-10-14 | Stainless steel clamps used for guard filter changeout.
557 | 8NYA CZ Harvest SUM level low | michael.knapp@sanofi.com | 8NYA 1403 | 2022-10-12 | 2022-10-14 | Harvest maintains required weight |  | Deviation
556 | Missing documentation of Mobile Phase A Lot/Expiry in OneLIMS/Empower | aimee.park@sanofi.com | 68NYA, QC Chemistry 2nd floor | 2022-08-21 | 2022-10-13 | Should have recorded reagent Lot/expiry at time of run | Yes | No Deviation
555 | 805015 Cystine/Tyrosine BOMs are not uniform | aidan.mula@sanofi.com | 8NYA, RM1350 | 2022-10-13 | 2022-10-13 | All of the BOM's(MES, Athena, and ASM) should be uniform. |  | No Deviation
554 | 100L SUB CGHS1022 Negative Pressure Reading | mark.lindblad@sanofi.com | 74NYA, CC2-01/03 | 2022-10-09 | 2022-10-13 | Potential to deviate from FBL-CST-000030/ FBL-ARA-000013 for Last Day Bioburden and or evidence of contamination.
553 | LIMS label not available for | michael.daly@sanofi.com | 8 NYA Room 1323 downstream ballroom | 2022-10-13 | 2022-10-13 | It is expected that all RPT samples have a corresponding LIMS label available |  | No Deviation
552 | AUG22 PMs went overdue 01SEP22 | john.mendenhall@sanofi.com | 51NYA, rooms 142, 88, 122B | 2022-09-01 | 2022-10-13 | FBL-SOP-000081. Step 6.4.9
551 | X2230 media transfer initiated and weight tared by recipe | duane.messier@sanofi.com | 74NYA / CC-201 | 2022-10-13 | 2022-10-13 | Per MBR final weight should be 415-425Kg.
550 | The incubation time for QCM Sample CBYHAS008 on attachment 4 of FBL-SOP-000607 and attachment 1 of FBS-SOP-000941was not documented | cleiton.lopes@sanofi.com | 68 NYA, QCM, Sample receipt | 2022-10-11 | 2022-10-12 | Attachments 1 of FBS-SOP000941 to have incubation date/time recorded and attachment 4 of FBS-SOP-000607. | Yes | Deviation
549 | out of service tote scanned and used for buffer | aidan.mula@sanofi.com | 8NYA | 2022-10-02 | 2022-10-07 | MES record should not allow OOS equipment to be utilized without generating a warning, exception, or otherwise stopping the operator from proceeding. |  | No Deviation
548 | 45NYA-Cold Room 174-APT CC-Product impact unknown-FB-0358 all 8 Hepa filters fail calibration due to the 7 of the 8 HEPA Fan units are not functional | robert.bauch@sanofi.com | 47-174 | 2022-10-11 | 2022-10-13 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
547 | Hair Found on 51 NYA Grade C RODAC B | jessica.kychler@sanofi.com | 68 NYA Room 4700 QC Micro | 2022-10-04 | 2022-10-12 | It is expected that PPE would prevent hair from being present in a grade C area
546 | 51NYA-Room 132-APT DS-product impact unknown-F-15531 UV easy cal found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 51-132 | 2022-09-20 | 2022-10-12 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
545 | 76NYA-Room 3104 Biosurgery-product impact unknown- F-21186-011 VHP Generator DPT3 enclosure pressure found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 76-3104 | 2022-10-11 | 2022-10-12 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
544 | Missing monthly (month of Aug) cleaning for Stainless Stairs and Decking | dena.elmesalamy@sanofi.com | 76NYA 3101 | 2022-08-31 | 2022-09-16 | Performing monthly cleaning on time
543 | 8NYA-Room 1503 PF8-product impact unknown-Temperature indicator FB-4310-005 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1503 | 2022-10-10 | 2022-10-12 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
542 | CU Bioburden counted on day 4 outside the range 5-7 days | Stefanie.Erickson@sanofi.com | QC Micro counting bench | 2022-10-06 | 2022-10-11 | for the plates for CU WFI bioburden to be counted within the incubation range of 5-7 days | No | Deviation
541 | New clamps being received on 1000L/3000L SUM bags | richard.manasian@sanofi.com | 1350 8NYA | 2022-07-28 | 2022-07-28 | Clamps should be secure when connected
540 | Athena expiry of TFF Diafiltered Retentate is not aligned with FBL-CST-000030 | craig.valery@sanofi.com | 74 NYA PU1-24 | 2022-10-11 | 2022-10-11 | Athena Expiration should align with FBL-CST-000030
539 | 8NYA-Cz- needed automation intervention at start of Harvest, Harvest Bleed pump not responding | john.lytle@sanofi.com | Rm 1434 | 2022-10-11 | 2022-10-11 | When the Harvest prompt is acknowledged, the automation acquires the Harvest SUM and associated pump
538 | Four (4) PMs not completed by due date | angela.sauro-havalotti@sanofi.com | 76-210 and 51-138, 51-153 | 2022-10-01 | 2022-10-06 | PMs are completed by due date or extended per FBL-SOP-000081
537 | Supplier deviation for 692169 3K SUM bag | richard.manasian@sanofi.com | 1350 8NYA | 2022-10-10 | 2022-10-10 | The tote bag should have four holes to connect to spider hoist | Yes | Deviation
536 | 8NYA FZ MCC Recipe End of Run-143 of target 144 eluates collected | michael.knapp@sanofi.com | 8NYA 1333 | 2022-10-11 | 2022-10-11 | The MCC is complete following the collection of the 9th eluate SUM. |  | No Deviation
535 | Cerezyme – CGHBN006– 8 NYA – Media Rate of Change (RoC) Parameter (EM-X2250-MED/MEDIA_ROC) went to zero for 7 hours 41 minutes, Alarm information not documented in iNova log | sheila.pandolfo@sanofi.com | 8 NYA Room 1403 | 2022-08-24 | 2022-10-10 | It is expected that the media rate of change (perfusion rate) would be within 1-2.5 RVD (Reactor Volumes Per Day) per FBL-CST-000029 version 14.0 (Control Strategy for Cerezyme Integrated Continuous Biomanufacturing (Cz-ICB) Process), Table 7.8, An excursions </= 6 hours is allowed. Per FBL-EWI-000004 version 8.0 (Alarm Response Procedure) it is expected that all quality critical alarms have the alarm cause and immediate actions documented in the iNova log. Once this is complete, the operator can acknowledge the DeltaV alarm and continue the recipe. | Yes | Deviation
534 | During SP blank run recipe failed to communicate with eluate SUM. | stanley.durand@sanofi.com | 8NYA room 1323 | 2022-10-10 | 2022-10-10 | It's expected for the recipe to communicate with the eluate SUM
533 | QA Ops Quarterly Walk-Through Missed at 8 NYA | Mackenzie.Gomes@sanofi.com | 8 NYA | 2022-10-10 | 2022-10-10 | QA Ops performs walk-throughs quarterly (MAR22, JUN22, SEP22, DEC22). JUN22 and SEP22 walk-throughs were missed. |  | Deviation
532 | Fz SP x3031 set up recipe in hold due to max stabilization timer exceeded. Will not resume. | robert.morawski@sanofi.com | 8NYA RM1503 | 2022-10-10 | 2022-10-10 | Setup recipe continued without alarm/hold.
531 | Missed to perform an initial 30-minute soak submerging for Collection Can during Aug (lot CBYHAM002) and Sweco System/Collection Can during Sep22 (lot CBYHAH007) | dena.elmesalamy@sanofi.com | 76NYA 3103 | 2022-08-30 | 2022-10-07 | perform an initial 30-minute soak submerging for all dirty equipment
530 | Ammonium Sulfate barrel has a leak on bottom outlet | robert.morawski@sanofi.com | 8NYA CNC | 2022-10-10 | 2022-10-10 | No liquid being observed at the bottom of the barrel.
529 | 8NYA Cryostat alarm not acknowledged within QCA time frame | michael.knapp@sanofi.com | 8NYA 1270 | 2022-10-05 | 2022-10-05 | Alarms are clearly identified and acknowledged within required timeframes.
528 | Counts on negative control plate 76NYA Dynamic EM | roodline.bertrand@sanofi.com | 76NYA, F-0704 | 2022-09-27 | 2022-10-05 | Per FBL-SOP-001496 and FBL-SOP-000755, negative control plates should have 0CFU.
527 | Active Asset # TCU X-2223001 (asset# 2009571) is Deactivated in INFOR | tanni.arif@sanofi.com | 8NYA | 2022-10-07 | 2022-10-07 | Active assets should not be deactivated in INFOR
526 | GEMBA: 74NYA - TFF - Buffer (790818) out of temperature spec | william.bradley@sanofi.com | 74NYA PU1-22 | 2022-10-09 | 2022-10-09 | Temperature in the record states the buffer to be between 3 and 15 degrees.
525 | Media hold on SUB 2220. Sub could not pull media due to unconnected Media tote. ewf | Allan.Riley@sanofi.com | 8NYA, Phase 2 - 1403 | 2022-10-09 | 2022-10-09 | FBL-CST-00009 - media tote is hooked up before needed
524 | 74 NY Ave Leak during FS1 batching FBL-MBR-000131 | greg.dankner@sanofi.com | 74 NY Ave room CC1-08 | 2022-10-09 | 2022-10-09 | Expected that the FS1 would not leak from the 100L Mix bag and SUM (Single Use Mixer)
523 | 45NYA GMP WFI LOOP GASKET LEAK IN THE LAB. | MACJONES.OSEIKUFFOUR@SANOFI.COM | 45NYA LOOP IN ROOM 168 | 2022-10-08 | 2022-10-08 | FBL-SOP-000767 FBL-SOP-000646
522 | 76NYA R&D DISTRIBUTION LOOP SHUTDOWN DUE TO LOW PRESSURE. WO# 2926974 | MACJONES.OSEIKUFFOUR@SANOFI.COM | BUILDING 76NYA ROOM 347 MECHANICAL ROOM | 2022-10-07 | 2022-10-08 | FBL-SOP-000365 / FBL-SOP-000870
521 | Power outage from external supply source affecting 45NYA and 5 MTN | kelly.alanmurphy@sanofi.com | 45NYA (entire building) | 2022-10-07 | 2022-10-07 | Power stays on consistently.
520 | FALCON: Temp and level alarms for cryo X9622 | nicole.lufkin@sanofi.com | 74 NYA CC1-18 | 2022-10-04 | 2022-10-04 | cryo level maintained
519 | 74NYA-Room UT1B-02-Falcon-CIP Skid X-9521 Return Conductivity Indicator (F-13284-001) and Delta V (F-13284-005) found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 74-UT1B-02 | 2022-10-06 | 2022-10-07 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
518 | X2230 Held for HiHi DO CTLR | patrick.hickey@sanofi.com | 74 NYA. CC2-01 | 2022-10-07 | 2022-10-07 | The DO is expected to stay within the LoLo and HiHi of 67-77%.
517 | GEMBA 74NYA - HyperCel Equilibration CGHHE019 | william.bradley@sanofi.com | PU1-08 | 2022-10-06 | 2022-10-06 | Sample is to be taken when flow totalizer reaches 615 L per FLB-MBR-000183 collect 50mL sample of column effluent from skid process waste drain
516 | 8NYA- ASM- Post Thaw sample check had lower than expected ViCell results | john.lytle@sanofi.com | Rm 1343 | 2022-10-06 | 2022-10-06 | greater than 0.3 VCD
515 | HA - CBYHAS008 - 76NYA Biosurgery - COVID positive employee on manufacturing floor | charlie.houle@sanofi.com | 76 NYA rooms | 2022-10-04 | 2022-10-04 | Employee should not enter facilities COVID positive
514 | Incorrect CV Spec used by QC on UFDF CV results | stanley.durand@sanofi.com | 8NYA room 1323 | 2022-10-05 | 2022-10-06 | It is expected that UF equipment cleaning spec is used per FBL-FRM-000776 | Yes | Deviation
513 | Slight damage to two G-Totes (5485-01)for Sepramesh 6642-01 lot #'sCBYMES030 and CBYMES033 | alan.rousseau@sanofi.com | 2-8C at 80 NYA Room 7CRT refrigerator | 2022-09-27 | 2022-09-27 | FBS-MBR-000088 FBS-SOP-000694
512 | Air in pouches US lot CBYSEP045 p/n 789446 making it difficult to close 10-pk cartons | alan.rousseau@sanofi.com | 74 NYA room 1017 | 2022-10-05 | 2022-10-05 | Pouches were expected to fit easily into the 10-pack cartons
511 | A280 Capture Eluate samples prepared and tested in duplicate | nancy.watkins@sanofi.com | QC Chemistry, 45NYA | 2022-10-04 | 2022-10-05 | per FBL-MTD-000050, step 6.5.5. Any other sample fractions will be run using the suggested dilution scheme without duplicates.
510 | Lab event discovery outside of testing window for ABYHAM005 18M Stability | kristina.erickson@sanofi.com | 76NYA RM 107 | 2022-06-14 | 2022-09-29 | Correlation coefficient >0.90 in FBS-MTD-000053 | Yes | Deviation
509 | CNC space was entered with improper gowning | kristine.caron@sanofil.com | 8NYA | 2022-10-05 | 2022-10-05 | Employees would have gowned per FBL-EWI-000002
508 | Antifoam Setup SUTs labels missing to perform take-out against the record | Lolav.ahmed@sanofi.com | 8NYA | 2022-10-05 | 2022-10-05 | N/A
507 | Cz CM blank run/ Confusion on what samples should be submitted to QCM | stanley.durand@sanofi.com | 8NYA room 1503 | 2022-10-03 | 2022-10-03 | TBD
506 | 8NYA-room 1323-product impact unknown-FB-2949-007 the pre filter flow on X-309001 Nano filtration cart found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1323 | 2022-10-04 | 2022-10-05 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
505 | Covid-74NYA-Room CC1-21-product impact unknown-FB-3654-005 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-CC1-26 | 2022-10-04 | 2022-10-05 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
504 | Capto column operation went into hold for high pressure during sanitiaztion phase | Peter.darcy@sanofi.com | 74 NYA PU1-09 | 2022-10-05 | 2022-10-05 | FBL-MBR-000182 step 8.11.4. Pressure should not exceed 30 psig | Yes | Deviation
503 | 66 CFU Bioburden result for 2022-CV-0585 | alvaro.rueda@sanofi.com | 74 NYA Falcon | 2022-10-22 | 2022-10-03 | < 10 CFU/100mL bioburden results
502 | Incorrect label was used for WIP label on BGF1S07 H6 | brian.mcgee@sanofi.com | 51NYA/Rm 140 | 2021-06-28 | 2022-10-04 | It is expected that the correct label is used at step 4.33.4 for WIP labeling
501 | Wave X2125 went into hold for loss of communication to OPC server | lindsey.hymer@sanofi.com | 74NYA/CC2-01 | 2022-10-03 | 2022-10-03 | Wave rocker should be properly communicating with Unicorn |  | No Deviation
500 | Balance Calibrated Incorrectly Resulting in Invalid Pass of Daily Calibration | kristina.erickson@sanofi.com | 76-107 | 2022-09-28 | 2022-10-03 | FBS-SOP-000461 step 6.3.7, acceptance weight criteria for 10mg-52g balance.
499 | SARS CoV2 - Lot TBD -74NYA - 2022-CV-0592 (74NYA Buffer Transfer Line TP-3111 to PU1-24 )TOC Action Level | Craig.Stevenson@sanofi.com | 74NYA, Falcon | 2022-09-30 | 2022-10-03 | Is it expected that TOC generate passing results below alert limit.
498 | 74NYA-Room CC1-07-Falcon-No as found data taken on floor scale F-14200 - unit is non functional | robert.bauch@sanofi.com | 74-CC1-07 | 2022-09-30 | 2022-10-03 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been able to be collected
497 | 8NYA-RM 1334-product impact unknown- 2 pipettes FB-3492 and FB-3493 overdue for calibration | robert.bauch@sanofi.com | 8-1334 | 2022-10-01 | 2022-10-03 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.7.3 Instrument should have been received into Metrology before it's due date
496 | 51 NYA- Thyrogen- Missing signatures during bioreactor inoculation | ian.odell@sanofi.com | 51 NYA Room 140 | 2022-05-24 | 2022-10-02 | All bioreactor inoculation steps will be performed / verified in MBR at the time of execution.
495 | 51 NYA- Thyrogen- Missing signatures during post bioreactor SIP activities. | ian.odell@sanofi.com | 51 NYA Room 140 | 2022-05-23 | 2022-10-02 | All bioreactor preparation steps will be performed prior to inoculation.
494 | 51 NYA- Thyrogen- Missing witness initials during bioreactor refeed step | ian.odell@sanofi.com | 51 NYA Room 140 | 2022-07-07 | 2022-10-02 | Expected that witness section would be complete with initials and date.
493 | 76NYA-Room 3104 Biosurgery-product impact unknown- F-21186 VHP Generator is overdue for calibration (9 items) | robert.bauch@sanofi.com | 76NYA Room 3104 | 2022-10-01 | 2022-10-03 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 units should have been calibrated before their due date.
492 | No reportable bioburden and LAL results for site 51-DI-1 for month of September as site was unable to be sampled | marissa.lentoni@sanofi.com | 51 NYA, Room 129A, Parts Washer Mechanical Space | 2022-09-30 | 2022-10-30 | 51-DI-1 would function as expected and samples would be collected and tested for this site.
491 | Failed seal on NRSR sample line tee, ATF A/ Cz | john.lytle@sanofi.com | Rm 1434 | 2022-10-02 | 2022-10-02 | integral seal
490 | Cz SUBx2220 backup vent filter set (PN 762421) not installed | john.lytle@sanofi.com | Rm 1434 | 2022-09-25 | 2022-10-02 | FBL-EWI-000011, Section "Install Vent Filter", Step #5 instructs the operator to "connect both sets of vent filters to SUB if not already attached and clamp off all filters."
489 | Cz SUB x2220 experiencing agitator speed control issues leading to process holds | john.lytle@sanofi.com | Rm 1434 | 2022-10-02 | 2022-10-02 | agitator speed adjusting to SP 80rpm
488 | Missing document during quarterly batch record audit | william.lee@sanofi.com | 74NYA | 2022-09-30 | 2022-09-30 | The batch record should be available and documented in the quarterly audit.
487 | Cz Leak identified during tube sealing | valerie.staras@sanofi.com | 8NYA Room 1403 | 2022-09-30 | 2022-09-30 | it is expected that the seals remain integral
486 | Missing to fill out FBS-LBK-000029 attachment 3 (Process Hose Usage Log for HA Ferm area for CBYFRM029 and CBYFRM030) | dena.elmesalamy@sanofi.com | 76NYA 364 | 2022-08-01 | 2022-09-30 | Process Hose Usage Log for CBYFRM029 and CBYFRM030
485 | Clean Utilities Team member performing Preventive Maintenance and Work Orders without completed the SOPs. | brizaida.medina-vega@sanofi.com | All site | 2022-09-29 | 2022-09-29 | Complete the SOPs in iLearn before start performing any job associated with his responsibilities.
484 | CBYSEP068 - Hal sheet only lot. not enough full sheets to perform "In process film Weight" | fatima.fragaalvarez@sanofi.com | 233 | 2022-09-30 | 2022-09-30 | Weight 160 half sheets (where 2 half sheet is equivalent to 1 full sheet)
483 | (4) DP transmitter sensors cannot have the final ‘As Found’ calibration performed | elizabeth.couturier@sanofi.com | QC Molecular Biology lab, Waltham | 2022-10-01 | 2022-10-01 | It is expected that as-found data is able to be collected for instruments
482 | Non reportable for Falcon MMC POU Buffer 827753 (497690) and MMC POU 790804 (49770) | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2022-09-09 | 2022-09-28 | It is expected that samples are delivered to QC Micro with ample testing time and not received if samples are expired | Yes | Deviation
481 | Wrong lot of Illustration labels 21MN00704 p/n 803141 were used in packaging of Seprafilm lots CBYSEP041 and CBYSEP043 lots | alan.rousseau@sanofi.com | 1017 74 NYA | 2022-09-23 | 2022-09-29 | The expected Results were that QA and MFG would issue and verify the correct lot of illustration labels to be used.
480 | Hyaluronic Acid - 9BYHAS006 - 76NYA - Endotoxin was not submitted to be tested within the 36M Testing Window | Craig.Stevenson@sanofi.com | 76 New York Ave, Room 107, Quality Control (QC) Biopolymers (QCBP) | 2022-09-17 | 2022-09-29 | Per FBL-PTM-000044, Endotoxin should be submitted to be tested at the 36M Stability Time Point.
479 | Incorrect logbook version in use in CC manufacturing | adam.hildebrand@sanofi.com | 74 NYA/ CC2-01 | 2022-09-29 | 2022-09-29 | It is expected that the correct version of logbooks be in use at all times on the MFG floor
478 | Missing Vesi Retain Sample H9 CGF1B03 (rhTSH) | daniel.bertman@sanofi.com | 51 NYA/108B | 2022-05-17 | 2022-09-28 | The samples are to remain in freezer until required disposal. The test request should remain with the samples.
477 | A280 of Thyrogen Incorrect Cuvette selected/reported and samples retested without notifying management | john.kapaniris@sanofi.com | 45 New York Avenue Quality Control Chemistry Lab | 2022-09-27 | 2022-09-27 | Analyst was required to inform QC Management of the error prior to re-running samples
476 | Gemba for TOC CV Rinse Incorrect Acid Rate and Oxid Rate Used | nicholas.koester@sanofi.com | QC Micro | 2022-09-20 | 2022-09-28 | Per FBL-SOP-000595 Step 6.4.8, SARS-COV2 samples may only be tested in Standard Mode using the below parameters: Acid Rate of 0.3, Oxid Rate of 0.7
475 | High high alarm for DO was incurred on 15SEP22 @ 1952 on X2301. | adam.hildebrand@sanofi.com | 74 NYA | 2022-09-15 | 2022-09-28 | It is expected that the bioreactor controls limits as outlined in FBL-CST-000030
474 | A280 stability sample re-tested before LE was written and reviewed | nancy.watkins@sanofi.com | QC Chemistry, 45NYA | 2022-09-27 | 2022-09-27 | BQ-GOP-000246 step 6.27 : Upon QC Management approval of the LE assessment, the sample(s) are eligible for fresh analysis testing and samples allocated for testing.
473 | DeltaV cannot communicate with Unicorn. Cannot load recipes on wave reactors | adam.rowe@sanofi.com | 74NYA, CC2-01 | 2022-09-27 | 2022-09-27 | Should be able to load a recipe
472 | During MFC checks within CGHPB034 step 9.5.8 was out of range with a reading of 0.20 (range is 0.23 - 0.27) | paul.mclaughlinjr@sanofi.com | 74NYA | 2022-08-12 | 2022-09-08 | 0.23 - 0.27
471 | 500L SUB in hold during inoc of the 2kL reactors for HIHI DO. | paul.mclaughlinjr@sanofi.com | 74NYA | 2022-09-26 | 2022-09-26 | O2 would be stable during transfer of product
470 | High temp on indirect water condenser. 74NYA. WO#2922234 | MACJONES.OSEIKUFFOUR@SANOFI.COM | 74NYA | 2022-09-26 | 2022-09-26 | N/A REFERENCE WO#2922234
469 | 100L SUB gas and DO shut down due to HIHI pressure | paul.mclaughlinjr@sanofi.com | 74NYA | 2022-09-21 | 2022-09-26 | DO would be stable at 72%
468 | COVID+ Personnel Entry Into 68NYA/51NYA from 09Sep2022 to 22Sep2022 | tanni.arif@sanofi.com | 68NYA / 51NYA | 2022-09-23 | 2022-09-24 | ?
467 | Missing information from credo cooler log 2022QCC0063 | michael.montiverdi@sanofi.com | unknown | 2022-08-12 | 2022-09-23 | End time to be entered.
466 | pCO2 spike on the 100L SUB | paul.mclaughlinjr@sanofi.com | 74 NYA | 2022-09-26 | N/A | pCO2 would be stable
465 | Cz Q Membrane filter leaked while performing priming | robert.morawski@sanofi.com | 8NYA RM1503 | 2022-09-26 | 2022-09-26 | The filter primed with the buffer, and continued on with the recipe. | Yes | No Deviation
464 | 45NYA-QCC45 Room 1130-product impact unknown- F-21139 and F-21140 Chloride meters overdue for calibration | robert.bauch@sanofi.com | 45-1130 | 2022-04-01 | 2022-09-26 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.7.3 Instrument should have been received into Metrology before it's due date | Yes | No Deviation
463 | X2230 Hold due to pCO2 High High Alarm followed by hold due to DO high high alarm | sean.gormley@sanofi.com | 74NYA/CC2-01/X2230 | 2022-09-25 | 2022-09-26 | Probes will stay within limits set in DeltaV
462 | Two instances of Goods issuance not performed for Thyrogen SP1 Column | Ryan.Geraghty@sanofi.com | 51NYA, Room 138 | 2022-09-11 | 2022-09-26 | Per FBL-SOP-000897, Section 3, Manufacturing is responsible for "Athena component and material goods issue transactions shall be completed prior to use on the day of the operation." |  | No Deviation
461 | Gemba: No reportable endotoxin results for 51-WFI-2, 74-RO-1, 56 Pure, 74HWFI-25 sampled on 10May22 and 51-DI-4 sampled on 16Jun22 | marissa.lentoni@sanofi.com | 45 NYA, 51 NYA, 74 NYA | 2022-05-10 | 2022-09-25 | Samples would be tested and generate valid results. | No | Deviation
460 | Prompt to perform transfer out of 10L wave reactor X2141was acknowledged before 20L wave was equilibrated. | Brandon.Steff@sanofi.com | 8RM1403 | 2022-09-24 | 2022-09-24 | Selecting "NO" on "Confirm ready to transfer" prompt will return wave recipe to growth state.
459 | Leak during media transfer into wave reactor X2151 due to clamp on transfer line | Brandon.Steff@sanofi.com | 8RM1403 | 2022-09-24 | 2022-09-24 | Transfer line is free from obstructions prior to transfer.
458 | GEMBA: 74NYA - CAPTO MMC sample leak (COV_MMC_827588_EQUIL_BIOB) | william.bradley@sanofi.com | 68NYA in QC. The sample came from 74NYA PU1-09 | 2022-09-24 | 2022-09-24 | Samples are to be completely sealed.
457 | High High pressure "30PSI" on Capto MMC Column | alvaro.rueda@sanofi.com | 74 NYA Falcon | 2022-09-23 | 2022-09-23 | HI HI pressure is not exceeded
456 | Weight controller high high alarm on bioreactor X2210 | Brandon.Steff@sanofi.com | 8RM1333 | 2022-09-23 | 2022-09-23 | Harvest and media pumps maintain SUB weight of 100kg
455 | incorrect date of manufacture, resulting in incorrect expiry for VP5 in Athena | nicole.lufkin@sanofi.com | 74 CC1-07 | 2022-09-21 | 2022-09-22 | correct manufacture date and expiry
454 | 51NYA-APT CC room 152A-product impact unlikely-FB-1584-002 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 51-152A | 2022-09-22 | 2022-09-23 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
453 | Nano equil cut short to preserve chase buffer | zachary.rickard@sanofi.com | 8NYA 1503 | 2022-09-22 | 2022-09-22 | buffer should not have run short. |  | Deviation
452 | Valve not working during Nano Equil | zachary.rickard@sanofi.com | 8NYA 1503 | 2022-09-22 | 2022-09-22 | Valve should be closing.
451 | Witness step for COVID batch record CGHUB015/FBL-MBR-000185 signed by the same operator that performed the step | megan.botti-marino@sanofi.com | 74 NYA | 2022-09-02 | 2022-09-22 | According to "BQ-SOP-000269 (Good Documentation Practices), step 5.13.3, "If the document requires a witness signature, that signature must be made at the execution of the process. "Approved By," "Checked By", "Reviewed By", "Verified by" and Witnessed by" signatures must not be made by the same person who performed the activity. Additionally, according to chart 2 of "BQ-GOP-000248 (Phenix Deviation Process)," a deviation is required for a missing witness signature.
450 | Air filter line burst due to asepticquick connection not removed | alejandro.martinez2@sanofi.com | 8RM1350 | 2022-09-22 | 2022-09-22 | Determine if new SUT needed to proceed
449 | Aseptic qual failure - 51 NYA | elizabeth.couturier@sanofi.com | 51 NYA room 140 | 2022-09-13 | 2022-09-21 | It is expected that aseptic qualifications are passing
448 | MCC X-2560 Cleaning PQ TOC rinse failures at 3 out of the 4 rinse locations. | valerie.staras@sanofi.com | 8 NYA | 2022-08-20 | 2022-09-03 | It is expected that TOC be within acceptance criteria (.50ppm)
447 | Missing Logbooks at 51 and 45 NYA during LB auditing process | Subhashini.Satishkumar@Sanofi.com | 51 and 45 | 2022-09-21 | 2022-09-21 | Logbooks should have been found upon search
446 | No reportable results for Viable Air, Rodac A, and Rodac B for Sites 131-1 and 130B-2 in 51NYA (multiple biolims numbers) | marissa.lentoni@sanofi.com | 51 NYA, Room 131 and Room 130B, Site 131-1 (grade C) and Site 130B-2 (grade C) | 2022-05-09 | 2022-09-20 | Samples would be included in EM collection batch that biolims associates them with and results would be recorded for these sites. | No | Deviation
445 | Extension paperwork upload missing from INFOR for work order 2526858 | angela.sauro-havalotti@sanofi.com | 51-132 | 2022-03-31 | 2022-09-06 | FBL-SOP-000128, step 6.10.15, scan associated documents into "document tab" if applicable, prior to closure of work order
444 | No reportable viable air results for routine EM site CC2-77 in 74 NYA (biolims 491278) | marissa.lentoni@sanofi.com | 74 NYA, Room CC2-01/03, Site CC2-77 Viable Air (Grade D) | 2022-09-16 | 2022-09-19 | Plate would be included with the remaining samples collected that day and incubated upon arrival back to the micro lab. | Yes | Deviation
443 | Incomplete pH Data Upload into OneLims | nancy.watkins@sanofi.com | 68NYA, QC Chemistry Lab | 2022-09-08 | 2022-09-20 | Temperature only recorded , no requirements stated in ALF-MTD-000045.
442 | Filter for 768506 while being moved after transfer to FIT | aidan.mula@sanofi.com | 8NYA RM1350 | 2022-09-20 | 2022-09-20 | Determination on path forward
441 | sample 2200045526 retested before LE was reviewed | Stefanie.Erickson@sanofi.com | QC Micro water chemistry bay, TOC | 2022-09-16 | 2022-09-20 | for the LE to be reviewed before the TOC sample was repeat tested.
440 | Meeting to Discuss missing RQ documentation | DANIEL.BATI@SANOFI.COM | 51 NYA ROOM | 2022-09-08 | 2022-09-19 | It is expected that the RQ and associated documentations be completed and return the RQ document to validation. FBL-SOP-000942
439 | Raw NaCl Part# GZ1335-04 is hard to scoop | aidan.mula@sanofi.com | 8NYA RM1364 | 2022-09-20 | 2022-09-20 | NaCl should not be hard and require extra effort to scoop.
438 | 74 CU 502512 74HWFI-25 double tested for LAL by Nexus instrument | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2022-09-17 | 2022-09-17 | It is expected that one reportable results is available for a sample
437 | Missing RQ documentation for sensor array X-3350 located in room 138 | sylvester.kegbeh@sanofi.com | 51 NYA ROOM RM room | 2022-09-08 | 2022-09-19 | BL-SOP-000942
436 | Buffer temperature not chilling for zinc operation | kelvin.leo@sanofi.com | 8 NYA | 2022-09-19 | 2022-09-19 | Buffer will chill and recipe will not get in hold |  | No Deviation
435 | No report non-viable result for NBO_BFH-NVC Center (sample 364576) | marissa.lentoni@sanofi.com | NBO, NBO 1204, Site NBO-BFH-NVC Center | 2022-05-13 | 2022-09-16 | The sampler would observe that a failing flow rate was generated and would resample the site to ensure valid results were generated for the site.
434 | No reportable LAL result for PW-2 for week of 22May (sample 375594) | marissa.lentoni@sanofi.com | NBO, Room 1206, Site PW-2 | 2022-05-25 | 2022-09-16 | The site would tested for LAL and valid results would be generated.
433 | GEMBA 74NYA - COLUMN PACKING - FOUND BUG IN BAG | william.bradley@sanofi.com | 74NYA PU1-08 | 2022-09-19 | 2022-09-19 | Material to be clean
432 | Missing verification Steps FBS-MBR-000093 (CBYHAH005 Steps 7.1.16; 7.11.17; and 7.11.18); FBS-LBK-000030 (3 steps in page 10/47); and FBS-LBK-000033 (2 steps in page 2/27) | dena.elmesalamy@sanofi.com | 76NYA 3101/ 3103 | 2022-06-06 | 2022-08-06 | Completed verification steps on time
431 | Nanofilter Wetting recipe on hold due to filter differential pressure controller low-low alarm | alejandro.martinez2@sanofi.com | 8NYA RM:1323 | 2022-09-17 | 2022-09-17 | Automation intervention to be able to restart recipe and proceed with nanowetting.
430 | HA - CBYFRM034 / CBYHAH007; FBS-MBR-000067, step 7.2.7.6; Component not within specified range | charlie.houle@sanofi.com | 76 NYA, room 364 | 2022-08-30 | 2022-09-16 | Per FBS-MBR-000067, step 7.2.7.6 requires 1251 grams with an acceptable range of 1238.49g to 1263.51g.
429 | Stability predictive adverse trend : HA Lot BBYHAH001 adverse trend for LOD Vac | elizabeth.couturier@sanofi.com | N/A | 2022-09-16 | 2022-09-16 | It is expected that predicted stability is within specification per stability protocol | Yes | Deviation
428 | Abnormal sparge rates for production reactor X2301 oxygen | paul.mclaughlinjr@sanofi.com | 74NYA/COVID | 2022-09-16 | 2022-09-16 | All dissolved gas readings are normal for both offline and inline readings
427 | Automation intervention needed during MMC chrom skid CV | michael.daly@sanofi.com | 8 NYA Room Downstream Ballroom 1503, CNC 1212 | 2022-09-16 | 2022-09-16 | It is expected the correct valves
426 | No reportable results for weekly Critical Utility (CU) WFI sites 51-WFI-1A and 51-WFI-1C at 51NYA | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2022-09-10 | 2022-09-16 | It is expected that 51-WFI-1A and 51-WFI-1C are collected and tested on a weekly basis per
425 | 45NYA-Rm 1136-QCC45-product impact unknown- FB-2513 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 45-1136 | 2020-07-16 | 2020-07-16 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
424 | Bioreactor X2301 went on hold due to high high DO ( 68.8 ) the range is 55.0 - 65.0 | banet.simon@sanofi.com | 74 NYA /CC2-01/ X2301 | 2022-09-15 | 2022-09-15 | After Inoculation, the DO should be at 60 range 55.0 - 65.00
423 | GEMBA: Incorrect dilution documented for TSH FP, 36M TP | Kaylyn.Manoogian@sanofi.com | 45NYA, QC Chemistry Lab | 2022-08-02 | 2022-09-14 | The dilution should have been 250uL of sample for a 4x dilution per FBL-MTD-000050.
422 | WFI expired on MMC skid for CV | cory.lefort@sanofi.com | 8 NYA Downstream phase 2 | 2022-09-15 | 2022-09-15 | Expected that WFI would not be expired during the entire operations/CV. | Yes | No Deviation
421 | EMS tag F076-MI-9526-005 for room 355 Humidity set incorrectly per GEODE IS document | Stephen.walker@sanofi.com | Framingham | 2018-09-26 | 2022-09-13 | Per the configuration document in Geode IS (CDS-109-06), the Low Alarm setpoint should have been set at 15% RH | Yes | Deviation
420 | No mold counts recorded for Biolims sample 354094 | kathryn.lentoni@sanofi.com | 74NYA Room EX1-28 Site EX1-28-76 | 2022-06-04 | 2022-09-12 | Mold counts would be recorded by counting analyst.
419 | No reportable conductivity results for 2200037559 (WFI 144) 2200037739 (WFI 144) 2200035540 (76-WFI-2) 2200037741 (Pure 219) | kathryn.lentoni@sanofi.com | 76NYA Room 347 Pure 219. 76NYA Room 352 WFI 144. 76NYA Room 420 (76-WFI-2) | 2022-06-28 | 2022-09-12 | Samples would be tested for conductivity and would generate valid results
418 | HA Ferm – CBYFRM027 – 76NYA – Incorrect times were recorded and changed without justification | dena.elmesalamy@sanofi.com | 76 NYA- room 364 | 2022-06-21 | 2022-09-12 | Provided the source of the corrected times.
417 | No reportable results for WFI 144 for 11Sep22 | marissa.lentoni@sanofi.com | 76 NYA, Room 352, Site WFI 144 (HA component prep) | 2022-09-11 | 2022-09-12 | Sampler would be able to enter the biosurgery space to collect the sample.
416 | 76NYA-Room 355 Biosurgery HAFL-EMS Humidity Indicator found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76NYA room 355 | 2022-09-12 | 2022-09-13 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
415 | Probes installed in a different order of operations than dictated by EWI during Cz SUB build operations | john.lytle@sanofi.com | Rm 1434 | 2022-06-23 | 2022-09-12 | probes to be installed per order of operations listed in the EWI, EBR
414 | no verifiable source for 2 aseptic challenge incubations | Stefanie.Erickson@sanofi.com | QC Micro counting | 2022-09-03 | 2022-09-11 | for the incubation to be recorded for F-15594 | Yes | Deviation
413 | Multiple (5) HA-CMC bottles were discarded by Quality Control Biopolymers prior to completed testing | Craig.Stevenson@sanofi.com | 76 New York Ave, Room 107, QCBP | 2022-09-12 | 2022-09-12 | It is expected that testing bottles are not discarded until all testing has been completed.
412 | Site 8-WFI1-5 has possible unreportable results due to discrepancies in the time section for stage 2 conductivity testing | kathryn.lentoni@sanofi.com | 8 NYA | 2022-09-08 | 2022-09-12 | A calibrated timer or BioLIMS component timer and determine the appropriate time increment of either 1 or 5 minutes. After each timed 1 or 5 minute increment record the conductivity, temperature, and time information of all subsequent readings on FBL-FRM000413. | Yes | Deviation
411 | CZ SUB X-2250 Bleed Capacitance EM holds SUB recipe approximately every couple hours | lolav.ahmed@sanofi.com | 8NYA room 1403 | 2022-09-12 | 2022-09-12 | FBL=CST-000029
410 | Butyl Agitator Set to Turn On Prior to Ammonium Sulfate Addition Completion | mark.smith@sanofi.com | 8NYA-1333 | 2022-09-11 | 2022-09-11 | Agitator setpoint should remain at 0 until completion of ammonium sulfate addition.
409 | Buffer leak in CNC area | cory.lefort@sanofi.com | 8 NYA | 2022-09-10 | 2022-09-10 | N/A
408 | 2 bioburden samples with same biolims number | Stefanie.Erickson@sanofi.com | 8 NYA | 2022-09-10 | 2022-09-11 | for there to be 1 sample for 1 biolims number
407 | Missed Witness Step Buffer Prep PH Titration | Jeff.hill@sanofi.com | 45 NYA | 2022-07-31 | 2022-09-10 | FBL-MBR-000036, Step 4.15, Titration should be witnessed
406 | PC02 deviation alarm for X2210 | banet.simon@sanofi.com | 74 NYA /CC2-01/ X2210 | 2022-09-10 | 2022-09-10 | PCo2 probes must be within 7 points range probe A currently at 14.2 and probe B at 22.6
405 | Air Handler 7 down at 51NYA causing DP alarms for 149, 152 | brian.mcgee@sanofi.com | 51NYA/149, 152, 153 | 2022-09-09 | 2022-09-09 | It is expected for positive pressure to be maintained in each room and the pressure criteria to remain in the acceptable ranges |  | Deviation
404 | A NewAge filter assembly (NewAge Part# CA7622124, Sanofi P/N 691058 Batch 21GN01885) was used for Cerezyme-ICB and Fabrazyme-ICB GxP Manufacturing without Suitability for Use (SFU) being completed under CCR 2020FRACC0104, per FBL-SOP-001476 version 1.0 (System Suitability SOP for Framingham Biologics and Biosurgery). | benjamin.larson@sanofi.com | 8 NYA Downstream | 2022-05-10 | 2022-09-07 | It is expected the Suitability for Use (SFU) should be completed prior to Manufacturing use per FBL-SOP-001476 version 1.0 (System Suitability SOP for Framingham Biologics and Biosurgery), Step 6.1.1. | No | Deviation
403 | Incorrect Part # on MBR | alexandre.silva2@sanofi.com | 76 NYA | 2022-03-08 | 2022-09-09 | Per FBS-MBR-000074, Step 7.4.2.6 operators are required to use Filter # 5988-01, Comment Require P# 5989-01 and .P# 5889-01 was used.
402 | OneLIMS Bioburden sample 497692, COV_MMC_827588_POU827755_BIOB expired before delivery to QC Micro | paul.consolmagno@sanofi.com | 74 NYA | 2022-09-09 | 2022-09-09 | The Bioburden sample would be put on test/incubated before the 24 hour expiration
401 | UF/DF Skid was drained and then the UF/DF operation process was canceled | alyson.belleville@sanofi.com | 51 NYA Room 88 | 2022-09-09 | 2022-09-09 | FBL-MBR-000112 operations would be able to be performed as written |  | No Deviation
400 | On 08SEP22, X-3051 (MMC skid) experienced a Hi Hi pressure alarm (PI-305012A) causing Capto MMC production recipe to go into hold during sanitization phase. | craig.valery@sanofi.com | 74 NYA, PU1-09 | 2022-09-08 | 2022-09-08 | It is expected that the recipe would continue without any hold conditions.
399 | 45NYA-Rm 1136-QCC45-product impact unknown- no as found collect due to unit failure | robert.bauch@sanofi.com | 45-1136 | 2022-08-16 | 2022-09-08 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been able to be collected
398 | Temp probe on Fz SP Pooling sum is not functioning properly | robert.morawski@sanofi.com | 8NYA RM1503 | 2022-09-08 | 2022-09-08 | temperature probe reads properly. |  | Deviation
397 | Incorrect Bag size used during Thyrogen Blue fraction 8 collection | jacob.kallon@sanofi.com | 51 NYA room 138 | 2022-09-07 | 2022-09-08 | It is expected that bag part number GZ6049-03 be used to collect blue fraction 8
396 | Expired Sample bag (692165) used for NRSR 22SR8024 | jarad.kukla@sanofi.com | 8NYA / 1313 | 2022-09-07 | 2022-09-07 | Material should not be expired |  | Deviation
395 | COVID positive QC contractor | elizabeth.couturier@sanofi.com | see above | 2022-09-04 | 2022-09-07 | It is expected that COVID-19 positive personnel do not enter controlled/classified areas.
394 | 51 NYA - Thyrogen- Missing WIP Label verification in MBR (CGF4B04) | ian.odell@sanofi.com | 51 NYA Room 140 | 2022-05-27 | 2022-09-07 | Per FBL-MBR-00091, step 4.44.3 and 4.45.4, WIP labels attached to record and harvest bag require a secondary verification | Yes | Deviation
393 | Fz nano skid setup incorrectly | kelvin.leo@sanofi.fcom | 8 NYA | 2022-09-07 | 2022-09-07 | setup accurately |  | No Deviation
392 | Fz X-2210: BGA results indicates daily sample is unrepresentative | Lolav.ahmed@sanofi.com | 8NYA Room 1333/1334 | 2022-09-07 | 2022-09-07 | FBL-CST-000028 Table 7.5, pH should be between 6.8-7.0 on day 12-60.
391 | Fz Nano skid setup incorrectly | robert.morawski@sanofi.com | 8NYA RM1503 | 2022-09-07 | 2022-09-07 | Correct connection made, and set up procedure resumes. |  | No Deviation
390 | Missing swab and double rinse tested for 2022-CV-0527 | megan.botti-marino@sanofi.com | 68 NYA | 2022-08-27 | 2022-09-06 | Per documentation for 2022-CV-0257 and the associated test request, 8 swabs were submitted for testing. On the attached assay sheet, only 7 swabs were tested. It is expected that all 8 swabs would be tested. Only one rinse sample was submitted per the associated test request, but there are two results for rinse samples. It is expected there there would only be one result. | No | Deviation
389 | CZ SUB X-2220 pH Loss During PA Sparge | michael.knapp@sanofi.com | 8NYA 1403 | 2022-09-06 | 2022-09-07 | Process Air accurately controls pH
388 | Thyrogen Finished Product Sample Result incorrectly reported on COA for Oxygen Headspace | Kaylyn.Manoogian@sanofi.com | 45NYA, QC Chemistry Laboratory | 2022-01-24 | 2022-09-02 | Validated results would have been <0.2% | Yes | No Deviation
387 | Chase buffer ran out during Cz Nano filtration | Pavel.dobrovolskiy-wynes@sanofi.com | 1503 | 2022-09-05 | 2022-09-05 | Its expected that there will be enough chase buffer hooked up to complete the operation
386 | Hyaluronic Acid (HA) - lots unknown- 76 NYA rooms 3101 and 3102 COVID positive employee in HA purification | kenneth.morgan@sanofi.com | 76 NYA rooms 3101, 3102 | 2022-08-30 | 2022-09-06 | It is expected that personnel would not enter classified manufacturing space when they have an illness per FBS-SOP-000667.
385 | NewAge Sterility claim discrepancy | benjamin.larson@sanofi.com | NA | NA | 2022-09-02 | NA
384 | Upstream employee tested covid positive on 02Sep. 2 Weeks previous was19Aug2022. Occupational Health notified Friday 02Sep2022 will return to work 12Sep2022 | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2022-09-02 | 2022-09-02 | Covid + employees not enter controlled space |  | Deviation
383 | SUM1515 pump is not working, 760736 buffer for Q membrane in question mix timer of 10 hours exceeded in DeltaV resulting in Hold. MFG Alarm cleared | aidan.mula@sanofi.com | 8NYA, RM1350 | 2022-09-05 | 2022-09-05 | Determination if batch can be used passed 10 hour timer and if pump will function. |  | Deviation
382 | Bug found in Suite 2 Zinc row CNC space | stanley.durand@sanofi.com | 8NYA room 1212 | 2022-09-04 | 2022-09-04 | Per FBL-SOP-000512 step 6.3 placed pest in a container, treat area immediately with IPA, complete Attachment 3 and generate work order.
381 | 51 NYA - Thyrogen- Harvest storage discrepancy for LN CGF1S09B SP1 process | ian.odell@sanofi.com | 51 NYA Room 108C | 2022-08-10 | 2022-09-03 | All storage of harvest material should be recorded on BIRFs
380 | Agitator Motor Reading Failure on CZ X-2220 | michael.knapp@sanofi.com | 8NYA 1403 | 2022-09-03 | 2022-09-03 | Agitator speed readings are consistent and accurate.
379 | P/N 713985 Aquesyn 8-port manifold misaligned with 15ft extension line during Zinc blank run CNC operations. | stanley.durand@sanofi.com | 8NYA room 1212 | 2022-09-02 | 2022-09-02 | It is expected for the Aquesyn valve to be properly connected and aligned to the manifold prior to removing the tab | Yes | No Deviation
378 | Tote set up performed in wrong MO RF550C013 and RF550C014 | richard.manasian@sanofi.com | 1350 8NYA | 2022-09-02 | 2022-09-02 | Correct record should be used to set up tote bag for transfer
377 | Contractor entered CNC with improper gowning through MAL 1422 | lolav.ahmed@sanofi.com | 8NYA CNC | 2022-09-02 | 2022-09-02 | Per FBL-EWI-000002
376 | P/N 6921693 - 3K SUM BAG LEAK | alexis.tello@sanofi.com | 8NYA | 2022-09-02 | 2022-09-02 | path forward
375 | End of run wave stage Endotoxin samples leaked | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2022-09-01 | 2022-09-01 | Vial should be secure and not leaked
374 | Butyl CNC buffers mixed together due to operator removing clamps prior to valves being closed | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA1210 | 2022-09-01 | 2022-09-01 | Its expected that claps stay on while manifold valves are still open
373 | Bioburden action level for 2022-CV-0517 (Chrom 1 Xfer Line) | alvaro.rueda@sanofi.com | 74 NYA Falcon | 2022-09-25 | 2022-09-01 | 0 CFU
372 | QA Observation missed for Nanofilter change out | Mackenzie.gomes@sanofi.com | 8 NYA | 2022-07-10 | 2022-09-01 | QA Operations would observe filter changeout and document observation
371 | Doc #: RQ - F051-SV-51NYA-CQ-0002, RQ Physical Testing for COP of Slurry Vessel V-3373 was not executed by the month due of 31AUG22 due to Soil Unavailability | john.sicurella@sanofi.com | 51NYA | 2022-09-01 | 2022-08-25 | Doc #: RQ - F051-SV-51NYA-CQ-0002, RQ Physical Testing for COP of Slurry Vessel V-3373 should have been executed by the month due of 31AUG22 | Yes | Deviation
370 | Cerezyme – Lot CGHBN008 – 8 NYA – Offline pH reading out of range high and pO2 reading out of range low prior to initiating pH Control | sheila.pandolfo@sanofi.com | 8 NYA, Room 1403 | 2022-09-01 | 2022-09-01 | It is expected pH and pCO2 results would be in range giving a valid result. | No | Deviation
369 | Meeting to Discuss FBL-MBR-000114 Step 4.2.1.1 | sylvester.kegbeh@sanofi.com | 51 NYA ROOMS 88 and 87 | 2022-09-01 | 2022-09-01 | FBL-MBR-000114
368 | UFDF load protein concentration is out of spec. Spec is 368-920 actual concentration is 937.96 | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA 1323 | 2022-09-01 | 2022-09-01 | Its expected that load concentration will be within range of 368-920
367 | 51NYA-room 124-Dept APT SS-product impact unknown-GMP Critical pipette F-17853 overdue for calibration | robert.bauch@sanofi.com | 51-124 | 2022-09-01 | 2022-09-01 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.7.3 Instrument should have been received into Metrology before it's due date
366 | 45NYA Loss of Compressed air (utilities) | brizaida.medina-vega@sanofi.com | 45NYA Manufacturing & GLP Distribution | 2022-08-31 | 2022-08-31 | Dewpoint ≤ -4 ⁰F
365 | 51 Annual WFI Cooler PMs Unable to Be Completed in August | marc.fages@sanofi.com | 51 NYA | 2022-09-01 | 2022-08-05 | PM completed in August
364 | Unable to locate CV form for 2022-CV-0536 | alvaro.rueda@sanofi.com | 68NYA | 2022-08-27 | 2022-08-27 | Form | Yes | Deviation
363 | NBO Nitrates samples not tested within the 7 day testing window | kathryn.lentoni@sanofi.com | 68NYA 4700, QCM | 2022-08-22 | 2022-08-31 | The samples are expected to be tested within 7 day window
362 | No reportable results for UPLC concentration Fz eluate | Brandon.Steff@sanofi.com | 8RM1334 | 2022-08-28 | 2022-08-31 | concentration should be greater than 3.1 from EWI this IPC FBL CST-000028
361 | Failed bioburden for 2022-CV-0510 (DSP Process Hoses) | alvaro.rueda@sanofi.com | 68/74NYA | 2022-08-20 | 2022-08-30 | Passing Bioburden results
360 | 51 NYA - Thyrogen- Pest found in classified area (room 140) | ian.odell@sanofi.com | 51 NYA Room 140 | 2022-08-30 | 2022-08-30 | Expected that the classified areas be free of pests.
359 | 2 pests discovered in 8NYA Room 1301, outside of Room 1301 and 1400 | Craig.Stevenson@sanofi.com | 8NYA Room 1301, 1250 and 1400 | 2022-08-30 | 2022-08-30 | It is expected that pests are not in classified area
358 | X-2230 recipe hold when transferring out due to DO levels | david.loik@sanofi.com | 74NYA | 2022-08-30 | 2022-08-30 | DO levels would be maintained at setpoints during transfers.
357 | Discrepancy on filter Lot# entered on BR and consumed on Athena. | alexandre.silva2@sanofi.com | 76 NYA (Purification). | 2022-05-02 | 2022-08-29 | It's expected that Athena transaction reflect the same Material/Lot# used on the BR. | Yes | Deviation
356 | Unable to Run Chrom 2 Transfer Line CIP due to Recipe running on V-2561 | alvaro.rueda@sanofi.com | 74NYA | 2022-08-30 | 2022-08-30 | CIP recipe to run to completion
355 | Wave lot#CGHCR065 was sampled over the 78hour required time | adam.rowe@sanofi.com | 74NYA, CC2-01 | 2022-08-29 | 2022-08-29 | Waves need to be sampled no later than 78 hours after inoculation
354 | Urgent - Seprafilm - CBYSEP055- 76 NYA 231 - Affected supplier part number is 6164-49 lot 22MN00178 | fatima.fragaalvarez@sanofi.com | 76 NYA room 231 | 2022-08-29 | 2022-08-29 | BQ-GOP-000247 material was contained and production was stopped.
353 | 8NYA CZ2 POU Buffer (468234 CZ2_PHNL_761220_POU760728_BIOB) sampled and submitted in duplicate to QCM | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2022-08-29 | 2022-08-29 | Samples should be sampled per FBL-RPT-000538. If a sample requires a resample, samples need to be resampled per OneLIMS SOP instructions and not just labels reprinted.
352 | 2k Reactor Hold post media load for HIHI Temperature | paul.mclaughlinjr@sanofi.com | 74NYA | 2022-08-29 | 2022-08-29 | Temperature would fall within 27.0 +/- 2 during stabilization time per FBL-CST-000030
351 | Not Completing Hot Sanitization at 74NYA Due to Safety Concerns as Part of Investigation Steps for FRAD22E0705 | tanni.arif@sanofi.com | 74NYA | 2022-08-26 | 2022-08-26 | Per FBL-SOP-000517, Step 6.4.3: hot water sanitization is required if invasive work is required at or downstream of RO Membrane system.
350 | HA-CMC - Lot #(s) CBYCMC171, CBYCMC172; AHU 2A shut down due to power blip | charlie.houle@sanofi.com | 76NYA / rooms 224, 227, 229 | 2022-08-26 | 2022-08-26 | AHU-2A should run continuously.
349 | P/N 714031 1/2" PUMP TBG W/SGT & ASEPTQCK.,BIPRENE ruptured during transfer of Capto MMC Eluate into V-2561 | alvaro.rueda@sanofi.com | 74 NYA Falcon | 2022-08-29 | 2022-08-29 | No leak/rupture to occur
348 | ONC Pouch Label 849140 does not meet spec NBO-SPC-000138 | alan.rousseau@sanofi.com | room 1009 | 2022-08-29 | 2022-08-29 | NBO-SPC-000138 | Yes | No Deviation
347 | MMC Skid X3304 failed CV | Pavel.dobrovolskiy-wynes@sanofi.com | 8NYA Room 1503 | 2022-08-28 | 2022-08-28 | Its expected that all swab samples will pass TOC test
346 | Broken pins on fitting connection of temperature indicator TI-29102 | edgar.vale@sanofi.com | 51NYA, room 139 | 2022-08-27 | 2022-08-27 | Pins on fitting would be intact to allow temperature readings
345 | X2230 500L sub lot#CGHS5019 | duane.messier@sanofi.com | 74NYA/CC1-07 | 2022-08-27 | 2022-08-27 | Recipe should run without operator intervention
344 | Short Electrical Interruption at 74 NYA due to storm | amanda.sorel@sanoif.com | 74 NYA | 2022-08-26 | 2022-08-26 | There would be no electrical interruptions | Yes | Deviation
343 | Short Electrical Interruption at 45 and 51 NYA due to storm | Alyson.Belleville@sanofi.com | 51 NYA | 2022-08-26 | 2022-08-26 | There would be no electrical interruptions | Yes | No Deviation
342 | GEMBA: 74NYA - DSP Depth Filtration Chill Water Valve and Pump | william.bradley@sanofi.com | 74NYA CC1-03 | 2022-08-26 | 2022-08-26 | Process is to run as designed
341 | Daily sporklenz cleaning was not documented for 25AUG22 per FRAD22E0398. | adam.hildebrand@sanofi.com | 74 NYA/ CC1-03 | 2022-08-25 | 2022-08-26 | It is expected that the daily cleanings would be performed and documented accordingly
340 | Incorrect dilution/titer approved and used by QC and QCTS for the Fz ICB CHO-S coat 20TR0497 | paul.consolmagno@sanofi.com | 68 NYA, QCTS and QCM | 2020-12-08 | 2022-08-24 | The material would be prepared at a concentration of 5ug/mL (32 uL in 11mL) | Yes | Deviation
339 | No UV trend visible on DeltaV during elution of Q CGF4B04A process | alyson.belleville@sanofi.com | 51 RM 131 | 2022-08-26 | 2022-08-26 | It is expected that the UV trend would be visible on DeltaV while the product is coming off the column and that it would record a normal trend appropriately in PI. | Yes | Deviation
338 | Followup GEMBA for leaking ATF B Guard Filter | john.lytle@sanofi.com | Rm 1434 | 2022-08-25 | 2022-08-25 | integral system
337 | MCC Skid X2543 UV Cal Column 3: Low Post Col UV measurements | lolav.ahmed@sanofi.com | 8NYA Room 1333 | 2022-08-25 | 2022-08-25 | Column 3 should reads within the spec of col 1 & 2
336 | Sodium Hypochlorite disposed of during weekly cleanings in rooms 3102, 3103, 224, 227, and 229 | nicholas.macdonald@sanofi.com | 76 NYA rooms3102, 3103, 224, 227, and 229. | 2022-08-23 | 2022-08-23 | NO BLEACH solutions are to be disposed of into any floor drain in rooms: 3102, 3103, 224, 227 and 229. Bleach can only be discharged to the floor drains in rooms: 364 and 3101. | Yes | Deviation
335 | ATF B guard filter is leaking from filter outlet for Cz | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2022-08-25 | 2022-08-25 | Gruard filter should not be leaking
334 | No reportable results for multiple (3) Environmental Monitoring (EM) Non-Viable (NV) sites at 74NYA | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2022-04-27 | 2022-08-24 | It is expected that the climet have successful parsing where all the required data is transferred successfully into BioLIMS.
333 | 8NYA Supplier Deviation Needed: 692169 recirc line assembled incorrectly | alejandro.martinez2@sanofi.com | 8NYA Solution Prep, 1350 | 2022-08-25 | 2022-08-25 | Supplier Deviation Needed
332 | 100L SUB Hold due to HIHI DO | paul.mclaughlinjr@sanofi.com | 74NYA | 2022-08-24 | 2022-08-24 | FBL-CST-000030
331 | N2 shortage for 8 NYA. Fabrazyme SUB using N2 for pH control | Bhargav.Patel@sanofi.com | 8 NYA / 1333 | 2022-08-24 | 2022-08-24 | N2 gas should not be run out
330 | Wrong part number received for Bulk gas GZ1331-02 GZ6691-01 received | steven.servais@sanofi.com | 80 New York Av. Framingham, MA. | 2022-08-23 | 2022-08-23 | NBO-SOP-000122 Testing and Dispositioning Bulk Liquids and Gases. Step 6.1
329 | Expired FAD used to generate ATTO-TAGTM CBQCA protein assay for Butyl, SP, Zinc, and WFI eluate for Fabrazyme (Fz), at Eurofins (FAD13-160, version 3) | damian.barbato@sanofi.com | 8NYA | 2022-08-22 | 2022-08-22 | Expired FAD would not be used for GMP purposes | Yes | No Deviation
328 | Low step recovery of 92% for TSH Formulated Bulk CGF2B02, range is 93-104% | alyson.belleville@sanofi.com | 51 NYA | 2022-08-22 | 2022-08-23 | It is expected that the recovery for the operation would be within range specified in batch record. | Yes | Deviation
327 | HA - FBL-SOP-000941 Attachments used in 'Approved' state | charlie.houle@sanofi.com | 76 NYA, HA Fill Areas | 2022-04-13 | 2022-08-15 | All data should be entered on most current effective version.
326 | Total mix time for pH adjust not met CGHMS012 FBL-MBR-000131 PN 794568 | dianne.moustafa@sanofi.com | 74NYA CC1-08 | 2022-07-02 | 2022-08-21 | It is expected batch mixes 10 minutes after adding 2N HCL during pH adjust
325 | MCC skid 2811 on hold over 7 hours, Column 3 elution send to drain. | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2022-08-22 | 2022-08-22 | Automation, MSAT and MFG team trouble shoot pressure issue, put column 2 by pass flow through buffer to ensure buffer path is clear. put column back in inline and run manual buffer through column. Monitor pressure and flow through. Once it stabilize, clear the hold from recipe and re-start recipe.
324 | Process Temperature is identified to be out of range during chromatogram review from PI | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 132 | 2022-08-13 | 2022-08-22 | It is expected that the process temperature is between 6-10C per FBL-MBR-00011.
323 | SUB ATF Changeout Automation Intervention Required | michael.knapp@sanofi.com | 8NYA 1403 | 2022-08-20 | 2022-08-21 | The ATF's are above the alarm limit threshold and ATF changeout is possible without Automation intervention.
322 | MCC X-2811 Buffer Flow Alarm and Air in Bubble Trap | michael.knapp@sanofi.com | 8NYA 1403 | 2022-08-21 | 2022-08-21 | MCC system maintains proper flow for duration of campaign. Equilibration buffer is not allowed to run dry.
321 | High Conductivity Alarm on RO Supply and Return loop at 74 NYA. | Brizaida.MEDINA-VEGA@sanofi.com | 74 | 2022-08-21 | 2022-08-21 | The alarm set point in Delta-V is 1.1 and a quality Critical is set point is 1.3 and the system must operate below that set point or it will be in alarm per FBL-SOP-000316 attachment 3
320 | Clean expired kit used for process | william.lee@sanofi.com | 74NYA | 2022-07-11 | 2022-08-20 | Media Prep Tank kit should have been used within 30 days of cleaning per FBL-FRM-000469. COP kit batch ID 20220608.075704 was used on 11JUL22 | Yes | Deviation
319 | Cz Zinc Product transfer filter is clogged | robert.morawski@sanofi.com | 8NYA RM1503 | 2022-08-20 | 2022-08-20 | Cz Zinc load transfer completes without pump stopping.
318 | On 19AUG22, X-3051 (MMC skid) experienced a Hi Hi pressure alarm (PI-305012A) causing Capto MMC production recipe to go into hold during sanitization 2 phase. | craig.valery@sanofi.com | 74 NYA, PU1-09 | 2022-08-19 | 2022-08-19 | It is expected that the recipe would continue without any hold conditions.
317 | FZ SUB X-2210 ATF Failure | michael.knapp@sanofi.com | 8NYA 1333 | 2022-08-19 | 2022-08-19 | The ATFs continue to operate without failure. |  | Deviation
316 | CZ Thaw Sample Vicell Error | michael.knapp@sanofi.com | 8NYA 1343 | 2022-08-19 | 2022-08-19 | Vicell samples with the correct cell type and Sample ID.
315 | FZ Thaw Sample Vicell Error | michael.knapp@sanofi.com | 8NYA 1343 | 2022-08-19 | 2022-08-19 | Vicell samples with the correct cell type and Sample ID.
314 | Multiple HiHi DO alarms during transfer out on X2230 | david.loik@sanofi.com | 74NYA | 2022-08-19 | 2022-08-19 | DO would not cause an alarm.
313 | Missed QA signature for vial receipt at 8NYA | Adriana.Alban@sanofi.com | 8NYA | 2022-08-05 | 2022-08-18 | It is expected that QA witnesses all vial receipts | Yes | Deviation
312 | Product impact unknown-51NAY- FFO3 Room 146-TOC F-11538 found out of tolerance during a scheduled calibration. | robert.bauch@sanofi.com | 51-146 | 2022-08-15 | 2022-08-18 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
311 | Falcon BCA Samples Double Tested | john.kapaniris@sanofi.com | 68 NYA QC Chemistry | 2022-07-23 | 2022-08-02 | Falcon SDS/Western Blot samples were used in Falcon BCA Assay
310 | X2301 HiHi Temperature alarm after media load | Kevin.devanna@sanofi.com | 74NYA / CC-201 | 2022-08-18 | 2022-08-18 | Per FBL-CST-000030, Pg. 44, PSFM Equilibration Temperature 27 C +/- 1 C
309 | Product impact unknown-74NYA-room PU1-05-Floor HEPA FB-1301 found out of tolerance during a scheduled calibration for excessive Penetration | robert.bauch@sanofi.com | 74-PU1-05 | 2022-08-16 | 2022-08-18 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
308 | Trainer countersigned for performed by and verified steps, Missing trainer signatures | rebecca.Cunningham@sanofi.com | 76 NYA RM 507 | 2022-06-21 | 2022-08-18 | Trainer counter signs for all tasks performed by trainees and a separate person verifies | Yes | Deviation
307 | Plate readers annual PV missed required QA approval before being put back into services | kin.mui@sanofi.com | 68 New York Avenue, Room 3700, QCTS lab | 2022-02-14 | 2022-08-09 | Per FBL-SOP-000757 step # 6.10.4 it is expected to have QA's approval within 30 days of QC review of data. | Yes | Deviation
306 | 76 A/N 2022-R-021 Number of copies requested wasn't filled out | Subhashini.Satishkumar@Sanofi.com | 76 WFI Steam | 2022-07-08 | 2022-08-17 | A/N need to be released
305 | 51 NYA AHU 3 down | jacob.kallon@sanofi.com | 51NYA 130, 130A, 131 | 2022-08-17 | 2022-08-17 | It expected that the Air Handling Unit is properly functioning to allow for production to continue
304 | Product impact unknown-74NYA-room 74-Cryostorage GCB department-F-14465-300 EMS for Temperature channel A no As found data | robert.bauch@sanofi.com | 74-Cryostorage | 2022-08-16 | 2022-08-17 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been able to be collected |  | Deviation
303 | Thyrogen Drug Product Lot AW0145 24M Stability Time Point- No Result (NR) for Isoaspartic Acid and Free Subunit | yekaterina.zavgorodniy2@sanofi.com | 45 QCC | 2022-03-20 | 2022-08-16 | It is expected that Isoaspartic and Free Subunit Testing is requested at 24 month time point per FBL-PTM-000049
302 | 8 NYA - Employee Tested Positive for COVID-19 who entered the 8 NYA GMP Manufacturing Area | sheila.pandolfo@sanofi.com | 8 NYA Downstream, Rooms TBD | 2022-08-14 | 2022-08-15 | It is expected that COVID-19 positive personnel do not enter controlled/classified areas.
301 | On 16AUG22, X-3051 (MMC skid) experienced a Hi Hi pressure alarm (PI-305012A) causing Capto MMC production recipe to go into hold during sanitization phase. | craig.valery@sanofi.com | 74 NYA, PU1-09 | 2022-08-16 | 2022-08-16 | It is expected that the recipe would continue without any hold conditions.
300 | GMP CRIT Micro Balance Weight Set, FB-1518, SN 03304 is missing the 10mg weight. | NANCY.WATKINS@SANOFI.COM | 45NYA, QC Chemistry lab, room 1136 | NA | 2022-08-08 | Per BQ-GOP-000046 step 5.6.4 as found data will be collected prior to calibration date of 30Sep22.
299 | CZ SUB X-2250 high foam levels discovered | michael.knapp@sanofi.com | 8NYA 1403 | 2022-08-16 | 2022-08-16 | The automated antifoam dosing strategy is sufficient to maintain foam levels in the Bioreactor
298 | Expired Part (721472) used during Thyrogen UF/DF Operation | jacob.kallon@sanofi.com | 51 NYA/room 88 | 2022-08-12 | 2022-08-16 | It is expected that all parts be within expiration date at time of use | Yes | Deviation
297 | counts not recorded for CV bioburden | Stefanie.Erickson@sanofi.com | QC Micro counting | 2022-08-13 | 2022-08-15 | for there to be results recorded for the negative control associated with the bioburden run.
296 | FZ SUB Inoculation and Perfusion start timeline follow-up | michael.knapp@sanofi.com | 8NYA 1333 | 2022-08-12 | 2022-08-13 | 50L wave inoculation and perfusion start occur within expected time frames |  | Deviation
295 | Thyrogen 2073 Buffer, CGFBP007, 80 NYA room 102, 50L Carboy #1 out of 5 Overfilled | nicholas.shoenfelt@sanofi.com | 80-101 | 2022-08-14 | 2022-08-14 | It is expected that the scale will be tared prior correctly prior to adding EtOH to the carboy, and that the volume of the buffer will not exceed the carboy during filling.
294 | Validation Audit not Completed per FBL-SOP-000610 | jennifer.lopez@sanofi.com | Digital Operations Technology Dept. (Study Director: Hassan Imran) | 2022-08-12 | 2022-08-15 | Per FBL-SOP-000610 Step 6.5.1: Records Management issues a validation audit twice a year for studies not returned/submitted to RM. Per Step 6.5.12: Validation Audit needs to be completed within 30 business days of audit issuance.
293 | FZ X2200 ATF A Failure | michael.knapp@sanofi.com | 8NYA 1333 | 2022-08-14 | 2022-08-14 | The FZ ATFs are able to run for the duration of the campaign without failure.
292 | Incorrect fill JIT Buffer 2073 | kevin.ornell@sanofi.com | 80 NYA room 102 | 2022-08-14 | 2022-08-14 | Per FBL-MBR-000086 step 4.3.3 expected result is to fill to the required amount of 11.50 kg. The low target is 11.39kg and the maximum target is 11.62kg.
291 | On 16JUL22 recipe on X2321 held due to high high alarm for temperature post media load. | adam.hildebrand@sanofi.com | 74 NYA | 2022-07-16 | 2022-08-14 | It is expected that the recipe would be appropriate for warming this volume of media without exceeding temp limits.
290 | Mixer Failed to Start Post Ammonium Sulfate Addition | mark.smith@sanofi.com | 8NYA/1333 | 2022-08-14 | 2022-08-14 | Mixer should start via the recipe when mix timer begins.
289 | X2210 went into hold for high high DO alarm | adam.hildebrand@sanofi.com | 74 NYA | 2022-08-13 | 2022-08-13 | It is expected that the culture would maintain DO levels within Acceptable range post Inoculation
288 | 450L of Propylene Glycol available. Need to use 135L. Low Alarm Limit is 350L. Need to bypass | Thomas.loik@sanofi.com | 8 NYA 1350 | 2022-08-12 | 2022-08-12 | Low Alarm limit will be bypassed
287 | GEMBA:74NYA: X-3051 pump not working | william.bradley@sanofi.com | 74NYA PU1-09 | 2022-08-12 | 2022-08-12 | Refer to Minncare procedure
286 | QC Raw Material is running tests and the DI water isn’t the PM program | brizaida.medina-vega@sanofi.com | 80 NYA Room 80-124 | 2022-08-12 | 2022-08-12 | TBD
285 | Incorrect Bradford result for 2200036916 validated in iLIMS | john.kapaniris@sanofi.com | 45 New York Avenue Quality Control Chemistry Lab | 2022-08-14 | 2022-08-07 | FBL-MTD-000037 Step 6.9 Acceptance Criteria | Yes | No Deviation
284 | WFI overshot final QS on 760731 HES buffer | aidan.mula@sanofi.com | 8NYA RM1350 | 2022-08-12 | 2022-08-12 | Final volume must be within 1% of target per automation recipe. |  | No Deviation
283 | Multiple Critical Material Inventory Counts not performed for Q1 2022 | Nancy.Waktins@sanofi.com | 68NYA-Lab 2700 | 2022-04-30 | 2022-08-11 | Per FBL-SOP-000182, the Quarterly Inventory counts of Standards, Controls, and Critical Materials must be performed by the end of the month following the Quarter. | Yes | Deviation
282 | Unknown substance noticed in Cz Zinc Skid while performing set up | Michael.Daly@sanofi.com | 8NYA RM1503 | 2022-08-11 | 2022-08-11 | No substances being observed, and skid setup being completed. |  | Deviation
281 | QC Micro Balance settings were found to not match the configuration specification | elizabeth.couturier@sanofi.com | 68 QCM lab | 2022-08-07 | 2022-08-07 | It is expected that the configuration settings are not changed. | Yes | Deviation
280 | 2022-CV-0445 failed TOC swab 2 results during final release Chrom Load Transfer Line to TP-3001 | craig.valery@sanofi.com | 74 NYA, PU1-08 | 2022-07-26 | 2022-08-11 | TOC are expected to not exceed passing criteria
279 | 8NYA-RM 1403 CC8-product impact unknown-FB-4221-016 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 8-1403 | 2022-08-11 | 2022-08-11 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
278 | Extended downtime of AHU7A | john.mendenhall@sanofi.com | 76 NYA | 2022-08-05 | 2022-08-09 | AU would be running
277 | Malfunctioning pCO2 probe caused HiHi alarm on X2321 | david.loik@sanofi.com | 74NYA | 2022-08-08 | 2022-08-08 | The pCO2 probe would not malfunction
276 | The printouts for 45 autoclave is printing the incorrect year | jesse.demers@sanofi.com | 45 NYA RM 169 | 2022-08-10 | 2022-08-10 | Date and time on printout should match the display on the autoclave
275 | CZ SUB X-2220 High Pressure Event 10AUG22 | michael.knapp@sanofi.com | 8NYA 1403 | 2022-08-10 | 2022-08-10 | The vent filters do not wet out.
274 | Ineffective ER FRAD21R0115 for QA Review of Work Orders | tanni.arif@sanofi.com | N/A (INFOR Work Order System) | 2022-08-10 | 2022-08-10 | FBL-SOP-000081
273 | P/N 761222, Lot CGHBY008, Reserve samples 25589 was stored incorrectly. | paul.consolmagno@sanofi.com | QC Chemistry | 2022-04-07 | 2022-08-09 | P/N 761222, Lot CGHBY008, Reserve samples should have been frozen at ≤-60°C
272 | Sparge line was not available in CI for bioreactor Set up 2210 | Bhargav.Patel@sanofi.com | 8 NYA / 1333 | 2022-08-09 | 2022-08-10 | Sparge line should use from CI |  | No Deviation
271 | Repeated Deviance Alarm on pH Probes for Bioreactor D2 | james.letourneux@sanofi.com | 51RM140 | 2022-08-08 | 2022-08-09 | Expected that the probes remain within 0.10 pH units from each other.
270 | FIT Test 3X Failure PN# 5093-01 Buffer CGFB2033 | jeff.hill@sanofi.com | 45 NYA Room 190 | 2022-08-08 | 2022-08-08 | Filter will pass post FIT following Buffer Prep | Yes | No Deviation
269 | 76NYA-Room 107- QCBP-Product impact unknown-F-21403 weight set found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76-107 | 2022-07-06 | 2022-08-09 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
268 | hTSH Low CV Cell Counts on E4 lot CGF4B10 during Growth Phase | brian.mcgee@sanofi.com | 51NYA/Rm 140 | 2022-08-08 | 2022-08-08 | It is expected for the cells to grow at a rate fast enough to enable the density to reach 6e6 c/mL by growth day 8 |  | Deviation
267 | THYROGEN- Unplanned Shutdown 51NYA Purification Rooms | christopher.holmes@sanofi.com | 51 NYA | 2022-08-08 | 2022-08-08 | It is expected that the air handlers and cold rooms would remain operational.
266 | Falcon 1X MES Running Buffer Preparation incorrect per Method | john.kapaniris@sanofi.com | 68 NYA QC Chemistry | 2022-07-28 | 2022-08-05 | 1:20 Dilution of 20X MES | Yes | Deviation
265 | GEMBA: 74NYA FIT Tester does not show a post test result on buffer filter | william.bradley@sanofi.com | 74NYA | 2022-08-07 | 2022-08-08 | Results should show pass of fail on system and verified in PI
264 | HA-CMC - CBYCMC136 - 76NYA Room 229 - AHU2 power blip caused AHU to shutdown. | jan.stepanik@sanofi.com | 76NYA/room 229 | 2022-08-05 | 2022-08-05 | AHU2 should not have gone down.
263 | Logbooks FBL-LBK-000094 EUR (Equipment Use Record) old version 2.0 were being used | dennis.bazinet@sanofi.com | 74NYA all of manufacturing's EURs | 2022-08-07 | 2022-08-07 | Current versions of logbooks are being used
262 | X2331 held due to HiHi media temp | david.loik@sanofi.com | 74NYA | 2022-08-07 | 2022-08-07 | Temperature would stabilize during stabilization period
261 | Missing entries for removal of harvests on BIRFs. | alyson.belleville@sanofi.com | 51 NYA loading dock | 2022-08-05 | 2022-08-07 | It is expected that entries would be completed at time of step execution.
260 | WORK ORDER #2904163. BURST PIPE LINE ON THE SUPPLY AND RETURN LINE ON THE 76-DI LOOP PROCESS ENG. VENDOR NEEDED TO REPAIR PREFERABLY DECCO TO COMPLETE THE JOB ASAP. | MACJONES.OSEIKUFFOUR@SANOFI.COM | BUILDING 76NYA ROOM 140 MEZZINNE | 2022-08-06 | 2022-08-06 | FBL-SOP-000365
259 | Hi pressure alarm Capto MMC column 74 NYA | kpana.jordan@sanofi.com | Building 74NYA, PU1-09 | 2022-08-07 | 2022-08-07 | The pressure should not exceed 30PSI | Yes | Deviation
258 | Leaking LAL for Daily Site | jessica.kychler@sanofi.com | 68 NYA 4th floor room 4700 | 2022-08-06 | 2022-08-06 | Per FBL-SOP-000964 step 6.1.6 cracked or leaking samples should not be received
257 | CZ CGHBN006 MCC X-2811 Global Hold Timer Exceeded | michael.knapp@sanofi.com | 8NYA 1403 | 2022-08-06 | 2022-08-06 | The system does not hold and product is not directed to waste.
256 | CZ CGHBN004 SUB high pressure event | michael.knapp@sanofi.com | 8NYA 1403 | 2022-08-05 | 2022-08-05 | The system does not wet out vent filters and build pressure. Antifoam dosing is sufficient to maintain foaming levels.
255 | Bioburden for CV (2022-51CV-0028) sample found cracked and leaking upon delivery to QC Micro | jacob.kallon@sanofi.com | 68 NYA, QC Micro Lab | 2022-08-05 | 2022-08-05 | It is expected that CV samples be intact when delivered to QC | Yes | No Deviation
254 | Unit of Measure error preventing liquid weigh out | richard.manasian@sanofi.com | 1350 8NYA | 2022-08-05 | 2022-08-05 | MES should work as designed for liquid weigh out. Unit of measure should convert | Yes | Deviation
253 | 8NYA Cold Room 1220 Temperature Excursion/ Cold room temperature went into HiHi Alarm for 2.5 hours on 04AUG22 from 16:15 to 18:30 | stanleydurand@sanofi.com | 8NYA Room 1220 | 2022-08-04 | 2022-08-04 | It is expected for temperature in Cold Room 1220 to maintain temp within set points of 2° to 8°C | Yes | Deviation
252 | Non-viable data for Dynamic operations missing in Onelims system | Adriana.Alban@sanofi.com | 51NYA 152A | 2022-05-16 | 2022-08-02 | It is expected that the Non-viable data is parsed to Onelims | Yes | Deviation
251 | Falcon Buffer record executed using calculations for wrong Batch Size | Ryan.Geraghty@sanofi.com | 74NYA | 2022-07-30 | 2022-08-05 | It is expected that materials weighed out and used in buffer creation would be based on calculations using the correct Batch size.
250 | Sample 327148 EM Hydrocarbon not in respective sampling batch | andres.huerta@sanofi.com | 68 NYA | 2022-06-04 | 2022-08-04 | OneLIMS samples are expected to be in the EM sapling plan, as referenced in the EM plan component part of each unique sample.
249 | 76 NYA Clean Utilities Overdue PMs | wayne.fecteau@sanofi.com | 76-347 | 2022-08-01 | 2022-08-02 | PM would have been completed on or prior to the due date or extended due date as described on FBL-SOP-000081
248 | Buffer totes are out of calibration for Zinc chrop operation | Pavel.dobrovolskiy-wynes@sanofi.com | CNC space suite 2 | 2022-07-24 | 2022-08-03 | Its expected that buffer totes will remian within calibration during proccessing
247 | CZ- High SUB pH on X2220 | kelvin.leo@sanofi.com | 8 NYA | 2022-08-03 | 2022-08-04 | No hold |  | No Deviation
246 | 76-DI PROCESS ENGINEERING LOOP DISTRIBUTION PUMP FAULT/UV OFF. | MACJONES.OSEIKUFFOUR@SANOFI.COM | BUILDING 76NYA ROOM 140 | 2022-08-03 | 2022-08-03 | FBL-SOP-000365
245 | Spider was found in the room 140 next to D1 Bioreactor at 51NYA | daniel.bati@sanofi.com | 51NYA/ 140 CC | 2022-08-03 | 2022-08-03 | Scheduled a Gemba
244 | High Conductivity Alarm on RO Supply and Return loop at 74 NYA | brizaida.medina-vega@sanofi.com | 74 NYA | 2022-08-02 | 2022-08-02 | Less that 1.3 micro-siemens
243 | Leak detected on V2071 post media load | duane.messier@sanofi.com | 74NYA/CC1-07 | 2022-08-03 | 2022-08-03 | Vessel should not have had a leak post PHT
242 | MCC Skid 2811 Failed bubble trap top sensor | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2022-08-03 | 2022-08-03 | Top sensor should reading liquid when it fill up
241 | Tier 1 Action Notice unable to be located | timothy.pyne@sanofi.com | 70 NYA | 2021-12-15 | 2022-07-29 | it is expected that all action notices are reconciled with records management
240 | THYROGEN- Overdue PM - Monthly QCA Review WO (5W2H) | christopher.holmes@sanofi.com | 51 NYA | 2022-08-01 | 2022-08-02 | It is expected that the PM would have been completed by its 31JUL due date |  | Deviation
239 | X2210 pCO2 Deviation alarm | kurt.marsocci@sanofi.com | 74 NYA. CC2-01 | 2022-08-03 | 2022-08-03 | pCO2 probes are expected to stay within 7 slpm of each other.
238 | Worker entered Grade D space improperly gowned (5W2H) | michael.knapp@sanofi.com | 8NYA 1403 | 2022-08-02 | 2022-08-02 | All entrants into the Grade D space are properly gowned per FBL-EWI-000004 |  | Deviation
237 | Worker entered Grade D space improperly gowned | michael.knapp@sanofi.com | 8NYA 1403 | 2022-08-02 | 2022-08-02 | All entrants into the Grade D space are properly gowned per FBL-EWI-000004
236 | Incorrect Values pulled into Batch Record | Timothy.Ryley@sanofi.com | 8NYA Room 1333/1334 | 2022-07-01 | 2022-08-02 | EBR should pull most recent data
235 | X-2210 in hold due to DO alarm following inoculation | david.loik@sanofi.com | 74NYA | 2022-08-02 | 2022-08-02 | DO would not reach the HiHi before stabilizing
234 | 51 NYA- Thyrogen- D1 Bioreactor(CGF1S09) pO2 out of range on Transition Day 5 | ian.odell@sanofi.com | 51 NYA Room 140 | 2022-07-30 | 2022-07-30 | Per FBL-CST-000007, the pO2 range is 30-131mmHg for Transition.
233 | EM Non-viable sample data parsing issue | Adriana.Alban@sanofi.com | 152A | 2022-05-15 | 2022-08-02 | Climet Non viable sample data should be parsed to Onelims after each operation.
232 | During the column rinse phase of the Capto MMC integrity check, a leak was observed at rupture disc PSE-305085. | craig.valery@sanofi.com | 74 NYA, PU1-09 | 2022-08-02 | 2022-08-02 | It is expected that no leaks are observed
231 | HA Components (AUT) - FBL-MBR-000078; Lot# CBYAUT041; Cycle parameter different from SOP | charlie.houle@sanofi.com | 76 NYA, room 352 | 2022-08-01 | 2022-08-01 | Parameters for cycle should be identical in SOP and autoclave printout
230 | 51 NYA- Thyrogen- E4 (CGFSB010) duration from 24L sample to bioreactor inoculation exceeded 2 hours. | ian.odell@sanofi.com | 51 NYA Room 140 | 2022-08-01 | 2022-08-01 | Per FBL-MBR-000108 and BQ-CST-000024, duration between BGA sample of 24L and bioreactor inoculation must be less than or equal to 7.5 hours. This is a Consistency Indicator | Yes | Deviation
229 | 68NYA-Metrology-no product impact-Metrology Standard S-238 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2022-07-22 | 2022-08-02 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
228 | OOS temperature recorded for ingoing equilibration buffer sample during Q process | jacob.kallon@sanofi.com | 51 NYA | 2022-07-28 | 2022-08-01 | Per FBL-MBR-000111 step 4.4.4.4, the temperature range is 21-23 degrees C
227 | Gemba RQ - F076-SWECO-X3571-76NYA-CQ-0007, Process Sim for Pur of Sterile Bulk HA for SWECO Vsl Sys FBS-MBR-000053 (Total Sterile Hold Time Not Met) | john.sicurella@sanofi.com | 76NYA / | 2022-04-25 | 2022-08-01 | X-3571 Total Sterile Hold Time (X-3571) should have been >=120 Hours
226 | Broken thermometer | kathryn.lentoni@sanofi.com | 68NYA 4700, QCM | 2022-07-24 | 2022-07-24 | Thermometer not broken and functions properly
225 | Procedure not followed during filter integrity test | jacob.kallon@sanofi.com | 51 NYA | 2022-03-09 | 2022-08-01 | Per FBL-SOP-001011, 3 consecutive FIT failures require a deviation
224 | Bioburden Settle Plate sampling Start/End Times not recorded on worksheet | Arpan.christian2@sanofi.com | 68 NYA, QC Microbiology, 4700, BSC F-16088 | 2022-07-15 | 2022-08-01 | Per FBL-SOP-000380, Step 6.2.4 the Sampling Start/End times should be recorded during assay occasion.
223 | 8-STM-7 not in the routine monitoring program | brendan.scanlon@sanofi.com | 8 NYA 1403 | 2022-08-01 | 2022-08-01 | All routine monitoring sites utilized for GMP processing should be added to the routine monitoring program. | Yes | Deviation
222 | Gemba Doc #: RQ - F051-COL-X3350-CQ-0002, RQ Physical Testing for CIP Sensor Array and Valve Circuit X-3350 (51-138) was not executed by the month due of 31JUL22 due to Run #1 Conductivity Failure | john.sicurella@sanofi.com | 51NYA/Room 138 | 2022-08-01 | 2022-08-01 | FBL-VMP-000006, Step #: 6.3.10.1, "...RQ physical test (24 month interval) will be by the end of the month..."
221 | ATF Exchange Rate below CST PAR prior to planned ATF Changeout | michael.knapp@sanofi.com | 8NYA 1403 | 2022-07-31 | 2022-08-01 | ATF exchange rate remains within PAR prior to ATF changeout.
220 | 51NYA-Room 130-APT DS-Product impact unknown-FB-1842 overdue for calibration | robert.bauch@sanofi.com | 51-130 | 2022-08-01 | 2022-08-01 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been able to be collected
219 | CZ SUB CGHBN006 Daily Sampling Probe Adjustments not performed | michael.knapp@sanofi.com | 8NYA 1403 | 2022-07-31 | 2022-08-01 | Sample activities for the day are completed.
218 | CZ CGHBN004 Harvest SUM High Temperature | michael.knapp@sanofi.com | 8NYA 1403 | 2022-07-29 | 2022-08-01 | The Harvest SUM maintains temperature |  | Deviation
217 | Hepa Vacuum PM's not completed on time | patrick.hosegood@sanofi.com | 8 NYA | 2022-08-01 | 2022-08-01 | Parts are available and PM completed
216 | Gemba for double tested LAL water | roodline.bertrand@sanofi.com | 68NYA,QC MICRO LAB | 2022-07-28 | 2022-07-31 | Per FBL-SOP-001050 the sample needs to have one valid data | Yes | Deviation
215 | Thyrogen - CGF2B02 - Cell Culture CGF2B02 - Potable Media and Usage Log Storage and Removal times don't agree. | Edward.kozatek@sanofi.com | 51NYA RM 109 | 2022-03-03 | 2022-07-31 | Storage and removal should follow chronological order.
214 | hTSH Low offline p02 on T5 of CGF1S09 at 51NYA | brian.mcgee@sanofi.com | 51NYA/Rm 140 | 2022-07-30 | 2022-07-30 | It is expected for the p02 to be within allowable range. Range is listed in FBL-MBR-000091, step 4.19.10
213 | Six Thyrogen 51NYA Cell Culture BSB FADs (FAD16-071 PR-107-19) Batch Records not located. Lot #s 6GFSV001, 6GFSV002, 6GFSV003, 6GFSX001, 6GFSX002, and 6GFSX003 | jennifer.lopez@sanofi.com | Records Management (80NYA and Iron Mountain) | 2022-07-25 | 2022-07-25 | Per FBL-SOP-000681 (Batch Production Record Issuance and Control for Framingham Biologics and Biosurgery, version 9.0), it is expected that all completed/reviewed batch records are filed in controlled designated RM locations until archived off-site per FBL-SOP-000204 (Records Management Archiving at Framingham Biologics, version 7.0).
212 | Growth on Negative Control for Bioburden | jessica.kychler@sanofi.com | 68 NYA QC Micro, 4700 | 2022-07-23 | 2022-07-29 | It is expected that there wuld be no growth on the negative control | No | No Deviation
211 | Action Notice 2022-I-040: "WO 2899619 Replace Gaskets (Change Out)" was released but the gaskets were never changed out | richard.quitadamo@sanofi.com | 08-1403 | 2022-07-22 | 2022-07-28 | Action notice 2022-I-040 would not be closed/released until WO 2899619 was completed, and that the work defined within the AN would be completed as described | Yes
210 | Five new climets were purchased as part of Phase II CerezymeICB project, were onboarded to Gage Insite for routine calibration but have not gone through proper commissioning and qualification. | rondik.ahmed@sanofi.com | 8 NYA | 2019-09-30 | 2022-06-23 | Equipment will be qualified per FBL-SOP-0001460
209 | 51 NYA - Thyrogen- Wave pCO2 control not active during 1L media equilibration | ian.odell@sanofi.com | 51 NYA Room 152 | 2022-07-28 | 2022-07-28 | Per FBL-RPT-001386 pH should be 6.6-7.2
208 | 51 NYA- Thyrogen- Bioreactor D1 pCO2 offline result out of range on Transition Day 3 (CGF1S09) | ian.odell@sanofi.com | 51 NYA Room 140 | 2022-07-28 | 2022-07-28 | Per FBL-MBR-00091 step 4.17.7, this must be 100-230mmHg | Yes | Deviation
207 | MCC 2811 Bubble Trap low level alarm | michael.knapp@sanofi.com | 8NYA 1403 | 2022-07-28 | 2022-07-28 | The bubble trap fills and does not alarm for low level following buffer prime
206 | FZ Low Viability on Vicell Sample Day 34 | michael.knapp@sanofi.com | 8NYA 1333 | 2022-07-28 | 2022-07-28 | Viability remains within range for the duration of the FZ campaign. |  | Deviation
205 | 68NYA-room 4700 QCM-product impact unknown-FB-2426 found out of tolerance for conductivity during a scheduled calibration | robert.bauch@sanofi.com | 68-4700 | 2022-07-25 | 2022-07-28 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance | Yes | Deviation
204 | MCC X-2811 Column 3 Leak at Bleed Bag During Blank Run | michael.knapp@sanofi.com | 8NYA 1403 | 2022-07-28 | 2022-07-28 | The column is attached during the Buffer Prime activities and at the completion of buffer prime the column pathways are opened and the bleed bag is removed.
203 | 51 NYA- Thyrogen- pCO2 offline result out of range at 12L wave stage (CGFSB010) | ian.odell@sanofi.com | 51 NYA Room 152 | 2022-07-27 | 2022-07-27 | In FBL-RPT-001386, this is a CI. | Yes | Deviation
202 | Doc #: RQ - F068-INC-FB0487-CQ-0001, RQ Physical Testing for Environmental Shaker 40-QC-0487 was not executed by the month due of 30JUN22 | john.sicurella@sanofi.com | 68NYA / QCTS Lab | 2022-07-01 | 2022-07-01 | FBL-VMP-000006, Step #: 6.3.10.1, "...RQ physical test (24 month interval) will be by the end of the month..."
201 | X2301 hold for HiHi temperature after media load | david.loik@sanofi.com | 74NYA | 2022-07-27 | 2022-07-27 | FBL-CST-000030
200 | TOC swab 1 for 2022-CV-0445 has a failing result of 3.02ppm. | craig.valery@sanofi.com | 74 NYA | 2022-07-27 | 2022-07-27 | TOC are expected to not exceed passing criteria
199 | CZ Bioreactor X2250 ATF A - Multiple Guard Filters Replaced due to ATF A Breach | michael.knapp@sanofi.com | 8NYA 1403 | 2022-07-27 | 2022-07-27 | Unplanned ATF changeout is not required and there is no breach into the harvest line. |  | Deviation
198 | QSTP Alert | Kelly.Spenard@sanofi.com | Third Party Supplier | 2022-07-26 | 2022-07-27 | No Warning Letter
197 | Cerezyme – Lot CGHBN006 – 8 NYA – During Buffer Prime the Bubble Trap liquid level exceeded the high level sensor and wet the vent filter on MCC Skid X2811 | sheila.pandolfo@sanofi.com | 8 NYA, Room 1403 | 2022-07-26 | 2022-07-26 | It is expected that the bubble trap sensor would not malfunction and the sensor would recognize there is liquid when it is present. It is expected the vent filter would not get wetted out with equilibration buffer during the priming operation.
196 | High Conductivity results in the 76 WFI R&D tank and loop | brizaida.medina-vega@sanofi.com | 76 NYA WFI R&D Tank and loop | 2022-07-24 | 2022-07-24 | Conductivity into the validated range
195 | Capto MMC Recipe went into HOLD due to Hi Hi pressure | alvaro.rueda@sanofi.com | 74 NYA | 2022-07-26 | 2022-07-26 | HI HI pressure is not exceeded
194 | Missed Entry for BSC Usage logbook Water bioburden | Pushti.Shah@sanofi.com | 68 NYA/ QC Micro | 2022-07-25 | 2022-07-26 | 6.5.1 - Perform pre-use BSC disinfection with Spor-klenz RTU daily prior to first use. Prior to each additional use for the day, perform pre-use sanitization with 70% IPA. | Yes | Deviation
193 | SUT Failure - SUM Bag Leak Around pH Probe Port | mark.smith@sanofi.com | 8NYA/1503 | 2022-07-26 | 2022-07-26 | Product to stay contained within bag
192 | Thyrogen - CGF2B02A - GEMBA TSK Column Operation CGF2B02A - Recorded Temps don't align with PI/Temperature out of specification during Elution Steps/ | Edward.kozatek@sanofi.com | 51 NYA RM132 | 2022-06-21 | 2022-07-25 | FBL-SOP-000788 Attachment 5 Temperature remains at 6 - 10 degrees Celsius for all steps except storage.
191 | On 25July2022 at 2225 X2230 was alarming due to low temp (36.7) for TIC-2230018 After some troubleshootings it was discovered that the wire connection to the heater for the filter was loose, after tightning it the TIC-2230018 spiked to 65 degrees, as a result X2230 went into hold for high temp. Shortly after the spike, the temp came back down to range, X2230 was re-started and has been running fine since | adam.hildebrand@sanofi.com | 74 NYA | 2022-07-25 | 2022-07-25 | It is expected that the vent filter assembly heater would work normally without interruption
190 | Product TBD - Lot TBD - 8 NYA - Employee Tested Positive for COVID-19 who entered the 8 NYA GMP Manufacturing Area | sheila.pandolfo@sanofi.com | 8 NYA Room 1323 & 1324 | 2022-07-25 | 2022-07-25 | It is expected that COVID-19 positive personnel do not enter controlled/classified areas.
189 | Max Allowable Low pH Timer Exceeded | mark.smith@sanofi.com | 8NYA/1323 Downstream Ballroom | 2022-07-25 | 2022-07-25 | 1. Max allowable low pH timer should not exceed 210 minutes. 2. SUM bag SUT should come with appropriate SUT to allow welding of sample bags.
188 | 45NYA WFI still generation system is not running due to CONTACTOR damage | brizaida.medina-vega@sanofi.com | 45NYA Manufacturing & GLP Distribution | 2022-07-24 | 2022-07-24 | Equipment in service
187 | Operators not trained on current version of FBL-SOP-000431 | alvaro.rueda@sanofi.com | 74 NYA | 2021-11-02 | 2022-07-25 | SOP should have been assigned to a curricular
186 | GZ7084-01 (Buffer bag substitute) does not come with the required connection per FBL-SOP-000504 | Jeremy.Wooden@Sanofi.com | 80.9902 | 2022-07-25 | 2022-07-25 | The bag utilized for the process would have a female quick connect attached by the manufacturer. | Yes | No Deviation
185 | CV provisionally released with a preliminary bioburden read of 1 CFU. | megan.botti-marino@sanofi.com | 74 NYA | 2022-07-24 | 2022-07-25 | It is expected that the associated equipment should not have been provisionally released and a new CV issued.
184 | Cracked PI Elisa Vials | joseph.napoli@sanofi.com | 68NYA / QCM | 2022-07-22 | 2022-07-25 | vial to remain intact with no cracks
183 | Temp Probe Failure During SP Pooling | mark.smith@sanofi.com | 8NYA/1323 Downstream Ballroom | 2022-07-25 | 2022-07-25 | Temperature display should match product temperature | Yes | Deviation
182 | Stability: Prediction Interval Assessments for Intrinsic Viscosity EP and Molecular Weight EP for Lot 9BYHAH012 at 24 Months is OOS | Craig.Stevenson@sanofi.com | 76 New York Ave, Room 107, QCBP | 2022-07-22 | 2022-07-22 | It is expected that the Prediction Interval Assessments for Intrinsic Viscosity EP and Molecular Weight EP for Lot 9BYHAH012 at 24 Months generates passing results | Yes | No Deviation
181 | Supplier deviation for 692169 3K SUM bag | richard.manasian@sanofi.com | 1350 | 2022-07-24 | 2022-07-24 | SUM bag p/n 692169 should have a pressure indicator line that can be connected to bag with appropriate fittings
180 | Pressure lost and leak observed during re-SIP of Cz CGHSN007 MCC Load Line during cooldown of Column 2 | john.lytle@sanofi.com | Rm 1434/ 8NYA | 2022-07-24 | 2022-07-24 | FBL-EWI-000032
179 | Failed post filter FIT test 3 times on Large Scale PSFM batch but passed after the SIP | dennis.bazinet@sanofi.com | 74NYA, CC1-07 | 2022-07-23 | 2022-07-23 | FBL-MBR-000162 step 9.6.21 FIT results are out of spec per FBL-SOP-000866
178 | Leak found during media load of the 500L SUB at 74NYA | david.loik@sanofi.com | 74NYA | 2022-07-23 | 2022-07-23 | The SUT would not leak. | No | Deviation
177 | Fz Shake Flask Low Growth- Follow up | john.lytle@sanofi.com | 8NYA Upstream Seed Prep Rm# | 2022-07-23 | 2022-07-23 | Viability would be >= 80%
176 | HairFound on Bioburden Filter upon Opening | jessica.kychler@sanofi.com | 68 NYA 4th floor room 4700 | 2022-07-22 | 2022-07-22 | Filters are expected to be sterile and free of contamination | Yes | Deviation
175 | Leak discovered on SIP'd CZ MCC Skid | michael.knapp@sanofi.com | 8NYA 1403 | 2022-07-22 | 2022-07-22 | There are no leaks
174 | FZ Shake Flask low growth | michael.knapp@sanofi.com | 8NYA 1273 | 2022-07-22 | 2022-07-22 | Growth is seen in the shake flask and scale up is possible
173 | Seprafilm - CBYSEP044- 76 NYA rooms 233. COVID positive employee in Seprafilm inspection | fatima.fragaalvarez@sanofi.com | 233 and mezzanine level | 2022-07-21 | 2022-07-22 | It is expected that personnel would not enter classified manufacturing space when they have an illness per FBS-SOP-000667
172 | Isolator-76NYA 3104 room Gowning procedure not followed | valeria.moreno@sanofi.com | 3104 | 2022-07-21 | 2022-07-21 | Follow FBS-SOP-000667
171 | HA-76NYA room 3102-Out of range pressure in V-3532 during Sterile filtration | valeria.moreno@sanofi.com | 3102 | 2022-02-09 | 2022-07-21 | FBS-MBR-000056, step 7.8.28 page 21/66, pressure in V-3532 must be between 5 - 10 psi | Yes | Deviation
170 | SUT leak | Thomas.loik@sanofi.com | 8 NYA 1350 | 2022-07-21 | 2022-07-21 | Observation recorded, Pre filtration, No impact to product
169 | No reportable results for 76-407-29 (76 R&D) sampled on 11May2022 | marissa.lentoni@sanofi.com | 76 NYA, Site 76-407-29, Room 29, Grade D (Monthly Sampling) | 2022-05-11 | 2022-05-26 | Counts would be recorded for the site.
168 | NBO raw material 328689 NBO_raw_GZ1335-03 tested without batch; potential non-reportable | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2022-04-02 | 2022-07-19 | OneLIMS samples are expected to be tested in biolims. Per BCP, if sample cannot be tested, sample should be tested on paper.
167 | Falcon: Virus wave in hold DO dropped to zero then 73.1 | nicole.lufkin@sanofi.com | 74 NYA CC1-07 | 2022-07-20 | 2022-07-21 | FBL-CST-000030
166 | 68NYA-Metrology-no product impact-Metrology Standard S-1636 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2022-07-14 | 2022-07-21 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been able to be collected
165 | 70% IPA wipe Lot# and Expiry Not Documented on Water Bath Maintenance Form and No Dry State Storage Logbook for Equipment | elizabeth.couturier@sanofi.com | 68 QCM lab | 2022-06-29 | 2022-07-20 | It is expected per ALF-SOP-000020 that the maintenance logbook is fully filled out and that dry state maintenance is recorded in logbook ALF-FRM-000139. | No | Deviation
164 | CZ Batch record not generating associated PCS alarms | michael.knapp@sanofi.com | 8NYA 1403 | 2022-07-19 | 2022-07-20 | Quality Critical alarms generated by the associated recipe are captured as PCS alarms in MES>
163 | Spider Found in the Lady Airlock at 51NYA | daniel.bati@sanofi.com | 51NYA/ Airlock | 2022-07-20 | 2022-07-20 | Scheduled a Gemba |  | No Deviation
162 | Multiple Products – Lots TBD – Multiple Buildings – Missing entries in Credo Cooler Sample Pick-up Log, FBL-LBK-000188 Attachment 1 | sheila.pandolfo@sanofi.com | TBD, Samples are transported from designated lay down areas and Quality Control (QC) testing laboratories at Framingham MA | 2022-02-09 | 2022-07-17 | FBL-SOP-000942 | No | Deviation
161 | Retain samples and NRSR sample stored in 108c | Adriana.Alban@sanofi.com | 51NYA 108c | 2022-03-05 | 2022-07-19 | Retain samples are to be delivered to QC within 24 hours or discarded
160 | Unable to sample compressed air site 76-CA-43 due to presence of water in the valve | marissa.lentoni@sanofi.com | 76 NYA, 76-CA-43, Room 347 | 2022-07-20 | 2022-07-20 | Water vapor would not be present in the compressed air site and sampling would be able to occur. | Yes | Deviation
159 | “FBS-MBR-000058, Lot BBYFRM050 steps 7.6.2.6, 7.6.3.6, and 7.6.4.3 was completed by a trainee. Trainer countersigned and also performed verification step.” | dena.elmesalamy@sanofi.com | 76NYA-364 | 2021-10-08 | 2022-06-07 | Performed verification should be a different operator than the trainer countersigned operator.
158 | “FBS-MBR-000058, Lot BBYFRM060 steps 7.6.2.6, 7.6.3.6, and 7.6.4.3 was completed by a trainee. Trainer countersigned and also performed verification step.” | dena.elmesalamy@sanofi.com | 76NYA-364 | 2021-10-08 | 2022-06-07 | Performed verification should be a different operator than the trainer countersigned operator.
157 | “FBS-MBR-000058, Lot BBYFRM060 steps 7.6.2.6, 7.6.3.6, and 7.6.4.3 was completed by a trainee. Trainer countersigned and also performed verification step.” | dena.elmesalamy@sanofi.com | 76NYA-364 | 2021-10-08 | 2022-06-07 | Performed verification should be a different operator than the trainer countersigned operator.
156 | 51 NYA, Thyrogen - Water observed underneath Bioreators D1 and D2, Downstream from reactor | sheila.pandolfo@sanofi.com | 51 NYA, Room 140 | 2022-07-19 | 2022-07-19 | It is expected there would not be a leak at this connection. | Yes | No Deviation
155 | 76NYA-Room 233-Biosurgery packaging area-no product impact-2 balances found out of tolerance | robert.bauch@sanofi.com | 76-233 | 2022-07-19 | 2022-07-20 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
154 | CZ MCC CV Rinse sample leaked and resample performed | michael.knapp@sanofi.com | 8NYA 1403 | 2022-07-19 | 2022-07-19 | Rinse samples are pulled at the first opportunity and there are no leaks | Yes | No Deviation
153 | CZ SUB weight Low Low following perfusion start | michael.knapp@sanofi.com | 8NYA 1403 | 2022-07-19 | 2022-07-19 | Media turns on |  | No Deviation
152 | 51NYA Myco Sample disocvered leaking upon receipt into QCM lab (L/N: CGF1S09, G-0). | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2022-07-19 | 2022-07-19 | It is expected that samples delivered to QCM are not leaking into the secondary bag. | Yes | Deviation
151 | 1 Expired Timer used for a manual glass wash 2 No timer report on work sheet | Robin.Loguidice@sanofi.com | 68 NYA first floor glass wash room | 2022-05-06 | 2022-07-19 | To use a in calibration timer and to record the timer at time of use | Yes | Deviation
150 | 51 NYA- Thyrogen- pO2 offline result out of range on 8L spinner (CGFSS009) | ian.odell@sanofi.com | 51 NYA Room 152 | 2022-07-18 | 2022-07-19 | FBL-CST-000005 per table 8, this is a CI.
149 | 76NYA-Room 347-WFI Still 2 Distillate Temp- product impact unknown-F-8740-006 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 76NYA room 347 | 2022-07-12 | 2022-07-19 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been in tolerance
148 | CZ SUB Rate of Change alarm | michael.knapp@sanofi.com | 8NYA 1403 | 2022-07-18 | 2022-07-19 | SUB media switchover is successful and there is no rate of change alarm
147 | Dual Conductivity testing performed for WFI site: 8-WFI1-2 (8NYA, Room 1170), generating two separate valid results | tae.kang@sanofi.com | 68NYA/4700 | 2022-06-24 | 2022-06-29 | Dual testing should not have been performed. | Yes | Deviation
146 | During deltaV server reset, the wave systems lost communication between unicorn and deltav. The rocker stopped for all waves, but only 2126 had cells, the stoppage lasted ~10min. Also, all weights were cleared from deltaV, all wave now read 0 | adam.rowe@sanofi.com | 74NYA, CC2-01 | 2022-07-18 | 2022-07-18 | The planned reset should not have caused any issues, we should still have the weight reading for all waves
145 | CZ ATF SUT failure during wetting | michael.knapp@sanofi.com | 8NYA 1403 | 2022-07-18 | 2022-07-18 | There is no leak along the ATF bulb
144 | Covid Positive Event 8 NYA Downstream | Valerie.staras@sanofi.com | 8 NYA Lab Suite 2 | 2022-07-14 | 2022-07-18 | It is expected that COVID-19 positive personnel do not enter controlled/classified areas.
143 | Deformed o ring found on TFF liner when installing conductivity probe AI01 | craig.valery@sanofi.com | 74 NYA PU1-24 | 2022-07-18 | 2022-07-18 | manufacturing defect, it is expected that no defects are present.
142 | Minimum Weight check scale reading missing FBL-LBK-000340, FB-3660 logbook 2021LBC0294 | William.Lee@sanofi.com | 74 NYA Media CC1-08 | 2022-06-17 | 2022-07-17 | FBL-SOP-000786 section 6.1 A two point balance weight check is required. Step 6.1.10 Record the results of the first point balance check on the weight check log sheet.
141 | Buffer Bag Leak (PN# GZ7084-01) from Inlet Tubing | jeff.hill@sanofi.com | 45 NYA Room 190 | 2022-07-16 | 2022-07-16 | Buffer Bag will fill without leaking.
140 | Incorrect Column used to perform the rinse and equilibration steps of the SP2 process | sylvester.kegbeh@sanofi.com | 51 NYA ROOM 132 | 2022-07-16 | 2022-07-16 | It is expected that the SP2 column be used during the SP2 process.
139 | Pall filter failure. Filter type 10” 0.2micron filter (NBO-SPC-000648) P/N 6574-01, Batch CGFB1243, Expiry 29FEB2024 | jeff.hill@sanofi.com | 45 NYA Room 190 | 2022-07-15 | 2022-07-15 | It is expected that a test would Pass post FIT at step 6.4.2.6 of FBL-SOP-001011
138 | GEMBA: 74NYA - High Pressure alarm during Sanitization on Capto MMC | william.bradley@sanofi.com | 74NYA PU1-09 | 2022-07-15 | 2022-07-15 | Isolate column, remove filter, prime, have automation reset the flow direction and continue sanitization
137 | Bulk Identity sample not received into QC. | timothy.pyne@sanofi.com | 74 NYA, PU1-24 | 2022-06-27 | 2022-07-15 | It is expected that the sample would have been received into QC and tested. | Yes | No Deviation
136 | GEMBA: 74NYA - DSP, Hi alarm on Capto | william.bradley@sanofi.com | 74NYA | 2022-07-15 | 2022-07-15 | Shouldn't reach hi alarm
135 | Action notice 2022-R-018 released prior to having a passing conductivity result | timothy.hoard@sanofi.com | 76NYA, WFI STILL 2 FINN-AQUA, critical utilities. | 2022-07-14 | 2022-07-14 | Critical utilities should not be released until proper test samples and results are confirmed. WFI with restrictions under action notice should not be used until released by QA.
134 | HA - 76NYA Room 364 - Failed HEPA Certification F-0704 | charlie.houle@sanofi.com | 76 NYA, room 364 | 2022-07-14 | 2022-07-14 | LFH would meet acceptance criteria per FBL-SOP-000156.
133 | X-2210 in hold due to pCO2 deviation alarm (F-074-AI-2210028) | david.loik@sanofi.com | 74NYA | 2022-07-14 | 2022-07-14 | Recipe would not go into hold. | Yes | No Deviation
132 | POU sampled under incorrect Biolims sample IDs | Lolav.ahmed@sanofi.com | 8NYA Upstream | 2022-07-12 | 2022-07-13 | FBL-RPT-000538 Table 7
131 | V-2641 surpassed clean hold expiry and is required for operations this evening | craig.valery@sanofi.com | 74 NYA, CC1-03 | 2022-07-13 | 2022-07-13 | Vessel is to be in the clean status before operation
130 | MCC Eluate PBM SUM 2570002 Temperature below 2 degree. | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2022-07-13 | 2022-07-13 | per FBL-CST-012345 the temperature must be between 2-8C. | Yes | Deviation
129 | MCC held due to Column 3 pH fluctuations during equilibration. As a result, product remained on Column 1 longer than the 4hour total allowed, which resulted in product loss | john.lytle@sanofi.com | Rm 1434 | 2022-07-13 | 2022-07-13 | no hold during MCC column equilibration operations |  | No Deviation
128 | TSH - Missing Bioburden Sample from NRSR 22SRF013 | Pamela.Lessard@sanofi.com | 68/4720 Sample Receipt office/Microbiology Lab | 2022-07-03 | 2022-07-12 | Bioburden results would be obtained | Yes | Deviation
127 | Parts washer at 51 NYA failed conductivity as founds during routine calibration | nicholas.shoenfelt@sanofi.com | 51NYA / room 142 | 2022-07-12 | 2022-07-12 | It is expected that as found will pass calibration for conductivity probes. | No | Deviation
126 | Transmitter display for DO probe on the control panel turned off without any operator intervention SUB (Single Use Bioreactor) X-2250 | sheila.pandolfo@sanofi.com | 8 NYA, Room 1403 (Suite 2) | 2022-07-12 | 2022-07-12 | It is expected that it would not turn off during operations. | Yes | No Deviation
125 | Cerezyme and/or Fabrazyme - Lot TBD - 8 NYA - Employee Tested Positive for COVID-10 who entered the 8 NYA GMP Manufacturing Area | sheila.pandolfo@sanofi.com | 8 NYA, Rooms 1333, 1403 & 1340 | 2022-07-10 | 2022-07-12 | It is expected that COVID-19 positive personnel do not enter controlled/classified areas. | Yes | Deviation
124 | No reportable result for weekly EM Site 801-2 RODAC B (423135). | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2022-07-07 | 2022-07-07 | It is expected that plates are not broken upon plate counting, which would allow for accurate weekly EM results. | Yes | Deviation
123 | Incorrect total mix time of Blue Eluate (Load adjust) | edgar.vale@sanofi.com | 51NYA, room 132 | 2022-07-04 | 2022-07-13 | It is expected that the total mix time would be between 10 to 14 minutes
122 | Seprafilm-Room 1017 74 NYA Chart recorder die not record from 01JUL22 to 05JUL22 | alan.rousseau@sanofi.com | Room 1017 74 NYA Labeling and Packaging | 2022-07-01 | 2022-07-05 | Chart would record data | Yes | No Deviation
121 | GEMZ: 74NYA wrong recipe loaded for FBL-MBR-000146 | william.bradley@sanofi.com | 74NYA PU1-22 | 2022-07-12 | 2022-07-12 | Step 8.3.7 " Perform loading and running recipe for Extended Dry Mix - Ambient Buffer Prep Vessel Formulation per FBL-SOP-000537"
120 | GEMBA: 74NYA Concern with not following gowning procedure | william.bradley@sanofi.com | 74NYA suits ( CC, PU1-xx) | 2022-07-11 | 2022-07-11 | Jewelry is to be left outside of the lab.
119 | Approved version of NRSR issued prior to effective date | Ryan.Geraghty@sanofi.com | 51NYA, Room 138 | 2022-07-02 | 2022-07-12 | Per BQ-GOP-000181, Step 5.2 it is expected that the effective version of a document will be executed.
118 | Versions 1.0 and 2.0 of FBL-FRM-000738 became effective with incorrect versioning in revision history table | Amy.McIntyre@sanofi.com | In the GEODE+ documentation system | 2019-06-07 | 2022-07-11 | In version 8.0 of BQ-GOP-000181, Attachment 1 and Attachment 5 (Checklists for Change Owner and for Document Control), both the change owner (workflow initiator) and document control (in the reviewer stage) must
117 | 100L SUB (X2210) held due to HI HI alarm on DO nKPP | adam.hildebrand@sanofi.com | 74 NYA | 2022-07-11 | 2022-07-11 | It is expected that the DO would maintain setpoint during process | Yes | No Deviation
116 | Nanofiltration chase pump local read-out did not match DV reading - Caused SUM weight to drop out of range | mark.smith@sanofi.com | 8NYA/1323 Downstream Ballroom | 2022-07-11 | 2022-07-11 | Pump readout should match between DeltaV and local readout. This ensures SUM weight level is consistant. |  | Deviation
115 | Thyrogen - Improper gowning at 51NYA Controlled Areas | eliana.mendenhall@sanofi.com | 51NYA, rooms 129, 143 and140, corridor 84 | 2022-07-11 | 2022-07-11 | It is expected to wear a mask while entering the area per FBL-SOP-000864 step 6.3.1 | Yes | Deviation
114 | Missed count of 1 CFU on 2200035192, Viable Air, Site 226, Room 364, 76 NYA Zone 2, 18Jun2022, Action Level | paul.consolmagno@sanofi.com | F-16387, 76NYA; QCM, 68 NYA | 2022-06-18 | 2022-07-08 | Per FBL-SOP-0001496, step 6.2.19 the plate would have been saved until the results were evaluated and per step 6.4.1, the Action level result for the viable air would have been submitted for identification. | Yes | Deviation
113 | 51 NYA- Thyrogen- Colder apparatus leak during media refeed of D2 bioreactor (CGF2B08) | ian.odell@sanofi.com | 51 NYA Room 140 | 2022-07-10 | 2022-07-10 | Per FBL-MBR-000091, no leaks should be detected during the refeed process. |  | Deviation
112 | CZ SUB %DO loss for short period of time | michael.knapp@sanofi.com | 8NYA 1403 | 2022-07-08 | 2022-07-08 | Per FBL-CST-000029 Table 7.8 Non-KPP %DO maintains at 15% to 50%. Excursions of </= 12 hrs are allowed. This excursion was for 20 minutes total. |  | No Deviation
111 | 51 NYA - Thyrogen- Incorrect version on FBL-MBR-000091 used for Bioreactor run (CGF2B02) | ian.odell@sanofi.com | 51 NYA Room 140 | 2022-07-24 | 2022-07-08 | Current effective version should be verified before use and used for production activities |  | No Deviation
110 | Non-reportable for Falcon Endotoxin 402913 never received in QCM | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2022-06-09 | 2022-07-08 | Samples placed in laydown areas should be transported to QCM and received by Sample Receipt. | Yes | Deviation
109 | Sepraflim sterile lot CBYSEP024 labeled incorrectly 1/4 sheets and 1/2 sheets | alan.rousseau@sanofi.com | room 1017 NYA | 2022-07-08 | 2022-07-08 | FBS MBR-000096 | Yes | No Deviation
108 | Leak discovered on FZ SUB Bleed Line | michael.knapp@sanofi.com | 8NYA 1333 | 2022-07-07 | 2022-07-07 | SUB system is integral and there are no leaks. | Yes | Deviation
107 | Bed detection failed during butyl column pack (Add into 575) | Pavel.dobrovolskiy-wynes@sanofi.com | Column Pack *RM1303 | 2022-07-07 | 2022-07-07 | FBL-EWI-00016 (column packing Step 9) | Yes | No Deviation
106 | GEMBA for negative lactate Cedex results due to reagent bottles misplaced | Brandon.Steff@sanofi.com | 8RM1333 | 2022-07-03 | 2022-07-04 | Lactate results negative, bottles 1 and 2 are in correct positions | Yes | No Deviation
105 | Unable to locate BioLIMS unreceived sample 410820-CGHHE008-COV_AX_827589_POU790804_ENDO | cindy.guerra@sanofi.com | 74NYA / PU1-08 | 2022-07-14 | 2022-07-07 | Samples would be received by QC and tested in BioLIMS | Yes | Deviation
104 | Area clearance was not complete for Rm. 168 in 45NYA | charles.hart@sanofi.com | 45NYA Rm 168 | 2022-07-03 | 2022-07-07 | Area clearance per FBL-SOP-000367 step 6.1.5, Are clearance is expected to be documented in MES .
103 | 100L SUB exceeded High high alarm limit for DO on 30JUN22 | adam.hildebrand@sanofi.com | 74 NYA | 2022-06-30 | 2022-07-07 | It's expected that DO would control to setpoint post inoculation
102 | 68NYA-Metrology-no product impact-Metrology Standard S-1570-002 found out of tolerance during a scheduled calibration | robert.bauch@sanofi.com | 68-1830 | 2022-07-01 | 2022-07-07 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 The unit should have produced passing readings
101 | No reportable results for biannual 76 NYA compressed gas sites 76-CA-39, 76-O2-8, 76-CO2-13, and 76-CA-43 | marissa.lentoni@sanofi.com | 76 NYA, 76-CA-39 Room 407, 76-O2-8 Room 407, 76-CO2-13 Room 407, 76-CA-43 Room 347 | 2022-07-01 | 2022-07-01 | Sites would be set up properly for sampling to occur. | Yes | Deviation
100 | Expired sterile gloves used during HA Fill on 06Jul2022 | Craig.Stevenson@sanofi.com | 76 NYA - HA Fill Core | 2022-07-06 | 2022-07-06 | It is expected that no expired sterile gloves be used during open processing
99 | Cerezyme Gen 1 vesivirus sample shipment from Resilience to Geel missing temptale data | elizabeth.couturier@sanofi.com | Geel QC Lab | 2022-06-30 | 2022-07-06 | It is expected that the temptale recorded the temperature for the whole duration of the shipment. | Yes | Deviation
98 | BIOLIMS 231194 and 239239 BGHBN006-CZ2_500_761217_POU760744_BIO and Endo samples not taken | Bhargav.Patel@sanofi.com | 8 NYA / 1403 | 2021-12-17 | 2022-07-07 | FBL-RPT-000538 | Yes | Deviation
97 | No Product – No Lot – 8 NYA – Semi Annual Preventative Maintenance on Watson Marlow Pump, Asset 2009996 at 8 NYA, was not completed prior to due date | sheila.pandolfo@sanofi.com | 8 NYA, Room 1503 | 2022-06-30 | 2022-07-05 | It is expected the PM would be completed by 30JUN2022.
96 | Drip column concentration results out of range (high) during Butyl resin prep testing. | mark.smith@sanofi.com | 8NYA/1303 | 2022-07-06 | 2022-07-06 | Passing result is 40% - 60% per EBR | Yes | Deviation
95 | SUB recipe stuck in transition following prompt to collect harvest material X2220 (Cz) | Brandon.Steff@sanofi.com | 8RM1403 | 2022-07-06 | 2022-07-06 | SUB is able to set harvest SUM to TCU control | Yes | No Deviation
94 | SARS-CoV2 Myco and Viral Samples Receipt Time in LIMS > 24 hours from dispense time | elizabeth.couturier@sanofi.com | 68 NYA Sample Receipt lab | 2022-06-10 | 2022-07-05 | It is expected that Mycoplasma and Viral samples are frozen at </= -60C within 24 hours of collection | Yes | Deviation
93 | SUT HU Not Record In Buffer Record. RF552C001 (8NYA Solution Prep) | Luis.Abrego@sanofi.com | Building 8 / Room 1350 / Department Solution Prep | 2022-06-02 | 2022-06-29 | Necessary information for SUT scanned and record into MES record.
92 | MCC X2510 Load UV meter not stable, unable to zero properly | Brandon.Steff@sanofi.com | 8RM1333 | 2022-07-06 | 2022-07-06 | UV meter is able to zero properly |  | No Deviation
91 | HA - No Lots - 76NYA Room 352 - Failed Leak Test on X-4071 | charlie.houle@sanofi.com | 76 NYA, room 352 | 2022-07-05 | 2022-07-05 | Vaccum Leak Test cycle print out would indicate passing results.
90 | PSFM leak | ivan.hernandez@sanofi.com | CC1-09 | 2022-07-05 | 2022-07-05 | Full 20L amount is expected at 2-8C
89 | Seprafilm - multiple lots - 76 NYA multiple rooms.COVID positive employee in Seprafilm DHT/Sealing | kenneth.morgan@sanofi.com | 76 NYA | 2022-07-02 | 2022-07-05 | It is expected that personnel would not enter classified manufacturing space when they have an illness per FBS-SOP-000667.
88 | Temp Set point on wave x2126 automatically changed to 28C, it should be 27C | tech182@yahoo.com | 74NYA, CC2-01 | 2022-07-02 | 2022-07-02 | Setpoint should never be anything but 27C.
87 | On 04JUL22, X-3051 (MMC skid) experienced a Hi Hi pressure alarm (PI-305012A) causing Capto MMC production recipe to go into hold during sanitization phase. | craig.valery@sanofi.com | 74 NYA, PU1-09 | 2022-07-04 | 2022-07-04 | FBL-MBR-000182 step: 8.11.4
86 | Sample 409688 had an invalid sample CV and failure was not detected during transcription. | roodline.bertrand@sanofi.com | 68NYA,QC MICRO LAB | 2022-06-16 | 2022-07-03 | Per FBL-SOP-000366 site 76AWFI-26A needs daily LAL data. | Yes | Deviation
85 | SP Eluate SUM temp probe put recipe into hold | mark.smith@sanofi.com | 8NYA/1323 Downstream Ballroom | 2022-07-03 | 2022-07-03 | Recipe should not enter hold due to temperature prior to elution start. | No | Deviation
84 | Prompts acknowledged in the incorrect order for Depth Filtration Process | alvaro.rueda@sanofi.com | 74 NYA | 2022-07-03 | 2022-07-03 | Prompts to be acknowledged in the correct order
83 | CV Plated On Wrong Media | gabrielle.cochrane@sanofi.com | 4790 | 2022-07-02 | 2022-07-02 | FBL-SOP-000019, Page 17 ALF-MTD-000018 Page 5, 9.1.8.1. through 9.1.8.16
82 | Atypical UV trend observed during elution of SP1 | subhashini.satishkumar@sanofi.com | 51NYA, Rooms 138 | 2022-07-02 | 2022-07-02 | TBD
81 | Leak from Colder connection during E4 harvest. | james.letourneux@sanofi.com | 51NYA Room 140 | 2022-07-02 | 1330 | No leaks should occur during harvest.
80 | Possible Contamination of Media used for Bioburden | jessica.kychler@sanofi.com | QC Microbiology, 68 NYA floor 4 room 4700 | 2022-07-02 | 2022-07-02 | It is expected that media used for bioburden testing will be free of contamination | Yes | No Deviation
79 | RTU-3 went down at 51 NYA affecting rooms 130, 130A,130B, cold room 131. | Sylvester.Kegbeh@sanofi.com | 51NYA- Rooms 130, 130A,130B, cold room 131. | 2022-07-02 | 2022-07-02 | Air handler unit working properly | Yes | Deviation
78 | Quarterly Audits issued by Records Management not performed on time | marissa.cooney@sanofi.com | 74 and 76 | 2022-07-01 | 2022-07-01 | FBL-SOP-000441 section 6.8, steps 6.8.9-6.8.19 | Yes | Deviation
77 | Bubble Trap HiHi alarm consistently during skid clean recipe prime step | mark.smith@sanofi.com | 8NYA/1303 | 2022-07-01 | 2022-07-01 | Priming should complete with bubble trap level being maintained between Hi and HiHi sensors via recipe.
76 | Failed Effectiveness Reviews for FRAD20R0091 & FRAD20R0080 | richard.quitadamo@sanofi.com | 74 and 8 NYA | 2022-06-30 | 2022-06-30 | BQ-GOP-000249, ERs expected to close with CAPAs being effective
75 | 76 PURIFICATION AREA BADGE ACCESS AUDIT not completed per FBL-SOP-000083 | jacob.kallon@sanofi.com | 76 NYA, JIT room | 2022-06-30 | 2022-07-01 | Per FBL-SOP-000083, a quarterly badge access audit should be performed for 76 NYA JIT room
74 | OOS ELISA Thyrogen Harvest Fluid H-01 Lot CGF4B07 | paul.consolmagno@sanofi.com | 68NYA, QCM | 2022-06-30 | 2022-06-30 | Harvest fluid tested per FBL-MTD-000042 would meet specification as defined in FBL-SPC-000090 | Yes | Deviation
73 | GEMBA: 74 NYA - The RQ for US1254 past due (30JUN22) | william.bradley@sanofi.com | 74NYA | 2022-07-01 | 2022-07-01 | RQ - F074-PH-74NYA-CQ-0001-Process hoses using Utility Station US-1254 Circuit 1-40 | Yes | Deviation
72 | 74NYA-74-CC1-27-GOWNING A/L REACTOR SUITE-Falcon-Product impact unknown- FB-0453 HEPA filters found out of tolerance for Penetration during a scheduled calibration | robert.bauch@sanofi.com | 74-CC1-27 | 2022-07-28 | 2022-06-30 | FBL-SOP-000698 Step 6.12.2
71 | Fz Nano Post viral MES record Issue | Michael.Daly@sanofi.com | 8NYA RM1313 | 2022-06-30 | 2022-06-30 | The record would have branched to the steps that needed to be executed. | Yes | Deviation
70 | Missing entry for parts washer logbook | william.lee@sanofi.com | 74NYA | 2022-05-12 | 2022-06-30 | The entry should be documented as a pass
69 | No reportable hydrocarbon results for compressed air site 74-CA-01 sampled on 26Mar2022 | marissa.lentoni@sanofi.com | 74 NYA, Room UT1B02 (Basement), 74-CA-01 | 2022-03-26 | 2022-06-30 | all samples would be collected and results entered for 74-CA-01 at time of sampling | Yes | Deviation
68 | 74 NYA PQ Gap on Nitrogen Line Modification | tracy.chaves@sanofi.com | 74 NYA CC2-01 | 2022-06-10 | 2022-06-29 | Line would have been PQ'd prior to use
67 | 8NYA-CC8 room 1333-product impact unknown-FB-2735-002,FB-2735-004,FB-2735-006,FB-2735-008 on the wave reactor had no as found data taken | robert.bauch@sanofi.com | 8-1333 | 2022-06-30 | 2022-06-30 | BQ-GOP-000046 Calibration Program Version 4 Section 5.6.8 as found should have been able to be collected | Yes | Deviation
66 | Take out step not performed in RF552C001 | richard.manasian@sanofi.com | 8 NYA/ 1350/ Solution Prep | 2022-06-02 | 2022-06-29 | RF552C001 tote bag take out in transfer PU |  | No Deviation
65 | No reportable results for 74 NYA routine EM performed on 09Jan2022 | marissa.lentoni@sanofi.com | 74 NYA PU1-05-09, PU1-05-10, PU1-05-11 | 2022-01-09 | 2022-06-29 | Paperwork would be completed as the same time as sampling occurred. Incubation of plates would be initiated same day as sampling. | Yes | Deviation
64 | Gemba for count on aseptic qualification | roodline.bertrand@sanofi.com | 68NYA,QC MICRO LAB | 2022-06-21 | 2022-06-29 | 0 CFU per FBL-SOP-00755 | Yes | Deviation
63 | Logbook FBL-LBK-000285 has same operator as performed by and witnessed by | dianne.moustafa@sanofi.com | 74NYA CC1-15 Equipment prep | 2022-05-03 | 2022-06-28 | Two operators sign for steps
62 | Thyrogen- Leak During Harvest on H14 | christopher.holmes@sanofi.com | 51-140 | 2022-06-28 | 2022-06-28 | It is expected that no leaks would be present during the harvest. | Yes | Deviation
61 | HA-CMC - CBYCMC136 - 76NYA Room 229 - Ethanol waste system shutdown during decant. | jan.stepanik@sanofi.com | 76NYA/room 229 | 2022-06-28 | 2022-06-28 | Ethanol waste stream should have remained functional during decant. | No | Deviation
60 | Thyrogen- pCO2 Out of Range for Bioreactor T5 | christopher.holmes@sanofi.com | 45-140 | 2022-06-28 | 2022-06-28 | It is exepcted that the value would return in the range of 100-230 |  | No Deviation
59 | Missing logbooks at 76 NYA | subhashini.satishkumar@sanofi.com | 76 NYA | 2022-06-16 | 2022-06-16 | It is expected that all logbooks are available for quarterly audits
58 | Ineffective Effectiveness Review for CAPA FRAD21A0423 found from Central Inventory Audit | kelly.alanmurphy@sanofi.com | 8NYA, 45NYA, 51NYA | 2022-06-27 | 2022-06-27 | Per FBL-SOP-000128 Version 8, step 6.11.9, If parts need to be returned to Central Inventory, change status to “Tech Complete – Parts to Return”. Return parts to pick up location and complete the Business Process form documenting reason for return. | Yes | Deviation
57 | 76-231 Low temp EMS | john.beaule@sanofi.com | 76-231 | 2022-06-27 | 2022-06-27 | Temp to be within 20-25C
56 | CB6 Media bag does not have 2 sterile sample bags although 2 are required for POU sampling. | adam.hildebrand@sanofi.com | 74 NYA | 2022-06-28 | 2022-06-28 | Its expected to have stronger guidance in the SOP how to pull the sample and the logistics of said samples. | Yes | Deviation
55 | Action Level result for Trainer during Aseptic Gowning Qualification in the Fill/Finish Area at 76NYA. | Craig.Stevenson@sanofi.com | 76 New York Ave | 2022-06-21 | 2022-06-28 | It is expected that the trainer generate passing and valid results. | Yes | Deviation
54 | 3D Tote bag 50L miss labeled Is not correct bag. Part 692187 Lot 21GN01385 74 NYA falcon media | dianne.moustafa@sanofi.com | 74 NYA Media CC1-08 | 2022-06-28 | 2022-06-28 | Material be labeled correctly
53 | Thyrogen- 45 NYA Buffer GZ5026 Post Use Filter Failure | christopher.holmes@sanofi.com | 45 NYA Room 190 | 2022-06-24 | 2022-06-27 | It is expected that a test would stabilize at step 6.4.2.6 of FBL-SOP-001011 | Yes | Deviation
52 | 8 NYA - Lots TBD - Employee Tested Positive for COVID-19 who entered 8 NYA | sheila.pandolfo@sanofi.com | 8 NYA, Rooms 1350, 1373 and 1290 | 2022-06-24 | 2022-06-27 | Covid positive personnel do not enter controlled areas.
51 | 74NYA-74-CO1-10 Women's air lock-Falcon-Product impact unknown- FB-0443 HEPA filter found out of tolerance for Penetration during a scheduled calibration | robert.bauch@sanofi.com | 74-CO1-10 Women's air lock | 2022-06-27 | 2022-06-28 | FBL-SOP-000698 Step 6.12.2
50 | Thyrogen – Lot CGFSB008 – 51 NYA - pCO2 BGA offline sample above normal operating range of 30-100mmHg during 12L Wave Reactor Day 1 | sheila.pandolfo@sanofi.com | 51 NYA Room 152 | 2022-06-25 | 2022-06-25 | It is expected the pCO2 BGA offline sample result would be within the normal operating range (30-100mmHg) on the 12L Wave (Day 1).
49 | Inconsistency between INFOR/EDMS/FBL-SOP-000336 Attachment 2 page 3 of 5 | brett.mendenhall@sanofi.com | 51 NYA Room 140 | 2022-06-24 | 2022-06-24 | FBL-SOP-000336 and Drawing F051-AR-5500.DWG would align with the asset information in INFOR
48 | conductivity tested twice for 2 sites | Stefanie.Erickson@sanofi.com | QC Micro water chemistry bay | 2022-06-25 | 2022-06-27 | for there to be 1 result for conductivity | Yes | Deviation
47 | 74 NYA - Contractor Tested Positive for Covid-19 who entered 74 NYA GMP/MFG Area | sheila.pandolfo@sanofi.com | 74 NYA, Room PU1-22 | 2022-06-22 | 2022-06-26 | Covid positive personnel do not enter controlled areas. | No | Deviation
46 | COVID positive QCM Employee - 76 NYA – Biologics R&D, 76 NYA – R&D. 8 NYA | Craig.Stevenson@sanofi.com | 76 New York Ave, 8 New York Ave | 2022-06-23 | 2022-06-27 | Covid positive personnel do not enter controlled areas. | No | Deviation
45 | COM6/COM7 pCO2 and DO trends out of range | adam.hildebrand@sanofi.com | 74 NYA | 2021-12-25 | 2022-06-27 | FBL-CST-000030/ 70mmhg to 120mmhg KPP; FBL-CST-000030/ DO 60% +/- 5% KPP | Yes | Deviation
44 | Buffer Record FBL-MBR-000142, Batch# CGHBP039 not fully reviewed prior to use | alvaro.rueda@sanofi.com | 74 NYA | 2022-06-22 | 2022-06-27 | All page reviews of the buffer batch record should have been reviewed prior to signing step 9.13.7 | Yes | No Deviation
43 | SWECO Moisture Sensor failure during HA Bulk lot CBYHAH007 | nicholas.macdonald@sanofi.com | 76-3103 | 2022-06-25 | 2022-06-25 | TBD
42 | pH Chromatogram out of parameter during SP1 (Thyrogen) | gregory.klotz@sanofi.com | 51 NYA room 138 | 2022-06-10 | 2022-06-24 | Expect the ph trace to be withing the parameter range defined by FBL-SOP-000788
41 | Incorrect calculation performed during execution of FBL-MTD-000014 for Fabrazyme 1G | guy.francois@sanofi.com | QC Lab 45 NYA | 2022-06-24 | 2022-06-24 | Per FBLMTD-000014 step 6.9.4 the R squared value generated from a no- linear Michaelis-Menten curve. The R2 value must be ≥ 0.985. | Yes | Deviation
40 | Falcon Samples for SDS/Western Blot Overfilled, caps open during freeze/storage and leaking during thaw. | john.kapaniris@sanofi.com | QC Chemistry Laboratory 2nd floor 68 New York Avenue | 2022-06-23 | 2022-06-23 | N/A | No | Deviation
39 | Non-reportable non-viable result for EM site: CC2-83 and CC2-85 from 08Jan22 | tae.kang@sanofi.com | 74NYA/ CC2-01/03 | 2022-01-08 | 2022-06-22 | Sites CC2-83 and CC2-85 (both in room: CC2-01/03) should have been sampled separately on 08Jan22. | Yes | No Deviation
38 | Myco/Viral Retain Sample Freeze time was not documented | elizabeth.couturier@sanofi.com | 68 NYA, Sample Receipt | 2022-06-12 | 2022-06-15 | expected that samples are frozen w/in 24 hours | No | No Deviation
37 | Non-reportable results for 74-HWFI-1B 09JAN22 | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2022-01-09 | 2022-06-22 | Labels and test requests match and state the correct sample type. | Yes | Deviation
36 | Phase 2 required: OOL Endotoxin CV, Cz MCC SKID X2811 CIP RINSATE GLYCOL WASTE OUTLET, QC#: 22MS01283, L/N CGHAA006 | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2022-06-02 | 2022-06-02 | Per FBL-FRM-000776, the result generated should be <0.25EU/mL | Yes | Deviation
35 | Temperature excursion for QCTS freezers 40-QC-0495 and 40-QC-0638 | jacqueline.legere@sanofi.com | 68NYA Lab 3700 | 2022-06-16 | 2022-06-16 | Per FBL-SOP-001217, the temperature must be between -50C and -90C. | Yes | No Deviation
34 | Bioburden results (CFU) not documented during plate counting for Bulk HA Pre-Filtration Bioburden sample (Part#: 5615-20, Lot#: CBYHAH006, Manual Test Request#: 22MS01330) | tae.kang@sanofi.com | 68NYA/4700 | 2022-06-11 | 2022-06-16 | Bioburden results (CFU) should have been documented contemporaneously | Yes | Deviation
33 | pH Meter F-17806 Incorrect Settings | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2022-06-08 | 2022-06-15 | FBL-SOP-000351, 20-1424_75-QC-0147_148_QCM_pH Meters CV Protocol_FBL-FRM-000937 | No | Deviation
32 | Mannitol Assay performed by Analyst Kamlesh Patel while not trained on the current version of HPLC Ops and Maint. SOP | john.kapaniris@sanofi.com | QC Chemistry Laboratory 45 NYA | 2022-06-09 | 2022-06-16 | Trained in iLearn on V 7.0 of FBL-MTD-000178 | No | No Deviation
31 | 8-WFI2-2 was sampled while under action notice | marissa.lentoni@sanofi.com | 8 NYA, Room 1170 | 2022-06-13 | 2022-06-16 | FBL-SOP-000336, Attachment 3 page 1 addresses access restrictions for a tier 2 CU action notice | Yes | Deviation
30 | Shipment of Cerezyme Lot CAL0137 to Geel for Vesivirus testing exceeded the 120 hour limit | QCComplianceSupport@sanofi.com | Sanofi, Geel QC lab | 2022-06-14 | 2022-06-14 | it is expected per ALF-SOP-000880 step 9.7.1 that the E200 shipper not exceed 120 hours. | Yes | Deviation
29 | Multiple testing was performed, generating multiple valid data for microbial identification for Isolate F22-0314-1 (from 22FID0314) (Sample description: LIMS#: 365861 COV_MMC_827588_ELUATE_BIOB SDA Plate) | tae.kang@sanofi.com | 68NYA/4700 | 2022-06-02 | 2022-06-13 | Isolate F22-0314-1 (from 22FID0314) should not have been tested multiple times, as the first run generated valid result. | Yes | Deviation
28 | Two Falcon samples were submitted to QC Micro lab with short expiration times | daria.oberlander@sanofi.com | 68NYA,/room 4700/QC Microbiology | 2022-06-10 | 2022-06-11 | FBL-SOP-000019, all samples have 24-hour expiration time | No | Deviation
27 | Bioburden sample received leaking | Pamela.Lessard@sanofi.com | 68NYA\4720 | 2022-06-10 | 2022-06-10 | Sample would be able to be received and put on test | No | Deviation
26 | Dual testing performed for microbial identification for Isolate F22-0283-2 (from 22FID0283) (Sample description: Swab, TP-2632-C, 22MS01075 (05May22)) | tae.kang@sanofi.com | 68NYA/4700 | 2022-06-08 | 2022-06-09 | Isolate F22-0283-2 (from 22FID0283) should not have been dual tested. | Yes | Deviation
25 | Heat Block Incubation time for rhTSH UFB Stability Day 10 6-10°C Lot BGF1S11 was not recorded within the specified timeframe of 120 ±10 minutes per FBL-MTD-000051 Carbohydrate Analysis | john.kapaniris@sanofi.com | 45 New York Avenue Quality Control Chemistry Lab | 2022-02-28 | 2022-06-07 | FBL-MTD-000051 Step 6.7.3 requires incubation time of 120 ±10 minutes for Hydrolysis | Yes | Deviation
24 | INC-02364/Collection Batch: 202201921-1 of Environmental Monitoring samples received more than 24 hours after sampling | paul.consolmagno@sanofi.com | F-16387, 51 NYA; QCM, 68 NYA | 2022-06-06 | 2022-06-06 | The EM samples would be put on test/incubated before the 24 hours expiration | No | Deviation
23 | Aseptic Qualification (2200032261-51 NYA LS Cell Culture Grade A Aseptic Qualification ) incubation time not recorded | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2022-05-28 | 2022-05-31 | FBL-SOP-000476 (Responsibilities-QC), FBL-SOP-000873 Attachment 2 | Yes | No Deviation
22 | COVID positive QCM Employee - 74 NYA 2nd floor lab area | elizabeth.couturier@sanofi.com | 74 NYA 2nd floor lab area | 2022-05-25 | 2022-05-26 | unknown | No | Deviation
21 | 2 QCM EM Analysts performed EM without EM OJT completed | Craig.Stevenson@sanofi.com | 11 Forbes Rd / QCRM | 2022-05-10 | 2022-05-25 | Per BQ-GOP-000009 V8.0 step 5.4.1, Personnel must complete and document training prior to independently performing any related GMP task | Yes | Deviation
20 | Phase 2 required: OOL LAL Fabrazyme Capto Butyl P/N: 762729 L/N: CGHBE013 OneLIMS: 363332 | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2022-05-18 | 2022-05-25 | Per FBL-LRA-000027 V4.0, the LAL alert level is >0.2EU/mL and should generate results below the alert level. | Yes | Deviation
19 | SDA Plate (GZ95001154), Lot# 22GN00763 , Expiry: 13Sep22 found contaminated with mold before opening the pack | tae.kang@sanofi.com | 68NYA/4700 | 2022-05-23 | 2022-05-23 | Per NBO-SPC-001133, the media is released based on meeting specification, which includes media meeting sterile claim | Yes | Deviation
18 | dual incubation not performed for 2 EM batches | Stefanie.Erickson@sanofi.com | 51 NYA and 74 NYA, QC Micro counting bench | 2022-05-20 | 2022-05-22 | for the incubation to be ended in F-15593 and started in F-15594 in biolims | Yes | No Deviation
17 | Incorrectly tested TOC samples 2200029353 and 2200029477 | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2022-05-12 | 2022-05-18 | FBL-SOP-000595 Step 6.4.8 incorrect parameters; step 6.4.7 correct parameters | Yes | Deviation
16 | Myco and viral listed below were not frozen within 24 hours | elizabeth.couturier@sanofi.com | 68 NYA, Sample Receipt room | 2022-05-18 | 2022-05-18 | it is expected per above SOPs that viral/myco retains are frozen within 24 hours | No | Deviation
15 | LAL Buffer Sample tested above minimum validated dilution | kyle.lebherz@sanofi.com | 68NYA QCM Lab | 2022-05-16 | 2022-05-17 | To test the sample at 1:40 per FBL-SOP-000417 table 2 matrix | Yes | No Deviation
14 | Missed bioburden verification checklist for Cz POU media 760744 | judith.keller@sanofi.com | 68 NYA / 4700/ QCM | 2022-05-14 | 2022-05-16 | ALF-MTD-000018 step 9.2.3.2- checklist needs to be filled out | Yes | No Deviation
13 | No reportable bioburden results for 51-RO-5 and no reportable LAL results for 8-WFI2-3 | marissa.lentoni@sanofi.com | 68 NYA/4700/QC Micro lab | 2022-01-19 | 2022-05-13 | 51-RO-5 requires monthly bioburden testing per FBL-SOP-000189. 8-WFI2-3 requires daily LAL testing per FBL-FRM-000819. | Yes | Deviation
12 | Incorrect amount of Raw material bioburden (Part #: 1336-02:) powder weighed out for testing. | tae.kang@sanofi.com | 68NYA/4700 | 2022-04-13 | 2022-05-12 | Per FBL-SOP-000937, Step 6.1: 1 gram ± 0.01 of the powder of the raw material must be weighed out. | Yes | Deviation
11 | Total Protein Content (BCA) Results were reported incorrectly | elizabeth.couturier@sanofi.com | 45 NYA QC Chem | 2022-05-09 | 2022-05-10 | It is expected that TPC is reported in microgram/mL for these samples. | Yes | No Deviation
10 | Bioburden 2022-CV-0201 (300L Portable Vessel V-3424 CIP) tested >24 Hours | Craig.Stevenson@sanofi.com | 68NYA / Room 4700 / QCM | 2022-04-30 | 2022-05-10 | Per FBL-SOP-000019, Bioburden must be completed within 24 hours of sampling. | Yes | No Deviation
9 | Seprafilm incubation end date not recorded at the time of counting | elizabeth.couturier@sanofi.com | 68 NYA, QCM Lab | 2022-04-30 | 2022-05-09 | to mark when it was removed from incubators | Yes | No Deviation
8 | LAL NEXUS sample run with TOC label | elizabeth.couturier@sanofi.com | 68 NYA, QCM Lab | 2022-05-08 | 2022-05-09 | for the label to be on TOC | Yes | No Deviation